Gel Documentation Form and Worksheet

HLA-A*30 (101.429-12/12u) 	Lot No: 38V	Expiry Date: 2016-August-01

Name:__________________

Sample ID:______________

DNA Conc.(ng/ul):_________

Test  Date:_______________

Tested By:_______________

Review Date:_______________

Reviewed By:_______________

Gel Picture

PHOTO DOCUMENT

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

Interpretation:_______________

Failed lanes:    _________

Comments:    ___________________

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 4, 12 and 14 may give rise to nonspecific amplifications, most pronounced in primer mixes 12 and 14.

Primer mix 24 may have a tendency of giving rise to primer oligomer formation.

Primer mixes 3 and 8 may give rise to a lower yield of HLA-specific PCR product than the other A*30 primer mixes.

Primer mix 32 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by control primer pairs. PCR product sizes range from 75 to 200 base pairs. The PCR product generated by the control primer pair is 430 base pairs.

1HLA-A*30 alleles in bold lettering are listed as confirmed alleles on the on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.14.0, October 2013.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3Primer mix 5: Specific PCR fragment of 150 bp in the A*30:04:01-30:04:02, 30:06, 30:17, 30:29, 30:46 and 30:77 and the A*02:52, 03:82, 24:66, 68:06 alleles. Specific PCR fragment of 245 bp in the A*30:19 alleles.

Primer mix 6: Specific PCR fragment of 80 bp in the A*30:06 allele. Specific PCR fragment of 185 bp in the A*30:07 allele.

Primer mix 18: Specific PCR fragment of 130 bp in the A*30:32 allele. Specific PCR fragment of 210 bp in the A*30:20 and 30:71 alleles.

Primer mix 19: Specific PCR fragment of 125 bp in the A*30:70N allele. Specific PCR fragment of 160 bp in the A*30:22 allele. Specific PCR fragment of 200 bp in the A*30:31 allele. Specific PCR fragment of 235 bp in the A*30:77 and the A*03:45 alleles.

Primer mix 20: Specific PCR fragment of 210 bp in the A*30:23 allele. Specific PCR fragment of 275 bp in the A*30:34 and the A*03:04:02-03:04:03 and 11:153:01 alleles.

Primer mix 21: Specific PCR fragment of 155 bp in the A*30:30 allele. Specific PCR fragment of 185 bp in the A*30:24 allele.

Primer mix 22: Specific PCR fragment of 180 bp in the A*30:36 and 30:73N alleles. Specific PCR fragment of 215 bp in the A*30:27N allele.

Primer mix 27: Specific PCR fragment of 145 bp in the A*30:56 allele. Specific PCR fragment of 195 bp in the A*30:45 allele.

Primer mix 29: Specific PCR fragment of 130 bp in the A*30:59N, 30:61 and 30:74 and the A*01:57N, 02:156, 02:338, 03:17:01, 03:171, 11:119:01-11:119:02, 24:92 and 68:103 alleles. Specific PCR fragment of 190 bp in the A*30:73N allele.

4This lot of the HLA-A*30 subtyping kit cannot distinguish the A*30:01:01-30:01:07 alleles and the A*30:78N allele.

’w’, might be weakly amplified.

Change in revision R01 compared to R00:

- Primer mix 24 does not amplify the A*30:78N allele. This has been corrected in the Specificity and Interpretation tables. Thus, this lot of the HLA-A*30 subtyping kit cannot distinguish the A*30:01:01-30:01:07 alleles and the A*30:78N allele.