PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® HLA-A*31

Prod. No:	101. 430-12/12u

Lot No:    	89N

Expiry Date:    	2015-January-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 32 primer mixes used for HLA-A*31 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified HLA-A*31 alleles3 | Other amplified HLA-A alleles4 |
| --- | --- | --- | --- | --- |
| 1 | 155 bp | 800 bp | *31:01:02-31:07, 31:09-31:43, 31:45-31:62 | *01:07, 02:185, 23:21, 24:124, 26:19, 29:14, 30:12, 30:18, 30:55, 34:04, B*15:82, B*40:186 |
| 2 | 215 bp | 800 bp | *31:02, 31:07-31:08 | *02:243, 24:82, 33:08, 33:53 |
| 3 | 155 bp | 800 bp | *31:03-31:04 | *02:309, 03:01:19, 26:43:01, 34:02:01, 34:02:03-34:04, 34:06-34:09, 66:06 |
| 49 | 165 bp | 1070 bp | *31:03-31:04, 31:06 | *01:06, 02:114, 02:246, 02:279, 03:01:30, 03:05, 03:42, 03:98, 03:105, 03:122, 11:24:01-11:25, 11:31, 11:35, 29:01:01:01-29:01:02, 29:02:01:01-29:04, 29:06-29:23, 29:25-29:27, 29:29-29:31, 29:34-29:35, 30:26, 32:30, 32:32, 33:18, 34:02:01-34:04, 34:07-34:09, 68:08:01-68:08:02, 68:63, 80:01-80:02 |
| 55,10 | 110 bp, 135 bp, 285 bp | 800 bp | *31:03, 31:26, 31:38 | *11:43, 32:36, 33:13 |
| 65,11 | 130 bp, 165 bp | 800 bp | *31:05, 31:39 | *23:03:01, 29:03, 29:33, 32:13, 33:10 |
| 79 | 505 bp | 1070 bp | *31:07-31:08, 31:10 | *02:81, 02:87, 02:112, 02:124, 02:129, 23:01:01-23:01:09, 23:03:01-23:13, 23:15-23:26, 23:28-23:33, 23:35-23:37, 23:39-23:51, 24:13:01, 24:18, 24:24, 24:94, 24:188, 29:13, 32:01:01-32:01:05, 32:01:07-32:01:13, 32:03-32:09, 32:11Q-32:21, 32:23-32:41 |
| 86,12 | 155 bp, 220 bp | 1070 bp | *31:09, 31:24, 31:27 | *11:01:28 |
| 95,6 | 75 bp | 1070 bp | *31:01:02-31:02, 31:05, 31:07-31:61 | *02:24:02, 02:65, 02:152, 23:03:01, 29:32, 32:01:01-32:01:06, 32:01:08-32:01:11, 32:01:13-32:03, 32:05-32:27N, 32:29, 32:31, 32:33-32:41, 33:01:01-33:01:04, 33:01:06, 33:03:01-33:17, 33:20-33:31, 33:33-33:59, 74:01-74:15 |
| 105,8,13 | 115 bp, 160 bp | 1070 bp | *31:11, 31:41 | *02:24:02, 29:33, 32:02, 32:06 |
| 1114 | 135 bp, 210 bp | 1070 bp | *31:12, 31:26, 31:60N | *32:36 |
| 12 | 245 bp | 1070 bp | *31:01:02-31:06, 31:09, 31:11-31:20, 31:22-31:32, 31:34-31:62 | *02:243, 11:43, 29:19, 33:01:01-33:01:06, 33:03:01-33:12, 33:14-33:16, 33:18-33:47, 33:49-33:50, 33:52-33:59, 68:29 |
| 135 | 85 bp | 1070 bp | *31:13 | *02:251 |
| 14 | 150 bp | 800 bp | *31:24 | *11:01:28 |
| 1515 | 150 bp, 225 bp | 1070 bp | *31:15, 31:25 | *32:26 |
| 165,16 | 80 bp, 165 bp | 1070 bp | *31:16, 31:46 | *29:12, 33:58 |
| 1717 | 155 bp, 235 bp | 1070 bp | *31:17, 31:25, 31:40 | *32:26 |
| 18 | 200 bp | 1070 bp | *31:18 | *23:43 |
| 195,8 | 110 bp | 1070 bp | *31:19 | *03:52 |
| 20 | 165 bp,325 bp | 1070 bp | *31:56 |  |
| 21 | 180 bp | 1070 bp | *31:21 | *01:07, 02:185, 30:55 |
| 2219 | 150 bp, 190 bp | 1070 bp | *31:22, 31:27, 31:55 |  |
| 2320 | 165 bp, 200 bp | 1070 bp | *31:23, 31:40 |  |
| 247,21 | 150 bp, 220 bp | 1070 bp | *31:28, 31:55 |  |
| 2522 | 135 bp, 245 bp | 800 bp | *31:29, 31:59 | *23:03:01, 29:07 |
| 265 | 130 bp | 1070 bp | *31:30, 31:39 | *29:28, 32:10, C*02:02:15 |
| 275 | 95 bp | 800 bp | *31:31, 31:44 |  |
| 285,23 | 110 bp, 175 bp, 205 bp | 1070 bp | *31:32, 31:38, 31:60N |  |
| 295,24 | 115 bp, 275 bp | 1070 bp | *31:36, 31:48 | *02:140, 33:15 |
| 307,25 | 160 bp, 245 bp | 800 bp | *31:33-31:34 |  |
| 315 | 75 bp | 1070 bp | *31:35 | *01:07 |
| 325,26 | 75 bp, 115 bp | 1070 bp | *31:37, 31:41 | *02:24:02, 29:33, 32:02, 32:06 |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-A*31 SSP typings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective lengths of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most wells, or a band of 800 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-A*31 subtyping.

In addition, wells number 2, 3, 5, 6, 14, 25, 27 and 30 contain the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3For several HLA-A*31 alleles 4th exon or intron nucleotide sequences are not available. In these instances it is not known whether some of the primers of the SSP sets are completely matched with the target sequences or not. We assume that unknown sequences in these regions are conserved within allelic groups.

4Due to the sharing of sequence motifs between HLA-A alleles non-HLA-A*31 alleles will be amplified by primer mixes 1 to 19, 21, 25, 26, 29, 31 and 32. In addition, primer mix 1 will amplify the B*15:82 and B*40:186 alleles and primer mix 26 will amplify the C*02:02:15 allele.

5Short specific PCR fragments are less intense and not as sharp as longer specific bands.

6Primer mixes 8 and 9 may give rise to a lower yield of HLA-specific PCR product than the other HLA-A*31 primer mixes, most pronounced for the A*31:09 allele in primer mix 8.

7Primer mix 24 and 30 may give rise to nonspecific amplifications.

8Primer mix 10 and 19 may give rise to primer oligomer formation.

9Primer mixes 4 and 7 may weakly amplify the A*34:01 allele.

10Primer mix 5: Specific PCR product of 110 bp in the A*31:38 allele. Specific PCR product of 135 bp in the A*31:26 and the A*32:36 allele. Specific PCR product of 285 bp in the A*31:03 and in the A*11:43 and A*33:13 alleles.

11Primer mix 6: Specific PCR product of 130 bp in the A*31:39 allele. Specific PCR product of 165 bp in the A*31:05 and in the A*23:03:01, 29:03, 29:33, 32:13 and 33:10 alleles.

12Primer mix 8: Specific PCR product of 155 bp in the A*31:24 and 31:27 and the A*11:01:28 alleles. Specific PCR product of 220 bp in the A*31:09 allele.

13Primer mix 10: Specific PCR product of 115 bp in the A*31:41 and the A*02:24:02, 29:33, 32:02 and 32:06 alleles. Specific PCR product of 160 bp in the A*31:11 allele.

14Primer mix 11: Specific PCR product of 135 bp in the A*31:26 and the A*32:36 allele. Specific PCR product of 210 bp in the A*31:12 and 31:60N alleles.

15Primer mix 15: Specific PCR product of 150 bp in the A*31:25 and the A*32:26 alleles. Specific PCR product of 225 bp in the A*31:15 allele.

16Primer mix 16: Specific PCR product of 80 bp in the A*31:46 allele. Specific PCR product of 165 bp in the A*31:16 and the A*29:12 and 33:58 alleles.

17Primer mix 17: Specific PCR product of 155 bp in the A*31:25 and 31:40 and the A*32:26 allele. Specific PCR product of 235 bp in the A*31:17 allele.

19Primer mix 22: Specific PCR product of 150 bp in the A*31:27 and 31:55 alleles. Specific PCR product of 190 bp in the A*31:22 allele.

20Primer mix 23: Specific PCR product of 165 bp in the A*31:40 allele. Specific PCR product of 200 bp in the A*31:23 allele.

21Primer mix 24: Specific PCR product of 150 bp in the A*31:55 allele. Specific PCR product of 220 bp in the A*31:28 allele.

22Primer mix 25: Specific PCR product of 135 bp in the A*31:29 and the A*23:03:01 and 29:07 alleles. Specific PCR product of 245 bp in the A*31:59 allele.

23Primer mix 28: Specific PCR product of 110 bp in the A*31:38 allele. Specific PCR product of 175 bp in the A*31:32 allele. Specific PCR product of 205 bp in the A*31:60N allele.

24Primer mix 29: Specific PCR product of 115 bp in the A*31:48 and the A*02:140 and 33:15 alleles. Specific PCR product of 275 bp in the A*31:36 allele.

25Primer mix 30: Specific PCR product of 160 bp in the A*31:34 allele. Specific PCR product of 245 bp in the A*31:33 allele.

26Primer mix 32: Specific PCR product of 75 bp in the A*31:37 allele. Specific PCR product of 115 bp in the A*31:41 and the A*02:24:02, 29:33, 32:02 and 32:06 alleles.

1The sequence of the A*31011 allele has been shown to be identical to A*31:01:02.

2HLA-A*31 alleles in bold lettering are listed as confirmed alleles on the on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.8.0, April 2012.

3The A*31:16 and A*31:46 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 16.

5The A*31:29 and A*31:59 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 25

6The A*31:36 and A*31:48 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 29.

7Primer mix 5: Specific PCR product of 110 bp in the A*31:38 allele. Specific PCR product of 135 bp in the A*31:26 and the A*32:36 allele. Specific PCR product of 285 bp in the A*31:03 and in the A*11:43 and A*33:13 alleles.

Primer mix 6: Specific PCR product of 130 bp in the A*31:39 allele. Specific PCR product of 165 bp in the A*31:05 and in the A*23:03:01, 29:03, 29:33, 32:13 and 33:10 alleles.

Primer mix 8: Specific PCR product of 155 bp in the A*31:24 and 31:27 and the A*11:01:28 alleles. Specific PCR product of 220 bp in the A*31:09 allele.

Primer mix 10: Specific PCR product of 115 bp in the A*31:41 and the A*02:24:02, 29:33, 32:02 and 32:06 alleles. Specific PCR product of 160 bp in the A*31:11 allele.

1Primer mix 11: Specific PCR product of 135 bp in the A*31:26 and the A*32:36 allele. Specific PCR product of 210 bp in the A*31:12 and 31:60N alleles.

Primer mix 15: Specific PCR product of 150 bp in the A*31:25 and the A*32:26 alleles. Specific PCR product of 225 bp in the A*31:15 allele.

Primer mix 16: Specific PCR product of 80 bp in the A*31:46 allele. Specific PCR product of 165 bp in the A*31:16 and the A*29:12 and 33:58 alleles.

Primer mix 17: Specific PCR product of 155 bp in the A*31:25 and 31:40 and the A*32:26 allele. Specific PCR product of 235 bp in the A*31:17 allele.

Primer mix 22: Specific PCR product of 150 bp in the A*31:27 and 31:55 alleles. Specific PCR product of 190 bp in the A*31:22 allele.

Primer mix 23: Specific PCR product of 165 bp in the A*31:40 allele. Specific PCR product of 200 bp in the A*31:23 allele.

Primer mix 24: Specific PCR product of 150 bp in the A*31:55 allele. Specific PCR product of 220 bp in the A*31:28 allele.

Primer mix 25: Specific PCR product of 135 bp in the A*31:29 and the A*23:03:01 and 29:07 alleles. Specific PCR product of 245 bp in the A*31:59 allele.

Primer mix 28: Specific PCR product of 110 bp in the A*31:38 allele. Specific PCR product of 175 bp in the A*31:32 allele. Specific PCR product of 205 bp in the A*31:60N allele.

Primer mix 29: Specific PCR product of 115 bp in the A*31:48 and the A*02:140 and 33:15 alleles. Specific PCR product of 275 bp in the A*31:36 allele.

Primer mix 30: Specific PCR product of 160 bp in the A*31:34 allele. Specific PCR product of 245 bp in the A*31:33 allele.

Primer mix 32: Specific PCR product of 75 bp in the A*31:37 allele. Specific PCR product of 115 bp in the A*31:41 and the A*02:24:02, 29:33, 32:02 and 32:06 alleles.

11This lot of the HLA-A*31 subtyping kit cannot distinguish the A*31:01:02-31:01:13, 31:42-31:43, 31:45, 31:47, 31:49-31:54, 31:57-31:58 and 31:61 alleles and the A*31:14N allele.

This lot of the HLA-A*31 subtyping kit cannot distinguish the A*31:01:02-31:01:17 and the A*31:20 alleles.

Change in revision R01 compared to R00:

- Primer mix 14 does not amplify the A*31:14N allele. This has been corrected in the specificity and interpretation tables. Thus, this lot of the HLA-A*31 subtyping kit cannot distinguish the A*31:01:02-31:01:13, 31:42-31:43, 31:45, 31:47, 31:49-31:54, 31:57-31:58 and 31:61 alleles and the A*31:14N allele.

Change in revision R02 compared to R01:

- Primer mixes 4 and 7 may weakly amplify the A*34:01 allele. This foot note has been reinserted in the specificity table.

Change in revision R03 compared to R02:

- Primer mix 8 gives rise to a significantly lower yield of the A*31:09 specific PCR product than the other HLA-A*31 primer mixes.  This has been changed in the foot notes to the Specificity Table.

Change in revision R04 compared to R03:

- Primer mix 20 does not amplify the A*31:20 allele. This has been corrected in the Specificity and Interpretation Tables. Thus, this lot of the HLA-A*31 subtyping kit cannot distinguish the A*31:01:02-31:01:17 and the A*31:20 alleles.