PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® HLA-A*33

Prod. No:	101.432-12/12u

Lot No:    	32G

Expiry Date:    	2011-October-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 16 primer mixes used for HLA-A*33 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified HLA-A*33 alleles | Other amplified HLA-A alleles3 |
| --- | --- | --- | --- | --- |
| 1 | 205 bp | 800 bp | *330101, 330301-3307, 3310-3320, 3322-3325 | *6829 |
| 2 | 205 bp | 800 bp | *330101, 3304, 3305, 3307, 3316 | *6604 |
| 3 | 155 bp | 1070 bp | *330101, 3305, 3307, 3316 | *6604 |
| 4 | 210 bp | 1070 bp | *330301-330303, 3306, 3308-3315, 3317, 3320-3325 | *0241, 0265, 0280, 9217, 9235, 9252, 2462, 250101-2508, 260101-2638, 310102-3102, 3105, 3107-3125, 320101-3203, 3205-3219N, 340101, 340102, 3405, 3406, 4301, 6601-6603, 6605-6610, 7401-7414N |
| 54 | 90 bp | 800 bp | *3304 |  |
| 64,7 | 105 bp, 175 bp | 1070 bp | *3305, 3320 |  |
| 74,6,8 | 75 bp, 105 bp, 500 bp | 800 bp | *3306, 3318 | *29010101, 29010102N, 290201-2919, 3106 |
| 84,9 | 125 bp, 235 bp | 1070 bp | *3307, 3324 | *2919 |
| 910 | 150 bp, 185 bp | 800 bp | *3308, 3309, 3314, 3322 | *0120, 0224, 9237, 2622, 6609 |
| 1011 | 140 bp, 215 bp, 285 bp | 800 bp | *3308, 3313, 3322 | *0120, 0224, 9237, 1143, 2482, 2622, 3102, 3103, 3107, 3108, 6609 |
| 115 | 165 bp | 1070 bp | *3310, 3325 | *230301, 2903, 3105, 3213 |
| 12 | 235 bp | 1070 bp | *3311 | *1143, 6829 |
| 134,5,12 | 95 bp, 165 bp | 1070 bp | *3312, 3325 |  |
| 144,13 | 115 bp, 335 bp | 1070 bp | *3315, 3319 | *0210, 021701, 021702, 0239, 9208, 9210, 9240, 9248, 0315, 0319, 2404, 2419, 2428, 2444, 2489, 2907 |
| 15 | 140 bp | 1070 bp | *3316, 3323 |  |
| 1614 | 210 bp, 245 bp | 1070 bp | *3317, 3321 |  |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-A*33 SSP typings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective length of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may  be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most wells, or a band of 800 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-A*33 subtyping.

In addition, wells number 2, 5, 7, 9 and 10 contain the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3Due to the sharing of sequence motifs between HLA-A alleles non-HLA-A*33 alleles will be amplified by primer mixes 1 to 4, 7 to 12 and 14. The A*3308, A*2622 and A*6609 give rise to identical amplification patterns with the HLA-A*33  subtyping kit. These three alleles can be distinguished by e.g. the HLA-A low resolution kit and/or the HLA-A*26 and HLA-A*66 subtyping kits.

4Short specific PCR fragments are less intense and not as sharp as longer specific bands.

5Primer mixes 11 and 13 may give to non-specific amplifications.

6Primer mix 7 may have tendencies of giving rise to primer dimers.

7Primer mix 6: Specific PCR fragment of 105  bp in the A*3320 allele. Specific PCR fragment of 175 bp in the A*3305 allele.

8Primer mix 7: Specific PCR fragment of 75 bp in the A*3318 and A*29010101, 29010102N, 290201-2919 and 3106 alleles. Specific PCR fragments of 105 and 500 bp in the A*3306 allele.

9Primer mix 8: Specific PCR fragment of 125 bp in the A*3307 allele. Specific PCR fragment of 235 bp in the A*3324 and A* 2919 alleles.

10Primer mix 9: Specific PCR fragment of 150 bp in the A*3314, 3322 and A*0120, 0224, 9237, 2622 and 6609 alleles. Specific PCR fragment of 185 bp in the A*3308 and 3309 alleles.

11Primer mix 10: Specific PCR fragment of 140 bp in the A* 3322 and A*0120, 0224, 9237, 2622 and 6609 alleles. Specific PCR fragment of 215 bp in the A*3308 and A*2482, 3102, 3107 and 3108 alleles. Specific PCR fragment of 285 bp in the A*3313 and the A*1143 and 3103 alleles.

12Primer mix 13: Specific PCR fragment of 95 bp in the A*3312 allele. Specific PCR fragment of 165 bp in the A*3325 allele.

13Primer mix 14: Specific PCR fragment of 115 bp in the A*3315 and A*9240 alleles. Specific PCR fragment of 335 bp in the A*3319 and A*0210, 021701, 021702, 0239, 9208, 9210, 9248, 0315, 0319, 2404, 2419, 2428, 2444, 2489 and 2907 alleles.

14Primer mix 16: Specific PCR fragment of 210 bp in the A*3321 allele. Specific PCR fragment of 245 bp in the A*3317 allele.

1The sequence of the A*3302 allele has been shown to be identical to A*3303.

2Primer mix 6: Specific PCR fragment of 105  bp in the A*3320 allele. Specific PCR fragment of 175 bp in the A*3305 allele.

Primer mix 7: Specific PCR fragment of 75 bp in the A*3318 and A*29010101, 29010102N, 290201-2919 and 3106 alleles. Specific PCR fragment of 105 and 500 bp in the A*3306 allele.

Primer mix 8: Specific PCR fragment of 125 bp in the A*3307 allele. Specific PCR fragment of 235 bp in the A*3324 and A* 2919 alleles.

Primer mix 9: Specific PCR fragment of 150 bp in the A*3314, 3322 and A*0120, 0224, 9237, 2622 and 6609 alleles. Specific PCR fragment of 185 bp in the A*3308 and 3309 alleles.

Primer mix 10: Specific PCR fragment of 140 bp in the A* 3322 and A*0120, 0224, 9237, 2622 and 6609 alleles. Specific PCR fragment of 215 bp in the A*3308 and A*2482, 3102, 3107 and 3108 alleles. Specific PCR fragment of 285 bp in the A*3313 and the A*1143 and 3103 alleles.

Primer mix 13: Specific PCR fragment of 95 bp in the A*3312 allele. Specific PCR fragment of 165 bp in the A*3325 allele.

Primer mix 14: Specific PCR fragment of 115 bp in the A*3315 and A*9240 alleles. Specific PCR fragment of 335 bp in the A*3319 and A*0210, 021701, 021702, 0239, 9208, 9210, 9248, 0315, 0319, 2404, 2419, 2428, 2444, 2489 and 2907 alleles.

Primer mix 16: Specific PCR fragment of 210 bp in the A*3321 allele. Specific PCR fragment of 245 bp in the A*3317 allele.

3The A*3308, A*2622 and A*6609 give rise to identical amplification patterns with the HLA-A*33  subtyping kit. These three alleles can be distinguished by e.g. the HLA-A low resolution kit and/or the HLA-A*26 and HLA-A*66 subtyping kits.