PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® HLA-A*33

Prod. No:	101.432-12/12u

Lot No:    	19S

Expiry Date:    	2015-November-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 24 primer mixes used for HLA-A*33 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified HLA-A*33 alleles3 | Other amplified HLA-A alleles4 |
| --- | --- | --- | --- | --- |
| 1 | 205 bp | 800 bp | *33:01:01-33:01:07, 33:03:01-33:07, 33:10-33:20, 33:22-33:37, 33:39-33:50, 33:52, 33:54-33:67 | *68:29 |
| 2 | 205 bp | 800 bp | *33:01:01-33:01:07, 33:04-33:05, 33:07, 33:16, 33:27, 33:32:01-33:32:02, 33:34, 33:40, 33:49-33:50, 33:64, 33:67 | *03:104, 66:04 |
| 3 | 155 bp | 1070 bp | *33:01:01-33:01:07, 33:05, 33:07, 33:16, 33:27, 33:32:01-33:32:02, 33:34, 33:40, 33:49-33:50, 33:64, 33:67 | *02:332, 03:104, 24:220, 66:04 |
| 4 | 210 bp | 1070 bp | *33:03:01-33:03:13, 33:06, 33:08-33:15, 33:17, 33:20-33:26, 33:28-33:31, 33:33, 33:35-33:37, 33:39-33:48, 33:51-33:63, 33:65-33:66 | *02:41, 02:65, 02:80, 02:117, 02:135, 02:152, 02:289, 02:304, 02:309, 03:103, 11:116, 23:45, 24:62, 25:01:01-25:22, 26:01:01-26:39, 26:41-26:43:02, 26:45-26:84, 29:32, 31:01:02-31:02, 31:05, 31:07-31:68, 31:70-31:71, 32:01:01-32:03, 32:05-32:29, 32:31, 32:33:01-32:40, 32:42-32:47, 32:49-32:57, 34:01:01-34:01:02, 34:05-34:06, 43:01, 66:01-66:03, 66:05-66:19, 74:01-74:17 |
| 55 | 90 bp | 800 bp | *33:04, 33:33 | *26:68 |
| 65,8 | 105 bp, 170 bp | 1070 bp | *33:05, 33:20, 33:58 | *29:12 |
| 75,9 | 105 bp, 230 bp | 1070 bp | *33:06, 33:51 | *11:98, 66:15, 68:04 |
| 85,7,10 | 160 bp, 235 bp | 1070 bp | *33:07, 33:24 | *02:243, 29:19, 29:48, 31:54 |
| 98,12 | 150 bp, 185 bp | 800 bp | *33:08-33:09, 33:14, 33:22, 33:58 | *01:20, 01:66, 02:24:01-02:24:02, 02:137, 02:243, 02:309, 03:95, 26:22, 29:12, 29:22, 29:48, 30:47, 66:09 |
| 1013 | 135 bp, 215 bp, 285 bp | 800 bp | *33:08, 33:13, 33:22, 33:42, 33:53 | *02:24:02, 02:243, 11:43, 24:82, 29:22, 29:48, 31:02-31:03, 31:07-31:08 |
| 116 | 165 bp | 1070 bp | *33:10, 33:25 | *23:03:01, 24:208, 29:03, 29:33, 31:05, 32:13 |
| 125,14 | 95 bp, 230 bp | 1070 bp | *33:11, 33:36, 33:42 | *11:43, 31:66, 68:29 |
| 135,6,15 | 95 bp, 165 bp | 1070 bp | *33:12, 33:25 |  |
| 145,16 | 115 bp, 335 bp | 1070 bp | *33:15, 33:19 | *02:10, 02:17:01-02:17:02, 02:39, 02:108, 02:110, 02:140, 02:148, 02:242, 02:244, 02:268, 02:300, 02:303, 02:398, 03:15, 03:19, 11:139, 24:04, 24:19, 24:28, 24:44, 24:89, 24:109, 24:129, 29:07, 29:49, 31:29, 31:48 |
| 157 | 140 bp | 1070 bp | *33:16, 33:23 |  |
| 165,17 | 95 bp, 210 bp, 245 bp | 1070 bp | *33:17, 33:21, 33:36 |  |
| 175,18 | 150 bp | 1070 bp | *33:18:01-33:18:02, 33:26 | *03:42, 03:133, 23:53, 29:01:01:01-29:04, 29:06-29:31, 29:34, 29:36-29:50, 31:03-31:04, 31:06, 32:30, 32:32 |
| 185 | 100 bp | 1070 bp | *33:27 |  |
| 195 | 120 bp | 1070 bp | *33:28 | *03:01:18, 03:22:02, 32:10, C*02:02:15 |
| 205,18 | 90 bp, 225 bp | 1070 bp | *33:29, 33:39 | *32:53 |
| 215 | 115 bp | 1070 bp | *33:30 | *02:342 |
| 225,19 | 115 bp, 255 bp | 1070 bp | *33:31, 33:44 | *02:241, 26:24 |
| 23 | 180 bp | 1070 bp | *33:32:01, 33:52 | *02:332, 24:220, B*51:136 |
| 245,20 | 95 bp, 205 bp | 1070 bp | *33:34, 33:54 | *03:01:18, 11:01:28, 24:208, 29:09, 29:33, 31:24, 32:33:01 |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-A*33 SSP typings.

When the primers in a primer mix can give rise to HLA-specific PCR products of more than one length this is indicated if the size difference is more than 20 base pairs. Size differences of 20 base pairs or less are not given. For high resolution SSP kits the respective lengths of the HLA-specific PCR product(s) are given for the alleles amplified by these primer mixes.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most wells, or a band of 800 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-A*33 subtyping. In addition, wells number 2, 5, 9 and 10 contain the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to allow kit identification.In the presence of a specific amplification the intensity of the control band often decreases.

3For several HLA Class I alleles 1st and/or 4th exon(s) and beyond, as well as intron nucleotide sequences, are not available. In these instances it is not known whether some of the primers of the SSP sets are completely matched with the target sequences or not. We assume that unknown sequences in these regions are conserved within allelic groups.

4Due to the sharing of sequence motifs between HLA-A alleles non-HLA-A*33 alleles will be amplified by primer mixes 1 to 12, 14, 17 and 19 to 24. In addition, the C*02:02:15 allele will be amplified by primer mix 19 and the B*51:136 allele will be amplified by primer mix 23.

5HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

6Primer mixes 11 and 13 may have tendencies of unspecific amplifications.

7Primer mixes 10 and 15 have a tendency to giving rise to primer oligomer formation.

8Primer mix 9 and 10 may give rise to a long unspecific amplification product of approximately 640bp. This should be disregarded when interpreting the HLA-A*33 typings.

9Primer mix 6: Specific PCR fragment of 105 bp in the A*33:20 allele. Specific PCR fragment of 175 bp in the A*33:05 and *33:58 and the A*29:12 alleles.

10Primer mix 7: Specific PCR fragment of 105 bp in the A*33:06 allele. Specific PCR fragment of 230 bp in the A*33:51 and the A*11:98, 66:15 and 68:04 alleles.

11Primer mix 8: Specific PCR fragment of 160 bp in the A*33:07 and the A*31:54 alleles. Specific PCR fragment of 235 bp in the A*33:24 and the A*02:243, 29:19 and 29:48 alleles.

12Primer mix 9: Specific PCR fragment of 150 bp in the A*33:14, 33:22 and 33:58 and the A*01:20, 01:66, 02:24:01-02:24:02, 02:137, 02:309, 03:95, 26:22, 29:12, 29:22, 30:47 and 66:09 alleles. Specific PCR fragment of 185 bp in the A*33:08 and 33:09 and the A*02:243 and 29:48 alleles.

13Primer mix 10: Specific PCR fragment of 135 bp in the A*33:22 and A*02:24:02 and 29:22 alleles. Specific PCR fragment of 215 bp in the A*33:08, 33:42 and 33:53 and the A*02:243 and 24:82, 29:48, 31:02 and 31:07-31:08 alleles. Specific PCR fragment of 285 bp in the A*33:13 and the A*11:43 and 31:03 alleles.

14Primer mix 12: Specific PCR fragment of 95 bp in the A*33:36 allele. Specific PCR fragment of 230 bp in the A*33:11 and 33:42 and the A*11:43, 31:66 and 68:29 alleles.

15Primer mix 13: Specific PCR fragment of 95 bp in the A*33:12 allele. Specific PCR fragment of 165 bp in the A*33:25 allele.

16Primer mix 14: Specific PCR fragment of 115 bp in the A*33:15 and the A*02:140 and 31:48 alleles. Specific PCR fragment of 335 bp in the A*33:19 and the A*02:10, 02:17:01-02:17:02, 02:39, 02:108, 02:110, 02:148, 02:242, 02:244, 02:268, 02:300, 02:303, 02:398, 03:15, 03:19, 11:139, 24:04, 24:19, 24:28, 24:44, 24:89, 24:109, 24:129, 29:07, 29:49 and 31:29 alleles.

17Primer mix 16: Specific PCR fragment of 95 bp in the A*33:36 allele. Specific PCR fragment of 210 bp in the A*33:21 allele. Specific PCR fragment of 245 bp in the A*33:17 allele.

18Primer mix 20: Specific PCR fragment of 90 bp in the A *33:29 allele. Specific PCR fragment of 225 bp in the A*33:39 and the A*32:53 allele.

19Primer mix 22: Specific PCR fragment of 115 bp in the A*33:31 and the A*02:241 and 26:24 alleles. Specific PCR fragment of 255 bp in the A*33:44 alleles.

20Primer mix 24: Specific PCR fragment of 95 bp in the A*33:34 and the A*03:01:18, 11:01:28, 24:208, 29:09, 29:33, 31:24 and 32:33:01 alleles. Specific PCR fragment of 205 bp in the A*33:54 allele.

1The sequence of the A*3302 allele has been shown to be identical to A*33:03:01.

The Sequence of the A*33:38 allele has been shown to be identical to A*33:44.

2HLA-A*33 alleles in bold lettering are listed as confirmed alleles on the on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.11.0, January 2013.

3The A*33:09 and the A*02:309, 26:22 and 66:09 give rise to identical amplification patterns with the HLA-A*33 subtyping kit. These alleles can be distinguished by e.g. the HLA-A low resolution kit and/or the HLA-A*02, HLA-A*26 and HLA-A*66 subtyping kits.

4The A*33:29 and 33:39 alleles may be distinguished by the different sizes of the specific PCR products generated by primer mix 20.

5The A*33:31 and 33:44 alleles may be distinguished by the different sizes of the specific PCR products generated by primer mix 22.

6The A*33:51 and A*66:15 alleles give rise to identical amplification patterns with the HLA-A*33 subtyping kit. These alleles can be distinguished by e.g. the HLA-A low resolution kit and/or the HLA-A*66 subtyping kit.

7The A*33:53 and A*31:02, 31:07 and 31:08  alleles give rise to identical amplification patterns with the HLA-A*33 subtyping kit. These alleles can be distinguished by e.g. the HLA-A low resolution kit and/or the HLA-A*31 subtyping kit.

8Primer mix 6: Specific PCR fragment of 105 bp in the A*33:20 allele. Specific PCR fragment of 175 bp in the A*33:05 and *33:58 and the A*29:12 alleles.

Primer mix 7: Specific PCR fragment of 105 bp in the A*33:06 allele. Specific PCR fragment of 230 bp in the A*33:51 and the A*11:98, 66:15 and 68:04 alleles.

Primer mix 8: Specific PCR fragment of 160 bp in the A*33:07 and the A*31:54 alleles. Specific PCR fragment of 235 bp in the A*33:24 and the A*02:243, 29:19 and 29:48 alleles.

Primer mix 9: Specific PCR fragment of 150 bp in the A*33:14, 33:22 and 33:58 and the A*01:20, 01:66, 02:24:01-02:24:02, 02:137, 02:309, 03:95, 26:22, 29:12, 29:22, 30:47 and 66:09 alleles. Specific PCR fragment of 185 bp in the A*33:08 and 33:09 and the A*02:243 and 29:48 alleles.

Primer mix 10: Specific PCR fragment of 135 bp in the A*33:22 and A*02:24:02 and 29:22 alleles. Specific PCR fragment of 215 bp in the A*33:08, 33:42 and 33:53 and the A*02:243 and 24:82, 29:48, 31:02 and 31:07-31:08 alleles. Specific PCR fragment of 285 bp in the A*33:13 and the A*11:43 and 31:03 alleles.

Primer mix 12: Specific PCR fragment of 95 bp in the A*33:36 allele. Specific PCR fragment of 230 bp in the A*33:11 and 33:42 and the A*11:43, 31:66 and 68:29 alleles.

Primer mix 13: Specific PCR fragment of 95 bp in the A*33:12 allele. Specific PCR fragment of 165 bp in the A*33:25 allele.

Primer mix 14: Specific PCR fragment of 115 bp in the A*33:15 and the A*02:140 and 31:48 alleles. Specific PCR fragment of 335 bp in the A*33:19 and the A*02:10, 02:17:01-02:17:02, 02:39, 02:108, 02:110, 02:148, 02:242, 02:244, 02:268, 02:300, 02:303, 02:398, 03:15, 03:19, 11:139, 24:04, 24:19, 24:28, 24:44, 24:89, 24:109, 24:129, 29:07, 29:49 and 31:29 alleles.

Primer mix 16: Specific PCR fragment of 95 bp in the A*33:36 allele. Specific PCR fragment of 210 bp in the A*33:21 allele. Specific PCR fragment of 245 bp in the A*33:17 allele.

Primer mix 20: Specific PCR fragment of 90 bp in the A *33:29 allele. Specific PCR fragment of 225 bp in the A*33:39 and the A*32:53 allele.

Primer mix 22: Specific PCR fragment of 115 bp in the A*33:31 and the A*02:241 and 26:24 alleles. Specific PCR fragment of 255 bp in the A*33:44 alleles.

Primer mix 24: Specific PCR fragment of 95 bp in the A*33:34 and the A*03:01:18, 11:01:28, 24:208, 29:09, 29:33, 31:24 and 32:33:01 alleles. Specific PCR fragment of 205 bp in the A*33:54 allele.