PHOTO DOCUMENT

Olerup SSP AB

Olerup SSPTM HLA-A*66

Prod. No:	101.427-06 – including Taq polymerase

101.427-06u – without Taq polymerase

Lot No:    	Y03

Expiry Date:    	2009-July-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ug/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 8 primer mixes used for HLA-A*66 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified HLA-A*66 alleles | Other amplified HLA Class I alleles3 |
| --- | --- | --- | --- | --- |
| 1 | 175 bp | 800 bp | 6601, 6604, 6605, 6606 | 250101-2505, 260101-2603, 2605-2608, 2610-2633, 2635, 4301 |
| 24 | 100 bp | 1070 bp | 6601, 6604, 6606 | 0113, 0117, 110101-1111, 1113-1116, 1120-1132, 2502, 2613, 2619, 2633, 3401weakly, 3402-3404, 3405weakly, 3408 |
| 3 | 425 bp | 1070 bp | 6601, 6602, 6604, 6606 | 0234-023503, 0256, 0262, 0278, 9203, 2419, 3401-3408, 680101-68020103, 6806-6814, 6816-6819, 6821-6830, 6832-6835, 6837-6840, 6901 |
| 4 | 165 bp | 1070 bp | 6602, 6603, 6606 | 3103, 3104, 3402-3404, 3406-3408 |
| 54,5 | 70, 105 bp | 800 bp | 6604 | 0255, 2603, 2606, 2621, 2630, 6805, 6815, 6820 |
| 64 | 80 bp | 800 bp | 6601, 6604, 6605, 6606 | 0113, 110101-1111, 1113-1116, 1119-1132, 2419, 2444, 2603, 2606, 2621, 3103, 3401-3408 |
| 7 | 560 bp | 1070 bp | 6603 | 0216 |
| 84 | 100 bp | 1070 bp | 6605 | 01010101-0104N, 0106, 0108-0112, 0114-0116N, 0118N-0127N, 1117, 1119, 2444, 250101-250102, 2503-2506, 260101-2612, 2614-2618, 2620-2629, 2631, 2632, 2634, 2635, 3103, 3313, 3601-3604, 4301, 7410, 8001 |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-A*66 SSP subtypings.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control band may sometimes be observed. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most tubes, or a band of 800 base pairs, for some tubes.

Tube number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-A*66 subtyping.

In addition, tubes number 5 and 6 contain the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3Due to the sharing of sequence motifs between HLA-A alleles non-HLA-A*66 alleles will be amplified by primer mixes 1 to 8.

4Specific PCR fragments shorter than 125 base pairs have a lower intensity than longer PCR bands.

5Primer mix 5: Specific PCR fragment of 70 bp in the A*0255, 2603, 2606, 2621, 2630, 6805, 6815 and 6820 alleles. Specific PCR fragment of 105 bp in the A*6604 allele.

1Primer mix 5: Specific PCR fragment of 70 bp in the A*0255, 2603, 2606, 2621, 2630, 6805, 6815 and 6820 alleles. Specific PCR fragment of 105 bp in the A*6604 allele.

‘w’, might be weakly amplified.