Gel Documentation Form and Worksheet

HLA-A*68 (101.418-12/04 -12u/04u) 	Lot No: 97Y	Expiry Date: 2018-May-01

Name:__________________

Sample ID:______________

DNA Conc.(ng/ul):_________

Test  Date:_______________

Tested By:_______________

Review Date:_______________

Reviewed By:_______________

Interpretation:___________	       Failed lanes: ____________	   Comments:_______________

Gel Picture

PHOTO DOCUMENT

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 1, 17, 29 and 31 may give rise to a lower yield of HLA-specific PCR product than the other HLA-A*68 primer mixes.

Primer mix 35 and 38 may have tendencies of unspecific amplifications.

Primer mix 32 has a tendency to giving rise to primer oligomer formation.

Primer mix 11 may faintly amplify the A*02:06:05, 02:06.09, 02:122, 11:70, 11:121 and 26:07:01-26:07:02 alleles.

Primer mix 48 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 430 base pairs.

1HLA-A*68 alleles in bold lettering are listed as confirmed alleles on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.21.1, August 2015.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

Primer mix 4: Specific PCR fragment of 180 bp in the A*68:30 allele. Specific PCR fragment of 225 bp in the A*68:03:01-68:03:02, A*68:05 and A*68:20 alleles. Specific PCR fragment of 180 and 225 bp in the A*02:06:05 allele.

Primer mix 5: Specific PCR fragment of 140 bp in the A*68:25 and the A*02:339, 02:407 and 02:449 alleles. Specific PCR fragment of 235 bp in the A*68:04 and the A*11:98, 33:51 and 66:15 alleles.

Primer mix 8: Specific PCR fragment of 160 bp in the A*68:07 allele. Specific PCR fragment of 240 bp in the A*68:46 allele.

Primer mix 9: Specific PCR fragment of 195 bp in the A*68:47 allele. Specific PCR fragment of 250 bp in the A*68:08:01-68:08:02 and 68:63 and the A*25:19:01-25:19:02, 25:30, 26:43:01-26:43:02, 26:112, 34:03?-34:04?, 34:06?-34:10N?, 66:06 alleles.

Primer mix 11: Specific PCR fragment of 160 bp in the A*68:17 allele. Specific PCR fragment of 200 bp in the A*68:10, A*68:14 and 68:120N and the A*02:358 alleles.

Primer mix 12: Specific PCR fragment of 200 bp in the A*68:38, 68:110 and 68:120N and in the C*06:39 alleles. Specific PCR fragment of 405 bp in the A*68:11N allele.

Primer mix 13: Specific PCR fragment of 210 bp in the A*68:12, 68:50, 68:117, 68:129 and 68:131 alleles. Specific PCR fragment of 240 bp in the A*68:51 and the A*26:18 alleles. Specific PCR fragment of 260 bp in the A*68:113 and 68:115 alleles.

Primer mix 15: Specific PCR fragment of 145 bp in the A*68:19 and the A*02:237, 03:17:01, 03:171, 24:18, 24:204 and 24:213 alleles. Specific PCR fragment of 200 bp in the A*68:14 and 68:80-68:81 and the A*02:121, 02:425, 02:517, 23:47, 24:141 and 24:234 and in the C*12:37w alleles.

Primer mix 18: Specific PCR fragment of 200 bp in the A*68:26, 68:65, 68:115 and 68:131 and in the 01:01:01:01-01:01:65, 01:01:67-01:04N, 01:06-01:29, 01:31N-01:33, 01:35-01:71, 01:73-01:78, 01:80-01:85, 01:87-01:101, 01:103-01:144, 01:146-01:166, 01:168-01:175, 01:178N-01:189, 03:18, 03:97, 03:122, 03:135, 03:167, 11:27, 11:38-11:39, 11:94, 11:209, 23:53, 24:17, 24:41, 24:208, 24:296, 29:03, 29:33 and 80:01:01:01-80:03 alleles. Specific PCR fragment of 225 bp in the A*68:21:01-68:21:02, 03:04:01-03:04:03, 11:153:01-11:153:02. Specific PCR fragment of 200 and 225 bp in the A*01:86 allele.

Primer mix 21: Specific PCR fragment of 170 bp in the A*68:28 and the A*02:49, 24:87, 24:285, 24:289, 24:299, 31:97, 66:19 and in the B*07:197, B*27:45 and 27:108 alleles. Specific PCR fragment of 215 bp in the A*68:24 and 68:123 alleles.

Primer mix 26: Specific PCR fragment of 100 bp in the A*68:49N allele. Specific PCR fragment of 190 bp in the A*68:33 and 68:80 and the A*24:141 and 33:03:17 alleles.

Primer mix 27: Specific PCR fragment of 105 bp in the A*68:18N allele.  Specific PCR fragment of 185 bp in the A*68:34 and the A*02:385 and 02:529 alleles. Specific PCR fragment of 235 bp in the A*68:43:01-68:43:02 and 11:219 alleles.

Primer mix 28: Specific PCR fragment of 200 bp in the A*68:35, 68:38 and 68:110 and in the C*06:139 alleles. Specific PCR fragment of 230 bp in the A*68:44 allele.

Primer mix 29: Specific PCR fragment of 80 bp in the A*68:42, 68:54, 68:61 and 68:63 and in the A*02:03:01-02:03:08, 02:26, 02:99, 02:117, 02:148, 02:171:02, 02:253, 02:258, 02:264, 02:281, 02:315, 02:323, 02:345, 02:355, 02:370, 02:393, 02:402, 02:412, 02:427, 02:431, 02:447, 02:463, 02:466, 02:480, 02:489, 02:505, 02:529, 02:541, 02:544, 02:557, 02:568, 03:01:38, 03:123:02, 11:25, 11:191 and 26:01:17 alleles. Specific PCR fragment of 250 bp in the A*68:36 and 68:37 and in the B*44:03:19-44:03:20 alleles.

Primer mix 30: Specific PCR fragment of 145 bp in the A*02:117, 02:135, 25:01:01-25:01:08, 25:02-25:07, 25:10-25:19:01, 25:20-25:27:02, 25:29-25:30, 26:01:01:01, 26:01:02-26:01:14, 26:01:16-26:01:25, 26:01:27-26:01:37, 26:02:01, 26:03:01-26:10, 26:12-26:13, 26:15-26:17, 26:19-26:27, 26:29-26:39, 26:41-26:43:02, 26:45-26:46, 26:48-26:72, 26:74-26:109, 26:111, 26:113, 34:01:01-34:01:02, 34:05, 34:11-34:12, 43:01, 66:01:01-66:01:02, 66:02-66:15, 66:17 and 66:19-66:21 alleles. Specific PCR fragment of 245 bp in the A*68:37 allele.

Primer mix 34: Specific PCR fragment of 105 bp in the A*68:18N allele. Specific PCR fragment of 165 bp in the A*68:55:01-68:55:02 and 68:124 and the A*03:163 allele.

Primer mix 35: Specific PCR fragment of 140 bp in the A*68:59N allele. Specific PCR fragment of 295 bp in the A*68:93 and 68:97 and the A*32:73 and in the B*44:129 alleles.

Primer mix 37: Specific PCR fragment of 145 bp in the A*68:67 and the A*34:12 alleles. Specific PCR fragment of 170 bp in the A*68:123 allele. Specific PCR fragment of 285 bp in the A*68:26, 68:65, 68:115 and 68:131 and the A*26:29, 26:49 and 66:10 alleles.

Primer mix 39: Specific PCR fragment of 95 bp in the A*68:106 allele. Specific PCR fragment of 180 bp in the A*68:68 allele.

Primer mix 42: Specific PCR fragment of 130 bp in the A*68:60 and 68:99 and the A*11:164 allele. Specific PCR fragment of 245 bp in the A*68:101 and the A*03:180 allele.

Primer mix 45: Specific PCR fragment of 150 bp in the A*68:41 and the A*01:126 alleles. Specific PCR fragment of 220 bp in the A*68:72 alleles. Specific PCR fragment of 255 bp in the A*68:86 and the A*26:57 and in the B*39:104 alleles.

4The following HLA-A*68 alleles can be distinguished by the different sizes of the HLA-specific PCR product:

| Alleles | Primer mix |
| --- | --- |
| A*68:68, 68:106 | 39 |
| A*68:99, 68:101 | 42 |

The HLA-A*68 subtyping kit cannot distinguish the following silent mutations: the  A*68:01:01:01-68:01:01:02, 68:01:03-68:01:05, 68:01:07-68:01:12 68:01:14, 68:01:18-68:01:20, 68:01:28 and 68:01:32, the A*68:01:02:01-68:01:02:02, 68:01:06, 68:01:13, 68:01:15-68:01:16 68:01:21-68:01:27 and 68:01:30-68:01:31, the A*68:02:01:01-68:02:10,  the A*68:03:01-68:03:02, the A*68:27:01-68:27:02 alleles or the A*68:43:01-68:43:02 alleles.

‘w, may be weakly amplified.

‘?’, nucleotide sequence information not available for the primer matching sequence.