Gel Documentation Form and Worksheet

HLA-B\*07 (101.512-24/03, -24u/03u) Lot No: 9G7 Expiry Date: 2021-07-01

Name:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

Sample ID:\_\_\_\_\_\_\_\_\_\_\_\_\_\_

DNA Conc.(ng/ul):\_\_\_\_\_\_\_\_\_

Test Date:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

Tested By:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

Review Date:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

Reviewed By:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

*Interpretation:\_\_\_\_\_\_\_\_\_\_\_ Failed lanes: \_\_\_\_\_\_\_\_\_\_\_ Comments:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_*

**Gel Picture**

|  |
| --- |
| PHOTO DOCUMENT |







‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

**Notes:**

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 21, 24, 33, 54, 56, 58, 69 and 73 have a tendency giving rise to primer oligomer formation.

Primer mixes 9, 11, 39, 67 and 81 may give rise to a lower yield of HLA-specific PCR product than the other B\*07 primer mixes.

Primer mixes 14, 28, 35, 38, 40, 44, 52, 53, 56, 69 and 76 may have tendencies of unspecific amplifications.

Primer mix 25 may give rise to unspecific amplification in B\*73 alleles.

Primer mix 32 may give rise to a faint band in many B\*14, B\*37, C\*04, C\*05, C\*08 and C\*18 alleles.

Primer mix 71 may give rise to an unspecific amplification product of approx. 900 bp in some HLA-B alleles. This band should be disregarded in the interpretation of HLA-B\*07 typings.

In primer mix 39 the positive control band may be weaker than for other HLA-B\*07 primer mixes.

Primer mix 89 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 430 base pairs.

   



 







  















**1**HLA-B\*07 alleles in bold lettering are listed as confirmed alleles on the IMGT/HLA web page [www.ebi.ac.uk/imgt/hla](http://www.ebi.ac.uk/imgt/hla), release 3.27.0, January 2017.

**2**Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page <http://hla.alleles.org/alleles/deleted.html>.

**3**The following HLA-B\*07 primer mixes have two or more product sizes:

|  |  |  |  |
| --- | --- | --- | --- |
| Primer Mix | Size of spec. PCR product | Amplified HLA-B\*07 alleles | Other amplified HLA Class I alleles |
| **7** | 100 bp  215 bp | \*07:04:01-07:04:02, 07:25, 07:52, 07:146  \*07:29, 07:186 | \*08:187, 40:136, 40:231, 40:279, 41:08, 42:05:01-42:05:02  \*08:123, 08:132, **C\*07:335** |
| **8** | 160 bp  235 bp | \*07:59  \*07:05:01:01-07:06:03, 07:32, 07:34, 07:69, 07:78, 07:80, 07:90, 07:97, 07:105, 07:112, 07:123, 07:137-07:138, 07:140, 07:176, 07:182N, 07:201N, 07:206-07:207, 07:209-07:210, 07:213-07:214, 07:222, 07:249, 07:258, 07:262, 07:264, 07:269-07:270, 07:278, 07:283, 07:287, 07:293, 07:304 | \*08:79, 08:187, 27:125, 27:164, 38:63, 40:01:01-40:01:20, 40:01:22-40:01:25, 40:01:27, 40:01:29-40:01:31, 40:01:33-40:01:50, 40:01:52-40:01:53, 40:07, 40:10:01:01-40:10:02, 40:12, 40:14:01-40:16, 40:21-40:23, 40:25, 40:30, 40:34, 40:36, 40:38, 40:43, 40:45-40:49, 40:54-40:55, 40:61-40:63, 40:65-40:67, 40:69, 40:72:01-40:74, 40:76-40:77, 40:79-40:81, 40:84, 40:87:01-40:88, 40:100-40:102, 40:106, 40:108, 40:112-40:114:02, 40:116-40:118N, 40:121, 40:123, 40:125-40:126, 40:128, 40:132, 40:134-40:141,  40:146-40:147, 40:149-40:156, 40:158, 40:160:01-40:160:02, 40:166, 40:168, 40:170-40:171, 40:175, 40:179, 40:182-40:183, 40:185-40:186:02, 40:188, 40:191-40:199, 40:204, 40:207-40:208, 40:210, 40:212-40:213, 40:215-40:218, 40:221-40:223, 40:227-40:228, 40:231, 40:233-40:242, 40:245, 40:247, 40:249-40:253, 40:257-40:265N, 40:272-40:273, 40:277-40:279, 40:281-40:282, 40:285-40:286N, 40:288, 40:298:01-40:301, 40:308, 40:310, 40:312, 40:315-40:316, 40:319, 40:321, 40:323-40:329, 40:332-40:333, 40:338N-40:339, 40:344, 40:346, 40:353, 40:355, 40:357-40:358, 40:360, 40:365, 40:367-40:368, 40:370, 41:08, 42:05:01-42:05:02, 44:31, 48:01:01:01-48:01:06, 48:03:01-48:03:02, 48:05-48:07, 48:09-48:12, 48:14-48:16, 48:18-48:22, 48:28-48:45, 81:01:01-81:08 |
| **10** | 285 bp  420 bp | \*07:07, 07:12, 07:14, 07:18:01-07:18:02, 07:77, 07:137, 07:162-07:163, 07:193, 07:214, 07:248, 07:255  \*07:47, 07:81 | \*15:138, 15:230, 15:241, 15:297, 18:35, 18:110, 27:125, 38:63, 40:01:01-40:01:17, 40:01:19-40:01:50, 40:01:52-40:01:53, 40:07, 40:10:01:01-40:10:02, 40:12, 40:14:01-40:14:03, 40:21-40:23, 40:25, 40:33, 40:36, 40:38, 40:42-40:43, 40:46-40:49, 40:52, 40:54-40:55, 40:59-40:63, 40:65-40:67, 40:69, 40:72:01-40:74, 40:76-40:77, 40:79, 40:81, 40:84, 40:87:01-40:88, 40:100-40:102, 40:106, 40:108, 40:112-40:114:02, 40:116-40:118N, 40:121, 40:123-40:126, 40:128, 40:132, 40:134-40:141, 40:146-40:147, 40:149-40:156, 40:158, 40:160:01-40:160:02, 40:166, 40:168, 40:170-40:172, 40:175, 40:179, 40:182-40:183, 40:185-40:188, 40:191-40:199, 40:204, 40:207-40:208, 40:210, 40:212-40:213, 40:215-40:218, 40:221-40:223, 40:227-40:228, 40:231, 40:234-40:242, 40:245, 40:247, 40:249-40:253, 40:257-40:265N, 40:272-40:273, 40:277-40:282, 40:285-40:286N, 40:288, 40:299-40:301, 40:308, 40:310, 40:312, 40:315, 40:319, 40:321, 40:323-40:329, 40:332-40:333, 40:338N-40:339, 40:344, 40:346, 40:352-40:353, 40:355, 40:357-40:358, 40:360, 40:365, 40:367-40:368, 40:370, 41:08, 44:31, 44:150, 46:06, 48:03:01-48:03:02, 48:17, 48:23, 48:40, **C\*02:23, C\*03:260, C\*04:77, C\*07:38:01-07:38:02**  \*08:52, 38:41, 38:67, 40:109?, 40:292?, 48:18 |
| **11** | 130 bp  210 bp | \*07:39  \*07:08:01-07:08:02, 07:10, 07:32, 07:38, 07:132, 07:207, 07:228:01-07:228:02 | \*08:94, 08:190w, 18:07:01-18:07:02, 18:29, 18:79, 18:136, 27:01-27:05:08, 27:05:10-27:11, 27:13-27:15, 27:17, 27:19-27:21, 27:23-27:28, 27:30-27:76, 27:78-27:84, 27:86-27:91, 27:93-27:100, 27:102-27:105, 27:107-27:118, 27:120-27:128, 27:130-27:145, 27:147-27:152, 27:154-27:156, 27:158-27:175, 35:50, 35:84, 35:162, 35:197, 35:231, 35:267, 35:269, 37:21, 37:60, 37:67, 38:06-38:07, 39:58, 39:83, 40:08, 40:13, 40:68, 40:162, 40:220, 40:232, 48:06, 48:31, 51:57, 51:103, 51:116, 51:218, 53:21, 73:01-73:02 |
| **13** | 85 bp  290 bp | \*07:31, 07:34, 07:43  \*07:12, 07:14, 07:18:01-07:18:02, 07:137, 07:162-07:163, 07:199, 07:225, 07:227, 07:248 | \*08:187, 18:110, 40:199, 41:08, 42:05:01-42:05:02, 48:37, **C\*07:38:01-07:38:02**  \*15:297, 40:30, 40:34, 40:59, 40:137, 40:160:01-40:160:02, 44:150, 48:17, **C\*03:260** |
| **15** | 385 bp  420 bp | \*07:54, 07:123  \*07:11, 07:47, 07:57, 07:75:01:01-07:75:01:02, 07:240, 07:253 | \*45:06  \*08:17, 08:47, 39:20, 39:124, 42:15 |
| **16** | 205 bp  230 bp | \*07:132  \*07:16, 07:27, 07:37:01-07:38, 07:50, 07:75:01:01-07:75:01:02, 07:91, 07:180, 07:207, 07:219, 07:228:01-07:228:02, 07:306 | \*13:14, 13:83, 14:14, 15:68, 15:71, 15:175, 15:381, 15:446, 18:01:01:01-18:01:01:13, 18:01:03-18:01:12, 18:01:14-18:01:22, 18:01:24-18:02, 18:05:01:01-18:15, 18:17N-18:24, 18:26-18:67, 18:69, 18:71-18:93, 18:95-18:106, 18:108-18:124, 18:126-18:131:01:02, 18:133-18:141, 18:143-18:149, 27:01, 27:18, 27:23, 27:29, 27:44, 27:92, 27:119, 27:131, 35:50, 35:84, 35:162, 35:197, 35:231, 35:267, 35:269, 37:01:01:01-37:36, 37:38-37:54, 37:56-37:59, 37:61, 37:62w, 37:63-37:71, 38:01:01:01-38:02:07, 38:04, 38:06-38:12, 38:15-38:25, 38:27-38:32, 38:34N-38:40, 38:42-38:71, 39:05:01:01-39:05:03, 39:07-39:08, 39:11, 39:13:01-39:13:02, 39:20, 39:37, 39:49, 39:55-39:56, 39:113, 39:118, 40:02:01:01-40:06:14, 40:06:16, 40:08-40:09, 40:11:01-40:11:02, 40:13, 40:18-40:20:01:02, 40:24, 40:26-40:29, 40:35:01-40:35:02, 40:37, 40:39-40:40, 40:44, 40:56, 40:64:01:01-40:64:01:02, 40:68, 40:70:01-40:71, 40:74-40:75, 40:78, 40:82-40:83, 40:85-40:86, 40:89-40:91, 40:93-40:99, 40:103-40:105, 40:107, 40:109, 40:111, 40:115, 40:119-40:120, 40:122, 40:127, 40:131, 40:133Q, 40:142N-40:145, 40:148-40:149, 40:157, 40:159, 40:161-40:162, 40:165, 40:167, 40:169, 40:173-40:174, 40:176-40:177, 40:180-40:181, 40:184, 40:189-40:190, 40:200-40:203, 40:205-40:206, 40:209-40:211, 40:214, 40:220, 40:224-40:226, 40:229, 40:232, 40:243-40:244, 40:246, 40:248, 40:254-40:255, 40:266-40:271, 40:274-40:276, 40:283-40:284, 40:287, 40:289-40:297, 40:302-40:307, 40:309, 40:311, 40:313-40:314, 40:317, 40:320, 40:322, 40:330-40:331, 40:334-40:335, 40:337N, 40:340-40:341, 40:343, 40:345N, 40:347-40:351, 40:354, 40:356, 40:359, 40:361N-40:364, 40:369, 40:371, 42:11, 44:103, 44:130-44:131, 44:212, 44:224, 48:01:01:01-48:21, 48:23-48:46, 49:31, 50:43, 51:57, 51:103, 51:116, 51:218, 53:21, 57:10, 81:08 |
| **17** | 345 bp  425 bp | \*07:46  \*07:11, 07:27, 07:57, 07:75:01:01-07:75:01:02, 07:236, 07:240, 07:253 | \*08:17, 08:47, 08:126, 08:131, 08:139, 38:17, 39:20, 39:124, 42:15 |
| **19** | 205 bp  250 bp  295 bp | \*07:20, 07:140, 07:153  \*07:20, 07:35, 07:140, 07:153  \*07:35 | \*15:14, 15:91, 15:131, 15:161, 18:56, 35:45, 35:71, 35:169, 35:309, 40:132, 40:178, 44:10, 44:14, 44:17, 44:31, 44:43:01-44:43:02, 44:45, 44:51, 44:62, 44:77, 44:82, 44:107, 44:123, 44:135, 44:184, 44:213, 44:230, 45:09, 45:14, 46:17, 51:23, 51:108, 52:15, 53:22, 57:07, 57:93, 58:07, 82:01:01:01-82:03  \*44:54, 44:106, 44:135, 44:158, 44:184, 44:213 |
| **20** | 160 bp  205 bp  275 bp | \*07:59  \*07:51, 07:220, 07:226  \*07:21 | **\***08:143, **C\*07:294, C\*07:526, C\*07:530, C\*08:08:01-08:08:02**  \*53:46 |
| **21** | 100 bp  165 bp | \*07:52  \*07:22:01-07:22:02 |  |
| **23** | 85 bp  270 bp | \*07:20, 07:24, 07:131, 07:220, 07:223  \*07:272N | \*15:230, 15:297, 40:63, 40:324, **C\*03:260** |
| **25** | 305 bp  425 bp | \*07:117  \*07:27, 07:236, 07:240 | \*40:250  \*08:126, 08:131, 08:139, 38:17, 40:188? |
| **26** | 95 bp  205 bp  235 bp  290 bp | \*07:28, 07:118  \*07:51, 07:220, 07:226  \*07:20, 07:24, 07:60, 07:100, 07:131, 07:220, 07:223  \*07:42 | **\***08:14, 08:87, 27:32, 40:180, 44:54, 44:106, 44:135  \*08:143, **C\*07:294, C\*07:526, C\*07:530, C\*08:08:01-08:08:02**  **\***08:21, 08:25, 08:175, 15:07:01:01-15:07:03, 15:45, 15:55, 15:68, 15:126, 15:197:02, 15:207, 15:324, 15:329, 15:331, 15:360, 15:383, 15:405, 15:431, 15:450, 35:05:01:01-35:05:03, 35:22, 35:51, 35:58, 35:72, 35:89, 35:97, 35:114, 35:199, 35:232, 35:368, 40:05:01:01-40:05:01:02, 40:71, 40:174, 44:54, 44:106, 44:135, 44:158, 44:184, 44:213, 46:12, 51:64, 51:148, 53:14, 55:51, 58:18, **C\*03:05, C\*03:25, C\*03:27, C\*03:167, C\*03:296:01-03:296:02, C\*03:335** |
| **27** | 165 bp  215 bp | \*07:41  \*07:29, 07:186 | \*08:123, 08:132, **C\*07:335** |
| **29** | 85 bp  200 bp | \*07:31, 07:34, 07:43  \*07:44N | \*08:187, 18:110, 40:199, 41:08, 42:05:01-42:05:02, 48:37, **C\*07:38:01-07:38:02** |
| **31** | 215 bp  280 bp | \*07:55, 07:100  \*07:19, 07:33:01-07:33:03, 07:53, 07:60, 07:100 | \*08:70, 15:07:01:01-15:07:03, 15:45, 15:68, 15:126, 15:207, 15:324, 15:331, 15:405, 15:431, 15:450, 46:12, 48:19, **C\*02:60**  \*08:103, 08:134, 08:175, 13:71, 13:108, 14:05, 14:13, 14:53, 15:07:01:01-15:07:03, 15:45, 15:55, 15:68, 15:126, 15:197:02, 15:207, 15:324, 15:329, 15:331, 15:360, 15:383, 15:405, 15:421, 15:431, 15:450, 18:14, 27:07:01-27:07:06, 27:11, 27:24, 27:32-27:34, 27:43, 27:138, 35:05:01:01-35:05:03, 35:22, 35:31, 35:51, 35:58, 35:72, 35:89, 35:97, 35:114, 35:199, 35:232, 35:368, 37:09, 38:19, 39:03:01:01-39:03:01:02, 39:14:01:01-39:14:01:02, 39:24:01-39:24:02, 39:29, 39:37, 39:76, 39:120, 40:02:01:01-40:03:01:02, 40:05:01:01-40:05:01:02, 40:08-40:09, 40:13, 40:18-40:19, 40:24, 40:27:01-40:27:02, 40:29, 40:35:01-40:35:02, 40:37, 40:39-40:40, 40:50, 40:56-40:58, 40:71, 40:78, 40:82, 40:85, 40:89-40:91, 40:94, 40:97, 40:104-40:105, 40:107, 40:111, 40:115, 40:119, 40:122, 40:133Q, 40:142N-40:145, 40:157, 40:164, 40:169, 40:173-40:174, 40:176, 40:180-40:181, 40:189, 40:200-40:203, 40:205-40:206, 40:209, 40:211, 40:214, 40:219-40:220, 40:224, 40:226, 40:229, 40:232, 40:246, 40:248, 40:254-40:255, 40:266-40:267, 40:271, 40:274, 40:276, 40:283-40:284, 40:287, 40:289-40:293, 40:296-40:297, 40:302-40:305, 40:309, 40:314, 40:320, 40:322, 40:331, 40:334-40:335, 40:337N, 40:345N, 40:347, 40:356, 40:359, 40:369, 40:371, 42:08, 44:54, 44:106, 44:135, 44:158, 44:184, 44:213, 46:12, 48:04:01-48:04:02, 48:13, 48:24, 51:64, 51:81, 51:148, 53:14, 55:04, 55:49, 55:51, 55:86, 58:18, 58:27, **C\*02:60, C\*03:05, C\*03:25, C\*03:27, C\*03:167, C\*03:296:01-03:296:02, C\*03:335, C\*03:386, C\*07:526, C\*07:530, C\*08:08:01-08:08:02** |
| **33** | 90 bp  350 bp | \*07:149  \*07:61 | \*08:02, 08:117, 13:36, 15:89, 15:115, 18:09, 38:02:01-38:02:06, 38:08, 38:15, 38:18, 38:23, 38:29, 38:35, 38:43-38:50, 38:62, 38:64, 44:12, 44:232, 44:263, 51:54, 51:78:01-51:78:02, 53:09, 53:11-53:13, 53:31, 53:36, 56:07, 59:08 |
| **35** | 105 bp  215 bp | \*07:63  \*07:143 | \*27:157, 40:166, 42:02:01:01-42:02:01:02, 42:09, 42:17-42:18, 45:06, 54:02, 55:16, 56:35 |
| **36** | 175 bp  250 bp | \*07:167N, 07:201N, 07:285N  \*07:64 |  |
| **38** | 190 bp  215 bp | \*07:103, 07:113  \*07:66 | \*08:64 |
| **40** | 265 bp  395 bp | \*07:56:01-07:56:02, 07:72, 07:206  \*07:83 | \*08:26:01-08:26:03, 08:56:01-08:56:02, 08:180, 14:54, 38:28, 39:63, 39:135, 40:21?, 40:58?, 40:137?, 40:248?, 40:282?, 40:325?, 40:351?, 40:364?-40:365?, 41:12, 41:36, 42:21, 49:03, 49:48 |
| **41** | 400 bp  425 bp | \*07:58  \*07:156 | \*08:191 |
| **42** | 195 bp  225 bp | \*07:67N  \*07:71 |  |
| **43** | 185 bp  270 bp | \*07:68:01-07:68:03, 07:214, 07:237, 07:262  \*07:272N | \*15:138, 15:230, 15:241, 15:297, 18:35, 18:110, 35:66, 40:77, 40:87:01-40:87:02, 40:121, 40:158, 40:222, 40:237, 40:368, 44:150, 48:12, 48:14, 48:26, 48:29 |
| **48** | 200 bp  280 bp | \*07:76  \*07:77, 07:193 | \*40:60, 40:245, **C\*07:38:01-07:38:02** |
| **50** | 75 bp  105 bp  165 bp | \*07:74  \*07:175  \*07:87 | \*08:77, 48:27  \*15:126 |
| **51** | 335 bp  470 bp | \*07:97, 07:316N  \*07:93 | \*08:114, 39:100  \*08:39, 38:27 |
| **54** | 160 bp  185 bp | \*07:238  \*07:98, 07:167N, 07:201N |  |
| **57** | 150 bp  215 bp | \*07:109  \*07:70 |  |
| **58** | 245 bp  440 bp | \*07:112  \*07:79 |  |
| **59** | 100 bp  190 bp | \*07:115, 07:182N, 07:246  \*07:89, 07:113 | \*08:69, 14:05, 38:19, 39:03:01:01-39:03:01:02, 39:24:01-39:24:02, 39:37, 39:76, 39:120, 40:71, 40:314, 41:18, 42:06, 42:20, 48:07, **C\*08:08:01-08:08:02**  **\***08:64 |
| **61** | 205 bp  325 bp  420 bp | \*07:88  \*07:316N  \*07:105 |  |
| **63** | 65 bp  230 bp | \*07:94  \*07:135N | **\***67:04, **C\*02:120, C\*03:124, C\*04:49, C\*05:15, C\*07:262, C\*07:417, C\*07:475, C\*08:153, C\*12:95** |
| **65** | 385 bp  420 bp  450 bp | \*07:54, 07:123  \*07:315N  \*07:111N | \*45:06  \*08:72N |
| **68** | 335 bp  365 bp  415 bp | \*07:148  \*07:101  \*07:13, 07:15, 07:160 | \*40:193?  \*08:193  \*08:16, 08:123, 08:145, 42:18, 67:02 |
| **69** | 130 bp  190 bp  245 bp  465 bp | \*07:136:01-07:136:02  \*07:106  \*07:108  \*07:102 | \*40:320  \*08:18, 73:01-73:02 |
| **70** | 125 bp  205 bp | \*07:104  \*07:114 |  |
| **71** | 245 bp  280 bp  390 bp | \*07:108  \*07:107  \*07:181N |  |
| **72** | 165 bp  235 bp | \*07:13, 07:110  \*07:315N | \*37:52, 38:11, 48:36, 54:02, 67:02, 81:06, **A\*26:142, C\*03:137, C\*03:182, C\*07:102w, C\*07:351, C\*12:82, C\*16:116** |
| **73** | 85 bp  240 bp | \*07:119, 07:169  \*07:49N | \*48:34 |
| **74** | 345 bp  425 bp | \*07:120  \*07:156, 07:282 | \*40:150, 48:45  \*40:278 |
| **75** | 105 bp  250 bp | \*07:126, 07:175  \*07:121 | \*15:126 |
| **76** | 130 bp  170 bp  285 bp | \*07:136:01-07:136:02  \*07:128  \*07:122 | \*40:320 |
| **77** | 335 bp  405 bp | \*07:148  \*07:116 | \*40:193? |
| **78** | 110 bp  160 bp  380 bp  415 bp | \*07:125  \*07:231N  \*07:129  \*07:282 | \*40:278 |
| **79** | 65 bp  95 bp | \*07:130  \*07:119 |  |
| **85** | 110 bp  160 bp | \*07:208  \*07:238, 07:285N |  |
| **87** | 210 bp  260 bp | \*07:251N  \*07:218, 07:220 | \*08:118, 14:05, 14:13, 14:53, 18:14, 40:174, 51:64 |

**4**The B\*07:50 and B\*48:08 alleles will give rise to identical amplification patterns. These alleles can e.g. be distinguished by the HLA-B low resolution kit and/or the HLA-B\*48 high resolution kit.

**5**The following HLA-B\*07 alleles can be distinguished by the different sizes of the HLA-specific PCR product:

|  |  |  |  |
| --- | --- | --- | --- |
| **Alleles** | **Primer mix** | **Alleles** | **Primer mix** |
| B\*07:42, 07:118 | 26 | B\*07:89, 07:246 | 59 |
| B\*07:49N, 07:169 | 73 | B\*07:94, 07:135 | 63 |
| B\*07:60, 07:100 | 31 | B\*07:102, 07:106 | 69 |
| B\*07:61, 07:149 | 33 | B\*07:121, 07:126 | 75 |
| B\*07:63, 07:143 | 35 | B\*07:129, 07:231N | 78 |

‘w’, might be weakly amplified.

‘?’, nucleotide sequence information not available for the primer matching sequence.