Gel Documentation Form and Worksheet

HLA-B*15 (101.516-24/04, -24u/04u) 	Lot No: 66V	Expiry Date: 2016-November-01

Name:__________________

Sample ID:______________

DNA Conc.(ng/ul):_________

Test  Date:_______________

Tested By:_______________

Review Date:_______________

Reviewed By:_______________

Gel Picture

PHOTO DOCUMENT

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

Interpretation:_______________

Failed lanes: _______________

Comments: ___________________

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mix 23, 38 and 66 may give rise to a lower yield of HLA-specific PCR product than the other HLA-B*15 primer mixes.

Primer mixes 12, 15, 32, 40, 59, 80, 82, 84 and 85 have a tendency to giving rise to primer oligomer formation.

Primer mixes 5, 11, 19, 30, 34, 81 and 83 may have tendencies of unspecific amplifications.

Primer mix 77 may give rise to a lower yield of HLA-specific PCR product than the other HLA-B*15 resolution primer mixes.

Primer mix 91 may give rise to a long unspecific amplification product of approximately 800 bp. This should be disregarded when interpreting the B*15 typings.

Primer mix 96 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by control primer pairs. PCR product sizes range from 75 to 200 base pairs. The PCR product generated by the control primer pair is 430 base pairs.

1HLA-B*15 alleles in bold lettering are listed as confirmed alleles on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.15.0, January 2012.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3Primer mix 4: Specific PCR fragment of 180 bp in the B*15:01:01:02N allele. Specific PCR fragment of 390 bp in the B*15:285 allele.

Primer mix 10: Specific PCR fragment of 300 bp in the B*15:37-15:38:02 and 15:185 alleles. Specific PCR fragment of 520 bp in the B*15:06, 15:27:01-15:27:03, 15:84, 15:109, 15:112 and 15:195 alleles.

Primer mix 13: Specific PCR fragment of 395 bp in the B*15:09-15:10:02, 15:18:01-15:18:06, 15:21, 15:23, 15:37, 15:44, 15:51-15:52, 15:66, 15:72, 15:80, 15:90, 15:93, 15:99, 15:108, 15:114-15:115, 15:119, 15:124, 15:133-15:134, 15:153, 15:161, 15:176, 15:186, 15:189, 15:197-15:198, 15:200, 15:221, 15:226N, 15:229, 15:238, 15:243, 15:252, 15:263, 15:275, 15:290 and 15:292-15:294N allele. Specific PCR fragment of 435 bp in the B*15:79N allele.

Primer mix 14: Specific PCR fragment of 465 bp in the B*15:09-15:10:02, 15:30, 15:37, 15:45, 15:48, 15:63, 15:83, 15:90, 15:99, 15:150, 15:188, 15:243, 15:248, 15:252 and 15:287 alleles. Specific PCR fragment of 515 bp in the B*15:110 allele.

Primer mix 16: Specific PCR fragment of 120 bp in the B*15:116 and 15:124 and the B*40:52, 40:59, 46:06, 46:43, 48:23 and 57:59 alleles. Specific PCR fragment of 190 bp in the B*15:12, 15:19, 15:270, 15:298 and 15:304N alleles.

Primer mix 20: Specific PCR fragment of 360 bp in B*15:105 allele. Specific PCR fragment of 505 bp in the B*15:16:01-15:16:03, 15:34, 15:62, 15:67, 15:85-15:86, 15:95, 15:137, 15:194, 15:222, 15:254 and 15:303 and the B*46:11, 46:18, 46:33 and 46:53 alleles.

Primer mix 21: Specific PCR fragment of 205 bp in the B*15:16:01-15:16:03, 15:67, 15:95, 15:222 and 15:254 and in the C*03:158 alleles. Specific PCR fragment of 425 bp in the B*15:43 allele.

Primer mix 24: Specific PCR fragment of 205 bp in the B*15:26N allele. Specific PCR fragment of 305 bp in the B*15:33, 15:72, 15:78:01-15:78:03, 15:107, 15:116, 15:141, 15:150, 15:188, 15:209N, 15:248, 15:287 and the B*46:06 and 46:45 alleles.

Primer mix 25: Specific PCR fragment of 315 bp in the B*15:46, 15:53, 15:106 and 15:212 alleles. Specific PCR fragment of 385 bp in the B*15:28 allele.

Primer mix 30: Specific PCR fragment of 320 bp in the B*15:40, 15:47:01-15:47:02, 15:49, 15:52, 15:114, 15:117, 15:124, 15:138, 15:238 and 15:241 and the B*46:06, 46:25 and 46:30 alleles. Specific PCR fragment of 355 bp in the B*15:75 allele.

Primer mix 31: Specific PCR fragment of 320 bp in the B*15:44, 15:50, 15:69, 15:83, 15:86, 15:93, 15:121, 15:186, 15:188, 15:199, 15:218Q, 15:224, 15:252, 15:265 and the B*46:03, 46:18, 46:21:01-46:21:02 and 46:29 alleles. Specific PCR fragment of 395 bp in the B*15:114 and 15:153 alleles.

Primer mix 32: Specific PCR fragment of 275 bp in the B*15:196 allele. Specific PCR fragment of 350 bp in the B*15:190N allele. Specific PCR fragment of 380 bp in the B*15:48, 15:108, 15:136, 15:235, 15:249 and the B*46:19 alleles.

Primer mix 33: Specific PCR fragment of 200 bp in the B*15:118 allele. Specific PCR fragment of 255 bp in the B*15:18:03, 15:42, 15:73, 15:86, 15:224 and 15:303 and the B*46:11 and 46:18 alleles. Specific PCR fragment of 340 bp in the B*15:192 and 15:281 alleles.

Primer mix 36: Specific PCR fragment of 360 bp in the B*15:82 and 15:260 and in the C*03:186:01 alleles. Specific PCR fragment of 435 bp in the B*15:36, 15:89, 15:115 and 15:256 alleles.

Primer mix 39: Specific PCR fragment of 315 bp in the B*15:60 and 15:119 and in the C*07:213 alleles. Specific PCR fragment of 385 bp in the B*15:95 and 15:285 alleles.

Primer mix 40: Specific PCR fragment of 220 bp in the B*15:71 and B*15:175 and in the C*06:77 alleles. Specific PCR fragment of 425 bp in the B*15:13:01-15:13:02, 15:16:01-15:17:02, 15:23-15:24:02, 15:67, 15:95, 15:157, 15:162, 15:168, 15:177, 15:196, 15:208, 15:216, 15:222, 15:230, 15:254, 15:268 and 15:273 alleles.

Primer mix 42: Specific PCR fragment of 345 bp in the B*15:02:01-15:03:04, 15:05:01-15:06, 15:09-15:10:02, 15:13:01-15:13:02, 15:16:01-15:18:06, 15:21, 15:23, 15:25:01-15:25:03, 15:29, 15:31, 15:36-15:37, 15:39:01-15:40, 15:44, 15:48, 15:52, 15:55, 15:61-15:62, 15:64:01-15:64:02, 15:67, 15:69, 15:72, 15:74, 15:80, 15:88w, 15:89-15:91, 15:93, 15:95, 15:98, 15:103, 15:106-15:108, 15:112, 15:114-15:115, 15:119, 15:121, 15:123-15:124, 15:127, 15:131-15:134, 15:136, 15:138-15:139, 15:151, 15:153, 15:155-15:156, 15:158, 15:161-15:162, 15:170, 15:173, 15:176w, 15:177, 15:185-15:186, 15:188, 15:194-15:198, 15:200w, 15:204, 15:208, 15:210, 15:213-15:214, 15:216, 15:219-15:220, 15:222-15:223, 15:226N, 15:229-15:230, 15:235, 15:238, 15:240, 15:242-15:243, 15:250w, 15:252-15:255, 15:263, 15:265-15:266, 15:268, 15:271, 15:273-15:275, 15:281-15:283, 15:288-15:294N, 15:297, 15:301-15:302N and the B*46:06, 46:08, 46:13:01-46:13:03, 46:19, 46:21:01-46:21:02, 46:25-46:26, 46:33 and 46:43 alleles. Specific PCR fragment of 460 bp in the B*15:73 and 15:303 alleles. Specific PCR fragments of 345 and 460 bp in the B*15:86 and 15:224 and the B*46:11 and 46:18 alleles.

Primer mix 50: Specific PCR fragment of 85 bp in the B*15:67 and the B*35:110 and in the A*68:56 alleles. Specific PCR fragment of 185 bp in the B*15:206 allele.

Primer mix 51: Specific PCR fragment of 135 bp in the B*15:98, 15:109 and 15:163 alleles. Specific PCR fragment of 175 bp in the B*15:277 and 15:278 and the B*35:122 and 56:29 alleles. Specific PCR fragment of 200 bp in the B*15:246N allele.

Primer mix 52: Specific PCR fragment of 195 bp in the B*15:104 alleles. Specific PCR fragment of 345 bp in the B*15:92 and 15:213 alleles.

Primer mix 54: Specific PCR fragment of 250 bp in the B*15:187 and in the C*12:36 alleles. Specific PCR fragment of 410 bp in the B*15:97 allele.

Primer mix 55: Specific PCR fragment of 175 bp in the B*15:214 and the B*07:120, 18:81, 40:150, 51:165 and 57:29 alleles. Specific PCR fragment of 255 bp in the B*15:102 and 15:284 and the B*51:30 and 73:01-73:02 alleles.

Primer mix 56: Specific PCR fragment of 70 bp in the B*15:103 allele. Specific PCR fragment of 260 bp in the B*15:227 and the B*27:69, 37:23 and 40:213 alleles.

Primer mix 59: Specific PCR fragment of 230 bp in the B*15:191 and 15:234 and in the C*06:33 alleles. Specific PCR fragment of 375 bp in the B*15:198 and the B*35:161 alleles. Specific PCR fragment of 420 bp in the B*15:94N and 15:294N alleles.

Primer mix 61: Specific PCR fragment of 160 bp in the B*15:175 allele. Specific PCR fragment of 205 bp in the B*15:125 and the B*44:22, 44:105 and 49:24 alleles.

Primer mix 64: Specific PCR fragment of 90 bp in the B*15:302N allele. Specific PCR fragment of 365 bp in the B*15:122 and in the C*03:48 and C*06:85 alleles.

Primer mix 67: Specific PCR fragment of 265 bp in the B*15:53, 15:183 and 15:212 and the B*46:10 alleles. Specific PCR fragment of 300 bp in the B*15:132 allele.

Primer mix 68: Specific PCR fragment of 215 bp in the B*15:134 allele. . Specific PCR fragment of 380 bp in the B*15:247 and in the C*05:03 alleles.

Primer mix 70: Specific PCR fragment of 90 bp in the B*15:208 allele. Specific PCR fragment of 185 bp in the B*15:138, 15:230, 15:241 and 15:297 and the B*07:68:01-07:68:03, 18:35, 35:66, 40:77, 40:87:01-40:87:02, 40:121, 40:158, 40:222, 40:237, 44:150, 48:12, 48:14, 48:26 and 48:29 alleles.

Primer mix 71: Specific PCR fragment of 315 bp in the B*15:205 allele. Specific PCR fragment of 400 bp in the B*15:139 allele.

Primer mix 72: Specific PCR fragment of 110 bp in the B*15:226N and the B*35:216N and in the B*39:87N alleles. Specific PCR fragment of 295 bp in the B*15:140 and the B*51:166 alleles. Specific PCR fragment of 390 bp in the B*15:201 allele.

Primer mix 74: Specific PCR fragment of 75 bp in the B*15:145 and  15:176 and the B*35:237 and 52:28 alleles. Specific PCR fragment of 230 bp in the B*15:262N allele.

Primer mix 75: Specific PCR fragment of 280 bp in the B*15:146 allele. Specific PCR fragment of 390 bp in the B*15:201 allele.

Primer mix 77: Specific PCR fragment of 185 bp in the B*15:174 allele. Specific PCR fragment of 230 bp in the B*15:123 and 15:220 alleles.

Primer mix 78: Specific PCR fragment of 110 bp in the B*15:142 and the B*07:29, 07:186, 08:01:11, 08:56:01, 35:218, 51:68 and 57:49 and in the C*07:02:30 alleles. Specific PCR fragment of 385 bp in the B*15:147 and the B*46:42 alleles.

Primer mix 79: Specific PCR fragment of 135 bp in the B15:270 alleles. Specific PCR fragment of 215 bp in the B*15:182N allele. Specific PCR fragment of 420 bp in the B*15:148 allele.

Primer mix 80: Specific PCR fragment of 85 bp in the B*15:149N allele. Specific PCR fragment of 225 bp in the B*15:263 allele.

Primer mix 82: Specific PCR fragment of 80 bp in the B*15:81 and the B*58:47 alleles. Specific PCR fragment of 200 bp in the B*15:246N allele. Specific PCR fragment of 245 bp in the B*15:277 and the B*35:222 and 40:76 and in the C*05:30 alleles.

Primer mix 84: Specific PCR fragment of 155 bp in the B*15:167 allele. Specific PCR fragment of 195 bp in the B*15:211 and in the C*03:168 alleles. Specific PCR fragment of 235 bp in the B*15:197 and 15:215 and the B*35:117 alleles.

Primer mix 85: Specific PCR fragment of 340 bp in the B*15:157 allele. Specific PCR fragment of 395 bp in the B*15:76, 15:101 and 15:255 and in the C*07:335 alleles. Specific PCR fragment of 420 bp in the B*15:164 allele.

Primer mix 86: Specific PCR fragment of 230 bp in the B*15:166 allele. Specific PCR fragment of 260 bp in the B*15:272N allele. Specific PCR fragment of 400 bp in the B*15:193 allele.

Primer mix 87: Specific PCR fragment of 85 bp in the B*15:71, 15:175, 15:204 and 15:225 and the B*18:29, 35:50, 35:231, 37:22, 40:248, 41:12, 46:14, 49:03, 51:151, 54:02, 55:16, 56:35, 57:39 and 58:41 alleles. Specific PCR fragment of 295 bp in the B*15:169 allele.

Primer mix 88: Specific PCR fragment of 240 bp in the B*15:171 allele. Specific PCR fragment of 445 bp in the B*15:158 allele.

Primer mix 89: Specific PCR fragment of 370 bp in the B*15:160 allele. Specific PCR fragment of 420 bp in the B*15:94N and 15:294N alleles.

Primer mix 90: Specific PCR fragment of 95 bp in the B*15:202 and 15:239 and the B*13:58, 18:12, 35:10, 35:13, 35:16, 35:28, 35:69, 35:80, 35:226, 37:01:01-37:01:05, 37:01:07-37:05, 37:07, 37:09-37:27, 37:29-37:47, 40:185, 40:209, 44:83, 44:134, 49:03, 52:01:01:01-52:01:20, 52:03-52:15, 52:17, 52:19-52:31:01, 52:33-52:36, 53:17:01-53:17:02, 53:28 and 78:05-78:06 alleles. Specific PCR fragment of 195 bp in the B*15:173, 18:63 and 50:11 alleles.

Primer mix 91: Specific PCR fragment of 135 bp in the B*15:184 and 15:203 alleles. Specific PCR fragment of 325 bp in the B*15:170 allele.

Primer mix 92: Specific PCR fragment of 95 bp in the B*15:266 and the B*13:37 and 38:49 alleles. Specific PCR fragment of 155 bp in the B*15:172 allele.

Primer mix 93: Specific PCR fragment of 255 bp in the B*15:258N and 15:304N alleles. Specific PCR fragment of 295 bp in the B*15:177 allele. Specific PCR fragment of 385 bp in the B*15:120 allele.

Primer mix 94: Specific PCR fragment of 260 bp in the B*15:272N allele. Specific PCR fragment of 285 bp in the B*15:123 allele. Specific PCR fragment of 325 bp in the B*15:178 allele.

Primer mix 95: Specific PCR fragment of 340 bp in the B*15:181N allele. Specific PCR fragment of 395 bp in the B*15:40, 15:47:01-15:47:02, 15:49, 15:52, 15:114, 15:117, 15:124, 15:138, 15:238 and 15:241 and the B*46:06, 46:25 and 46:30 alleles.

‘w’, might be weakly amplified.

‘?’, nucleotide sequence information not available for the primer matching sequence.

4The following HLA-B*15 alleles can be distinguished by the different sizes of the HLA-specific PCR product:

| Alleles | Primer mix | Alleles | Primer mix |
| --- | --- | --- | --- |
| B*15:92, 15:104 | 52 | B*15:166, 15:193 | 86 |
| B*15:120, 15:258N | 93 | B*15:169,15:225 | 87 |
| B*15:142, 15:147 | 78 | B*15:208, 15:230 | 70 |

The HLA-B*15 subtyping kit cannot distinguish these silent mutations: the B*15:01:01:01, 15:01:03, 15:01:06-15:01:13, 15:01:15-15:01:16, 15:01:18-15:01:30 and 15:01:32-15:01:34 alleles, the B*15:02:01-15:02:08 alleles, the B*15:03:01-15:03:02 and 15:03:04 alleles, the B*15:05:01-15:05:02 alleles, the B*15:07:01-15:07:03 alleles, the B*15:11:01-15:11:02 and 15:11:05-15:11:06 alleles, the B*15:13:01-15:13:02 alleles, the B*15:16:01-15:16:03 alleles, the B*15:17:01:01-15:17:02 alleles, the B*15:18:01 and 15:18:04-15:18:06 alleles, the B*15:24:01-15:24:02 alleles. the B*15:25:01-15:25:03 alleles, the B*15:27:01-15:27:03 alleles, the B*15:32:01-15:32:02, the B*15:38:01-15:38:02 alleles, the B*15:39:01-15:39:02 alleles, the B*15:47:01-15:47:02 alleles, the B*15:64:01-15:64:02 alleles or the B*15:78:01-15:78:03 alleles.

5This lot of the B*15 kit cannot distinguish the B*15:159 and the B*15:01:01:01, 15:01:03, 15:01:06-15:01:13, 15:01:15-15:01:16, 15:01:18-15:01:30, 15:01:32-15:01:34, 15:217, 15:231-15:232, 15:244, 15:251, 15:257, 15:261, 15:264, 15:267, 15:276, 15:279-15:280, 15:286 and 15:296 alleles or the B*15:260 and the B*15:82 and 15:256 alleles.

Change in revision R01 compared to R00:

- The B*15:228 allele may not be amplified in primer mixes 41 and 44 to 48. This has been corrected in the Specificity and Interpretation Tables.

Change in revision R02 compared to R01:

- Primer mix 4 contains the shorter, 800 bp, internal positive control band. This has been corrected in the Specificity Table and the Primer Specification Table.

Changes in revision R03 compared to R02:

- Primer mix 66 may give rise to a lower yield of HLA-specific PCR product than the other HLA-B*15 resolution primer mixes. A foot note has been added in the Specificity Table.

Changes in revision R04 compared to R03:

- Primer mix 83 does not amplify the B*15:159 and 15:260 and the B*07:23, 18:95, 40:79, 40:190, 41:14 and 46:52 alleles. Thus, this lot of the B*15 kit cannot distinguish the B*15:159 and the B*15:01:01:01, 15:01:03, 15:01:06-15:01:13, 15:01:15-15:01:16, 15:01:18-15:01:30, 15:01:32-15:01:34, 15:217, 15:231-15:232, 15:244, 15:251, 15:257, 15:261, 15:264, 15:267, 15:276, 15:279-15:280, 15:286 and 15:296 alleles or the B*15:260 and the B*15:82 and 15:256 alleles. This has been corrected in the Specificity and Interpretation Tables.

Changes in revision R05 compared to R04:

- Primer mix 77 may give rise to a lower yield of HLA-specific PCR product than the other HLA-B*15 resolution primer mixes. A foot note has been added in the Specificity Table.