HLA-B\*18 (101.519-12/12u) Lot No: 7R0 Expiry Date: 2027-01-01

Name:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

Sample ID:\_\_\_\_\_\_\_\_\_\_\_\_\_\_

DNA Conc.(ng/ul):\_\_\_\_\_\_\_\_\_

Test Date:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

Tested By:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

Review Date:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

Reviewed By:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

***Interpretation:\_\_\_\_\_\_\_\_\_\_ Failed lanes: \_\_\_\_\_\_\_\_\_\_\_ Comments:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_***

**Gel Picture**

|  |
| --- |
| PHOTO DOCUMENT |





Abbreviations

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

**Notes:**

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 17, 18, 20, 30 and 36 may have tendencies of unspecific amplifications.

Primer mixes 10, 18, 26 and 38 have a tendency giving rise to primer oligomer formation.

Primer mixes 29 and 37 may give rise to a lower yield of HLA-specific PCR product than the other B\*18 primer mixes.

Primer mix 42 contains a negative control, which will amplify a majority of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 200 base pairs.



















**1**HLA-B\*18 alleles in bold lettering are listed as confirmed alleles on the on the IMGT/HLA web page [www.ebi.ac.uk/imgt/hla](http://www.ebi.ac.uk/imgt/hla), release 3.29.0, August 2017.

**2**Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page <http://hla.alleles.org/alleles/deleted.html>.

**3**The following HLA-B\*18 primer mixes have two or more product sizes:

|  |  |  |  |
| --- | --- | --- | --- |
| Primer Mix | Size of spec. PCR product | Amplified HLA-B\*18 alleles | Other amplified HLA Class I alleles |
|  **5** | 115 bp 145 bp | \*18:72:01-18:72:03\*18:04:01, 18:24, 18:98:02, 18:125  | \*07:05:11, 08:04:02, 15:01:38w, 15:64:02, 35:01:57w, 35:02:02w, 35:20:02w, 35:323:01w, 39:39:02, 53:01:26w, **C\*06:02:101, C\*12:301w** |
|  **6** | 130 bp200 bp245 bp | \*18:05:01:01-18:05:01:02\*18:31\*18:60  | **\***39:193**\***40:67, 47:06**C\*02:177** |
|  **7** | 90 bp 150 bp 230 bp | \*18:41\*18:07:01-18:07:03, 18:79, 18:136\*18:74N | \*15:416\*08:176, 35:289, 35:487**C\*07:483N** |
|  **9** | 125 bp 255 bp | \*18:75\*18:08, 18:27  | **A\*24:220, C\*14:76** |
| **10**  | 85 bp175 bp | \*18:23N\*18:28 | \*44:108N\*15:250, 27:132, 40:226, 57:156 |
| **11** | 125 bp 610 bp  | \*18:09, 18:54, 18:67, 18:136, 18:194:01:01-18:195 \*18:36, 18:56-18:58, 18:101 | \*44:03:12, 53:01:26, **A\*24:02:37**\*15:01:38, 15:64:02, 35:02:02, 35:20:02, 35:323:01, **A\*26:01:16** |
| **12** | 80 bp 140 bp | \*18:41, 18:73\*18:12:01-18:12:02, 18:77 18:178, 18:195  | \*15:416\*15:545, 44:445, 48:28 |
| **13** | 85 bp 265 bp | \*18:13, 18:19, 18:30\*18:44:01-18:44:02 | \*07:218, 07:220, 14:03, 14:112, 15:38:01-15:38:02, 15:368, 44:275, 51:29, 51:61:01-51:61:02, 51:82, 51:217, 51:279, 52:21:01-52:21:02, 52:24, 52:97, **A\*66:04w, C\*01:30, C\*02:142, C\*05:216:01:01-05:216:01:02, C\*06:207, C\*08:51, C\*08:114, C\*08:246, C\*12:87,****C\*14:76**\*14:01:01:01-14:02:08, 14:02:10-14:04, 14:07N, 14:09, 14:11-14:12, 14:14-14:36, 14:38-14:45, 14:47-14:52, 14:54, 14:57-14:61, 14:63-14:94, 14:97-14:101N, 14:103-14:113N, **C\*01:30, C\*06:207, C\*12:87, C\*14:76** |
| **14** | 110 bp170 bp 260 bp | \*18:102, 18:107\*18:04:01-18:04:02, 18:25, 18:68, 18:107, 18:125\*18:14  | \*44:03:12\*08:04:02, 15:01:38, 15:64:02, 35:01:57, 35:02:02, 35:20:02, 35:323:01, 44:03:12, 53:01:26\*07:218, 07:220, 08:118, 14:05, 14:13, 14:53, 35:519, 40:174, 51:64 |
| **16** | 95 bp 360 bp | \*18:19, 18:164 \*18:17N | \*08:70, 08:281, 13:31, 13:41, 13:131, 13:164, 15:01:44, 15:07:03, 15:27:03, 15:38:02, 15:78:04, 15:269, 35:14:01, 35:14:03, 35:62, 35:163, 35:343, 35:390N, 35:413, 35:512, 37:66, 39:36, 44:146, 44:275, 44:299, 46:01:03, 46:01:16, 50:47, 51:61:02, 52:21:01-52:21:02, 53:70, 54:33, 54:40, 55:21, 56:43, 58:78, **A\*25:76, A\*26:68, C\*01:09:01, C\*02:02:14, C\*06:02:27, C\*06:02:95, C\*07:22, C\*07:138, C\*07:812, C\*07:847, C\*12:03:16, C\*12:03:30** |
| **18**  | 125 bp 150 bp | \*18:32 \*18:15, 18:21w, 18:30w, 18:57:01-18:57:02, 18:101, 18:152 | \*15:192, 15:281, 35:91, 57:113, 58:138\*07:09:01w-07:09:02w, 07:11w, 07:162w, 07:237w, 07:255w, 15:03:01:01-15:03:13, 15:16:01:01-15:16:05, 15:18:01:01-15:18:10, 15:23, 15:29:01:01-15:29:01:02, 15:39:01:01-15:40:02, 15:47:01w-15:47:02w, 15:49w, 15:61-15:62, 15:64:01-15:64:02, 15:65w, 15:67, 15:69, 15:72, 15:74, 15:80, 15:93, 15:95, 15:98, 15:103, 15:108, 15:115, 15:119, 15:127, 15:131-15:132, 15:134, 15:138, 15:153, 15:158, 15:161, 15:173, 15:176, 15:194-15:195, 15:197:01-15:198, 15:210, 15:220:01:01-15:220:02, 15:220:04, 15:221w, 15:223, 15:226N, 15:238, 15:242:01-15:242:02, 15:254-15:255, 15:263, 15:266, 15:268, 15:274, 15:281-15:282, 15:288, 15:290, 15:292-15:294N, 15:306-15:307, 15:310-15:311, 15:314, 15:335, 15:337, 15:351, 15:354, 15:362, 15:364, 15:369, 15:376, 15:379w, 15:380N, 15:382, 15:388-15:389, 15:397, 15:402, 15:408, 15:414, 15:426, 15:433-15:434, 15:436, 15:444, 15:447-15:448, 15:451, 15:475, 15:487N, 15:494, 15:497, 15:502, 15:504:01-15:504:02, 15:508-15:509, 15:511, 15:517, 15:519w, 15:526-15:527, 15:548, 15:550, 15:560, 15:570-15:571, 15:573, 15:582, 15:588, 15:602, 15:607, 15:610, 15:612, 15:614:01:01-15:614:01:02, 15:616Q-15:617, 15:626, 15:643, 35:11:01-35:11:02, 35:11:03w, 35:21, 35:58, 35:99w, 35:188, 35:232w, 35:277, 35:337:01w-35:337:02w, 35:439:01:01-35:439:01:02, 35:519, 46:03w, 46:08, 46:62w, 46:83, 46:93, 48:25, 49:04:01-49:04:02, 51:37, 51:45, 53:06, 53:08:01-53:08:02, 53:28, 55:14, 55:68, 56:32, 57:13w, 57:22, 58:06, 58:19, 58:46, 58:91, **C\*01:102w, C\*03:175, C\*03:199, C\*03:245, C\*05:215w, C\*08:113, C\*15:09:01:01-15:09:02, C\*15:19, C\*15:27, C\*15:61, C\*15:107, C\*15:126, C\*15:143w, C\*15:194** |
| **19** | 105 bp 160 bp 405 bp | \*18:06, 18:142w\*18:53\*18:39 |  |
| **22** | 130 bp 150 bp 270 bp | \*18:75, 18:77\*18:33\*18:22, 18:69w, 18:105, 18:149-18:150, 18:204w | \*14:10, 14:102, 15:645, 35:21, 35:24:01-35:24:02, 35:81, 35:96, 35:109, 35:157, 35:188, 35:190, 35:233, 35:287, 37:04:01w-37:04:02w, 40:28, 44:190, 44:203:01-44:203:02, 44:275, 44:290, 51:04:01-51:04:02, 51:46, 51:56:01-51:56:03, 51:139, 51:227, 51:281, 51:310, 52:88, 53:02, 53:06, 53:28, 57:14:01-57:14:02, 58:09, 58:76, **C\*03:384, C\*15:39, C\*15:130, C\*15:195** |
| **23** | 145 bp 160 bp515 bp | \*18:51\*18:53\*18:21, 18:40, 18:61, 18:189, 18:217  | \*07:05:11, 08:04:02, **C\*06:02:101** |
| **25** | 75 bp190 bp | \*18:48\*18:35, 18:54, 18:95, 18:110, 18:195 | \*07:68:01-07:68:03, 07:214, 07:237, 07:262:01-07:262:02, 07:424, 15:138, 15:230, 15:241, 15:297, 35:66, 40:77, 40:87:01-40:87:02, 40:121, 40:158, 40:222, 40:237, 40:368, 40:422, 44:150, 46:87, 48:12, 48:14, 48:26, 48:29 |
| **26**  | 90 bp115 bp 180 bp | \*18:42\*18:92\*18:04:01-18:04:02, 18:25, 18:43, 18:107, 18:125 | \*07:05:11, 15:01:38, 15:64:02, 35:01:57, 35:02:02, 35:20:02, 39:39:02, 44:03:12, 53:01:26, **A\*24:02:37, A\*26:01:16, C\*12:301** |
| **27** | 80 bp 135 bp175 bp | \*18:37:01-18:37:02, 18:73, 18:125\*18:46, 18:74N, 18:78\*18:94N  | \*07:109, 07:356, 08:176, 48:28\*13:10, 13:26:01-13:26:02, 13:72, 13:127, 13:160, **C\*12:120** |
| **28** | 100 bp 175 bp200 bp | \*18:45\*18:38, 18:74N, 18:78\*18:95  | \*08:83, 13:10-13:11, 13:71, 13:108, 38:38, 40:91, 40:197 |
| **29**  | 65 bp 120 bp255 bp | \*18:50\*18:47\*18:18:01:01-18:18:01:02  | **\***07:218, 13:15, 14:111, 40:113, 51:16, 51:31, 51:82, 52:08, 73:01:01:01-73:01:01:02, 73:03, 78:08, **C\*02:142, C\*03:384, C\*08:246****C\*07:516** |
| **30**  | 100 bp145 bp 420 bp | \*18:31\*18:34\*18:52, 18:105  | **\***51:228, 52:28\*35:01:57, 35:02:02, 35:20:02, 35:323:01, 53:01:26, **C\*06:02:101, C\*12:301** |

**4**The following HLA-B\*18 alleles can be distinguished by the different sizes of the HLA-specific PCR product:

|  |  |  |  |
| --- | --- | --- | --- |
| **Alleles** | **Primer mix** | **Alleles** | **Primer mix** |
| B\*18:05:01:01-18:05:01:02, 18:60 | 6 | B\*18:36, 18:67 | 11 |
| B\*18:18:01:01-18:18:01:02, 18:47 | 29 | B\*18:37:02, 18:46 | 27 |
| B\*18:23N, 18:28 | 10 | B\*18:38, 18:45 | 28 |
| B\*18:34, 18:52 | 30 | B\*18:40, 18:51, 18:217 | 23 |

Abbreviations

‘w’, might be weakly amplified.