# Gel Documentation Form and Worksheet

# HLA-B*35	Lot No: 1L4	Expiry Date: 2024-07-01

# (101.522-24/03, -24u/03u)

# Name:__________________

# Sample ID:______________

# DNA Conc.(ng/ul):_________

# Test Date:_______________

# Tested By:_______________

# Review Date:_______________

# Reviewed By:_______________

# Interpretation:___________	Failed lanes: ___________ 	 Comments:_________________

Gel Picture

## PHOTO DOCUMENT

Abbreviations

ICB: Internal Control Band

AmpS: Amplicon Size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 22, 35, 41, 46 and 58 may have a tendency to giving rise to primer oligomer formation.

Primer mixes 3, 6, 7, 8, 10, 30, 31, 33, 46, 50, 53, 57, 65 and 88 have tendencies of unspecific amplifications.

Primer mixes 3, 8, 17, 20 and 93 may give rise to a lower yield of HLA-specific PCR product than the other HLA-B*35 primer mixes.

Primer mix 7 may faintly amplify the B*44 alleles.

Primer mix 14 may faintly amplify the B*37 alleles.

Primer mix 96 contains a negative control, which will amplify the majority of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 200 base pairs.

1HLA-B*35 alleles in bold lettering are listed as confirmed alleles on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.24.0, April 2016.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3 The following HLA-B*35 primer mixes have two or more product sizes:

| Primer Mix | Size of spec. PCR product | Amplified HLA-B*35 alleles | Other amplified Class I alleles |
| --- | --- | --- | --- |
| 7 | 160 bp 330 bp | *35:84, 35:90, 35:150:01-35:150:02, 35:184, 35:201, 35:217, 35:441 *35:23 | *07:177w, 08:102, 08:122, 13:36, 14:48, 15:345, 18:80, 18:102, 18:107, 27:23, 27:92, 27:157w, 38:16, 38:46, 40:08, 40:25, 40:68, 40:106, 40:166w, 40:232, 40:313, 44:12, 44:232, 44:263, 45:22, 49:26, 49:63, 50:15, 51:103, 51:119, 56:36w, 73:01:01:01w-73:02w *15:06, 15:27:01-15:27:04, 15:84, 15:109, 15:195, 15:327, 15:344, 15:398, 18:27, 40:100, 44:10, 47:01:01:02-47:03, 47:06-47:10, 82:01:01:01-82:02:02 |
| 10 | 140 bp 205 bp 240 bp | *35:83, 35:382 *35:263 *35:07 | C*07:63, C*07:421, C*07:524 *18:134, 40:283? |
| 12 | 95 bp 125 bp | *35:10, 35:13, 35:16, 35:28:01:01-35:28:01:03, 35:69, 35:80, 35:226, 35:270, 35:318 *35:92 | *13:58, 15:202, 15:239, 15:376, 18:12:01-18:12:02, 37:01:01:01-37:01:05, 37:01:07-37:05, 37:07, 37:09-37:27, 37:29-37:59, 37:62-37:66, 37:68-37:81, 37:83-37:91, 40:185, 40:209, 40:308, 40:350, 40:365, 44:83, 44:134, 44:445, 49:03, 51:310, 52:01:01:01-52:01:48, 52:03-52:15, 52:17, 52:19-52:31:02, 52:33-52:72, 52:74-52:95, 53:17:01-53:17:02, 53:28, 53:38, 78:05-78:06 |
| 16 | 95 bp 190 bp | *35:19, 35:47, 35:154, 35:217, 35:274, 35:450 *35:77 | *07:133w, 08:155, 15:143, 18:48, 18:107, 39:69, 40:08, 40:25, 40:68, 40:106, 40:166w, 40:232, 40:313, 44:12, 44:232, 44:263, 49:63, 50:15, 51:112, 53:30, 53:39 |
| 18 | 105 bp 180 bp | *35:20:01-35:20:02, 35:28:01:01-35:28:01:03, 35:72, 35:95, 35:226, 35:296, 35:439:01:01-35:439:01:02 *35:35, 35:93, 35:188, 35:190, 35:262, 35:277, 35:343 | *13:58, 15:202, 15:239, 15:308, 18:01:01:01-18:03:01:02, 18:04:01-18:06, 18:08-18:15, 18:17N-18:28, 18:30-18:32, 18:34-18:36, 18:38-18:47, 18:49-18:65, 18:67-18:78, 18:80-18:106Q, 18:108-18:124, 18:126-18:135:01:02, 18:137-18:154N, 18:156:01:01-18:177, 18:179-18:181, 18:183-18:189, 37:01:01:01-37:01:05, 37:01:07-37:05, 37:07-37:20, 37:22-37:25, 37:27, 37:29-37:56, 37:58-37:59, 37:62-37:66, 37:68-37:78, 37:80-37:81, 37:83-37:91, 40:185, 40:209, 40:308, 40:350, 40:365, 44:83, 44:134, 44:445, 45:22, 49:03, 51:07:01-51:07:02, 51:214, 51:222, 51:251, 51:310, 52:01:01:01-52:01:12, 52:01:14-52:01:48, 52:03-52:05, 52:07-52:15, 52:17-52:21:02, 52:24-52:31:02, 52:33-52:69, 52:71-52:88, 52:90-52:95, 53:09, 53:17:01-53:17:02, 53:28, 53:38, 53:40, 78:05-78:06 *08:37, 08:89, 08:107, 14:56, 15:50, 15:69, 15:93, 15:186, 15:199, 15:348, 15:368, 15:449, 15:494, 15:497, 18:01:01:01-18:01:13, 18:01:15-18:01:17, 18:01:19-18:09, 18:11-18:15, 18:17N-18:19, 18:22-18:25, 18:27-18:34, 18:37:01-18:40, 18:42-18:49, 18:51-18:55, 18:59-18:60, 18:62-18:85, 18:87-18:96, 18:98:01-18:100, 18:102-18:105, 18:107-18:121, 18:123-18:147, 18:149-18:163, 18:167, 18:170-18:188, 37:35, 37:56, 38:52, 39:07, 39:43, 39:91, 41:51, 42:09, 46:03, 46:21:01-46:21:02, 46:29, 54:09, 54:14, 55:14, 55:23, 55:68, 56:18, 56:31-56:32, 57:26, 58:05, 58:35, 58:76, 58:84, C*01:102, C*05:215, C*06:89, C*07:123, C*07:173, C*07:294, C*08:113, C*12:113, C*15:09:01:01-15:09:01:02, C*15:19, C*15:27, C*15:61, C*15:107, C*15:126, C*15:143, C*15:194, C*16:20, C*16:109 |
| 22 | 90 bp 270 bp | *35:46, 35:311 *35:21, 35:24:01-35:24:02, 35:81, 35:96, 35:109, 35:157, 35:188, 35:190, 35:233, 35:287 | *13:36, 15:02:01:01-15:02:12, 15:08:01-15:08:02, 15:11:01-15:11:07, 15:13:01-15:13:02, 15:15:01:01-15:15:01:02, 15:21:01:01-15:21:01:02, 15:31:01:01-15:31:01:03, 15:44, 15:55, 15:76w, 15:88-15:89:02, 15:112, 15:121, 15:139, 15:144, 15:148, 15:170, 15:189, 15:191, 15:194, 15:209N, 15:213-15:215, 15:223, 15:265, 15:283-15:284, 15:291, 15:297, 15:301-15:303, 15:305, 15:319, 15:325, 15:328, 15:330, 15:341, 15:345, 15:350, 15:357-15:358, 15:367, 15:374, 15:378, 15:384, 15:394, 15:399, 15:404, 15:418, 15:420, 15:425, 15:437-15:438, 15:454N, 15:460, 15:479, 15:486, 15:491, 15:495, 15:505, 15:508, 15:513, 15:537, 15:557, 51:104, 55:35w, 56:39w *14:10, 18:22, 18:69, 18:105, 18:149-18:150, 37:04:01-37:04:02, 40:28, 44:190, 44:203:01-44:203:02, 44:275, 44:290, 51:04:01-51:04:02, 51:46, 51:56:01-51:56:03, 51:139, 51:227, 51:281, 51:310, 52:88, 53:02, 53:06, 53:28, 57:14:01-57:14:02, 58:09, 58:76, C*03:384, C*15:39, C*15:130, C*15:195 |
| 25 | 135 bp 220 bp | *35:27, 35:56, 35:185, 35:203, 35:355 *35:01:34, 35:03:11, 35:09:02-35:09:03, 35:18, 35:31, 35:64:01, 35:88, 35:114, 35:118, 35:205, 35:230, 35:331 | *18:49, 18:79, 18:108, 18:132, 18:142, 37:05, 78:06, 78:09-78:10 *07:02:01:01-07:02:13, 07:02:15-07:02:24, 07:02:26-07:02:34, 07:02:36-07:06:01:02, 07:06:03-07:67N, 07:68:03-07:99, 07:101-07:213, 07:215-07:234, 07:236, 07:238-07:261, 07:263-07:358, 07:360Q-07:383, 07:385, 07:387-07:399, 08:01:02, 08:01:13, 08:20:01, 08:53:02, 08:79, 08:156, 08:213, 13:01:11, 13:21, 13:71, 15:01:35, 15:02:03, 15:03:02, 15:03:06, 15:10:03, 15:33, 15:47:02, 15:78:02, 15:78:04, 15:107, 15:116, 15:141, 15:150, 15:188, 15:248, 15:352, 15:407, 15:421, 15:507, 18:21, 18:165, 18:189, 27:01-27:04:04, 27:04:06-27:05:26, 27:05:28-27:21, 27:23-27:75, 27:77-27:89, 27:90:02-27:90:03, 27:91-27:98, 27:100-27:123, 27:125-27:126, 27:128-27:152, 27:154-27:223N, 37:02, 37:07, 37:09, 38:22, 38:63, 39:136, 40:01:01-40:01:13, 40:01:15-40:01:49, 40:01:51-40:02:19, 40:02:22-40:02:25, 40:02:27-40:06:04:02, 40:06:06-40:13, 40:14:03-40:16:01:02, 40:18-40:25, 40:27:01-40:27:02, 40:29-40:35:01, 40:36-40:40, 40:42-40:70:02, 40:72:02-40:76, 40:78-40:81, 40:83-40:86, 40:87:02-40:94, 40:96-40:114:01:02, 40:115-40:120, 40:122-40:135, 40:137-40:157, 40:159-40:183, 40:185-40:236, 40:238-40:254, 40:256N-40:325, 40:327-40:367, 40:369-40:407, 40:409-40:451, 41:08, 44:03:14, 44:31, 44:53:02, 44:150, 44:169, 44:182, 44:192:02-44:192:04, 44:343, 46:01:16, 46:06, 47:01:01:02-47:10, 48:01:01:01-48:01:10, 48:03:01:01-48:11, 48:13, 48:15-48:24, 48:26-48:50, 49:01:14, 51:01:09, 51:01:21, 51:01:25, 51:10, 51:24:02-51:24:04, 51:81, 51:197, 51:217, 51:279, 52:54, 54:11, 55:04, 55:08, 55:49, 55:86, 56:20:02, 57:01:07, 57:13, 58:27, 67:03, 81:01:01:01-81:09 |
| 26 | 210 bp 370 bp | *35:50, 35:84, 35:162, 35:197, 35:231, 35:267, 35:269, 35:411, 35:423 *35:01:23, 35:42:01 | *07:08:01-07:08:02, 07:32, 07:38, 07:207, 07:228:01-07:228:02, 08:94, 08:190w, 18:07:01-18:07:02, 18:29, 18:79, 18:136, 27:23, 37:21, 37:67, 38:06-38:07, 38:74, 39:58, 39:83, 40:08, 40:13, 40:68, 40:162, 40:220, 40:232, 42:26, 48:06, 48:31, 51:57, 51:103, 51:116, 51:218, 53:21 *18:07:01-18:07:02, 18:29, 18:79, 18:136, 27:23, 37:21, 37:67, 40:08, 40:13, 40:68?, 40:106?, 40:162?, 40:220?, 40:232? |
| 27 | 110 bp 160 bp | *35:41, 35:255 *35:84, 35:90, 35:150:01-35:150:02, 35:184, 35:201, 35:217, 35:441 | *07:177w, 08:102, 08:122, 13:36, 14:48, 15:345, 18:80, 18:102, 18:107, 27:23, 27:92, 27:157w, 38:16, 38:46, 40:08, 40:25, 40:68, 40:106, 40:166w, 40:232, 40:313, 44:12, 44:232, 44:263, 45:22, 49:26, 49:63, 50:15, 51:103, 51:119, 56:36w, 73:01:01:01w-73:02w |
| 30 | 115 bp 325 bp | *35:15:01-35:15:02, 35:33, 35:66, 35:341 *35:25-35:26, 35:49:01-35:49:02, 35:198, 35:254, 35:361 | *07:02:13, 07:02:35, 07:04:02, 07:05:10, 07:68:02, 07:151:02, 07:209, 07:262, 08:40, 08:165, 08:171, 08:205, 13:01:01:01-13:02:25, 13:02:27-13:02:32, 13:07N-13:11, 13:13:01-13:21, 13:22:02-13:23, 13:25-13:45, 13:47, 13:49N, 13:51-13:52, 13:54-13:59, 13:61, 13:63N-13:70, 13:72-13:85, 13:87-13:119, 13:121-13:127, 13:129-13:130, 13:132-13:140, 14:46, 18:10, 18:35, 18:50, 27:90:01, 38:159, 39:47, 39:50, 40:01:46, 40:02:20, 40:06:05, 40:14:01, 40:72:01, 40:77, 40:95, 40:121, 40:136, 40:237, 41:19, 44:16, 44:21, 44:64:01-44:64:02, 44:124, 45:10, 48:12, 49:07, 50:31, 51:16, 51:31, 51:34, 51:82, 51:93, 52:08, 52:67, 53:56, 54:26, 55:09, 55:22, 55:24, 56:60, 58:37, 59:09, 78:08, C*01:10w, C*02:02:14w, C*07:15w, C*07:207w, C*17:01:11 *07:209, 07:271, 13:97, 15:506, 18:37:01-18:37:02, 18:125, 27:01-27:21, 27:23-27:51, 27:53-27:66N, 27:68-27:194, 27:197, 27:199-27:223N, 37:28, 40:46?, 40:93?, 44:40, 44:44, 44:130, 44:156, 44:210:01-44:210:02, 51:267, 52:16, 55:102, 55:107, 56:05:01, 58:44, 81:01:01:01-81:01:05, 81:03?-81:04N?, 81:05, 81:06?-81:08?, 81:09, 83:01 |
| 31 | 170 bp 210 bp | *35:25, 35:124-35:125, 35:142, 35:162, 35:267, 35:320 *35:50, 35:84, 35:162, 35:197, 35:231, 35:267, 35:269, 35:411, 35:423 | *07:08:01-07:08:02, 07:32, 07:38, 07:207, 07:228:01-07:228:02, 08:01:01:01-08:03, 08:05, 08:07-08:16, 08:18-08:25, 08:27-08:53:02, 08:55, 08:57-08:60, 08:62-08:64, 08:66-08:84, 08:86N-08:101, 08:103-08:109, 08:111-08:121, 08:124-08:131, 08:133-08:145, 08:147-08:179, 08:181-08:225, 08:227, 08:229-08:238, 08:240-08:251, 14:42, 15:29:01:01-15:29:01:02, 15:180, 18:07:01-18:07:02, 18:79, 18:136, 37:21, 37:67, 38:06-38:07, 38:74, 39:58, 39:83, 42:26, 48:06, 48:31, 51:116, 51:176, 51:179, 51:218, 53:43:01-53:43:02, C*06:147 *07:08:01-07:08:02, 07:32, 07:38, 07:207, 07:228:01-07:228:02, 08:94, 08:190w, 18:07:01-18:07:02, 18:29, 18:79, 18:136, 27:23, 37:21, 37:67, 38:06-38:07, 38:74, 39:58, 39:83, 40:08, 40:13, 40:68, 40:162, 40:220, 40:232, 42:26, 48:06, 48:31, 51:57, 51:103, 51:116, 51:218, 53:21 |
| 33 | 205 bp 310 bp | *35:45, 35:71, 35:169, 35:309, 35:452 *35:94 | *07:20, 07:140, 07:153, 15:14, 15:91, 15:131, 15:161, 15:518, 15:551, 18:56, 18:169, 38:158, 40:132, 40:178, 44:10, 44:14, 44:17, 44:31, 44:43:01-44:43:02, 44:45, 44:51, 44:62, 44:77, 44:82, 44:107, 44:123, 44:135, 44:184, 44:213, 44:230, 44:339, 45:09, 45:14, 46:17, 51:23, 51:108, 52:15, 53:22, 57:07, 57:93, 57:119, 58:07, 82:01:01:01-82:03 |
| 34 | 185 bp 230 bp | *35:214 *35:48, 35:433 | *15:334, 44:133, 46:70 *37:39, 37:48, 44:209, 53:07, 58:121, C*03:90, C*03:103, C*03:300, C*03:474, C*07:748 |
| 35 | 240 bp 365 bp | *35:40N *35:68:01, 35:151:02, 35:292 | *08:12:03, 08:107, 51:06:02, 53:47 |
| 36 | 110 bp 180 bp 230 bp | *35:79 *35:70, 35:220 *35:35, 35:468 | *40:236 *15:265, 37:40, 38:20:01-38:20:02, 39:42, 58:84, C*15:51 |
| 42 | 115 bp 155 bp 190 bp 250 bp | *35:97 *35:53N *35:91 *35:65Q | *51:275 *44:337, 57:113, C*03:347 |
| 43 | 205 bp 255 bp 400 bp | *35:54 *35:89 *35:390N | *08:151, 08:242, 14:42, 51:140, C*06:147 *08:135 |
| 44 | 110 bp 145 bp 300 bp | *35:78 *35:55 | *51:272 *13:58, 18:09, 37:10, 44:83, 44:134, 44:445, 51:54, 51:78:01-51:78:02, 52:20, 53:09, 53:11-53:13, 53:31, 53:36, 53:38 |
| 51 | 65 bp 155 bp 185 bp | *35:145N *35:100, 35:102 *35:214 | *18:158, C*07:719 *15:334, 44:133, 46:70 |
| 52 | 220 bp 250 bp | *35:141 *35:104 |  |
| 53 | 135 bp 220 bp | *35:103 *35:129Nw | *51:266 |
| 54 | 160 bp 190 bp | *35:106, 35:216N *35:110 | *58:77 |
| 55 | 205 bp 300 bp | *35:263 *35:108:01-35:108:02 | *18:134, 40:283? *37:91, 40:228?, 44:248, 51:255, 53:26 |
| 56 | 140 bp 225 bp | *35:128 *35:107, 35:158 | *08:179, 15:306 |
| 57 | 150 bp 185 bp | *35:149, 35:357 *35:111 | *15:398 |
| 58 | 110 bp 160 bp 200 bp | *35:139 *35:115, 35:260 *35:112 | *15:326, 40:321, 44:157, 55:74, C*05:30 *08:37, 08:89, 08:107, 15:51, 15:179:01-15:179:02, 15:199, 15:218Q, 15:518, 18:56, 18:113, 37:35, 37:56, 41:51, 42:09, 44:17, 44:230, 45:09, 53:22, 57:24, 57:119 |
| 59 | 75 bp 150 bp | *35:116 *35:127 | *15:359 *50:54, 55:100, 57:21 |
| 60 | 90 bp 135 bp | *35:121 *35:120 |  |
| 61 | 110 bp 175 bp 265 bp | *35:139 *35:122 *35:123 | *15:326, 40:321, 44:157, 55:74, C*05:30 *15:277-15:278, 40:339, 56:29 *50:67 |
| 62 | 140 bp 260 bp | *35:144 *35:132, 35:246 | C*15:56 |
| 64 | 90 bp 245 bp | *35:130N *35:134N | C*15:160N |
| 66 | 90 bp 115 bp | *35:133 *35:105, 35:113, 35:159, 35:397 |  |
| 67 | 105 bp 200 bp | *35:138 *35:191 | *15:413, 44:26, 47:07 *40:227, 49:32, 50:63 |
| 69 | 95 bp 325 bp | *35:143 *35:147 | *13:101, 51:73, 55:26 |
| 71 | 120 bp 160 bp | *35:126 *35:119 | *13:111 |
| 75 | 90 bp 135 bp | *35:176 *35:27, 35:56, 35:185, 35:203, 35:355 | *18:49, 18:79, 18:108, 18:132, 18:142, 37:05, 78:06, 78:09-78:10 |
| 76 | 145 bp 220 bp 375 bp | *35:271 *35:193 *35:161 | *18:66, 46:47 C*08:89N *15:198 |
| 77 | 135 bp 180 bp 385 bp | *35:227 *35:245, 35:389 *35:36 |  |
| 81 | 80 bp 120 bp 260 bp | *35:219 *35:228 *35:293 | *15:158, 58:42 |
| 83 | 95 bp 180 bp | *35:250 *35:245, 35:389 | *51:165 |
| 89 | 110 bp 240 bp | *35:381N *35:236, 35:259, 35:333Q | *53:48N, C*15:177N *37:16Q, 44:29, 44:89, 44:160Q, 44:285, C*03:215, C*15:96Q |

4The following HLA-B*35 alleles can be distinguished by the different sizes of the HLA-specific PCR product:

| Alleles | Primer mix |
| --- | --- |
| B*35:122, 35:123 | 61 |
| B*35:130N, 35:134N | 64 |
| B*35:132, 35:144 | 62 |
| B*35:219, 35:293 | 81 |
| B*35:236, 35:381N | 89 |

Abbreviations

w: might be weakly amplified.

?: nucleotide sequence of the primer matching region not known.