Gel Documentation Form and Worksheet

HLA-B*37 (101.541-06/06u) 	Lot No: 4D3	Expiry Date: 2018-07-01

Name:__________________

Sample ID:______________

DNA Conc.(ng/ul):_________

Test  Date:_______________

Tested By:_______________

Review Date:_______________

Reviewed By:_______________

Interpretation:_______________

Failed lanes:    _________

Comments:    ___________________

Gel Picture

PHOTO DOCUMENT

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 10 and 21 may have tendencies of unspecific amplifications.

Primer mixes 23 has a tendency of giving rise to primer oligomer formation.

Primer mix 26 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by control primer pairs. PCR product sizes range from 75 to 200 base pairs. The PCR product generated by the control primer pair is 430 base pairs.

1HLA-B*37 alleles in bold lettering are listed as confirmed alleles on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.22.0, October 2015.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3Primer mix 2: Specific PCR fragment of 95 bp in the B*37:02, 07:197, 15:342, 27:01-27:02-02, 27:04-01-27:04:05, 27:05:21, 27:08, 27:12, 27:15, 17:18, 27:23, 27:25-27:26, 27:31, 27:36, 27:40, 27:42, 27:44, 27:49, 27:52-27:53, 27:57, 27:61-27:63, 27:65N-27:69, 27:75, 27:77, 27:79, 27:83, 27:86, 27:89,  27:95, 27:100, 27:103, 27:105, 27:108-27:110, 27:112-27:115, 27:119-27:120, 27:126, 27:131, 27:134, 27:142 and the B*47:04-47:05 and 48:26 alleles.

Specific PCR fragment of 135 bp in the B*37:28 and the B*07:27, 07:236, 08:126, 08:131, 08:131, 08:139, 27:05:16, 27:07:01-27:07:04, 27:09, 27:14, 27:19, 27:27, 27:32, 27:34-27:35, 27:41, 27:43, 27:46, 27:70, 27:76, 27:78, 27:81, 27:127, 27:130, 27:136 and 38:17 alleles.

Specific PCR fragment of 95 and 135 bp in the 27:03, 27:05:02-27:05:15, 27:05:17-27:05:20, 27:05:22-27:05:31, 27:10, 27:13, 27:16-27:17, 27:28-27:29, 27:37-27:39, 27:45, 27:47-27:48, 27:50:01-27:51, 27:54-27:56, 27:58-27:60, 27:64N, 27:71-27:74, 27:80, 27:82, 27:84-27:85, 27:87-27:88, 27:90:01-27:90:03, 27:92-27:94N, 27:96:01-27:99, 27:101, 27:104, 27:111, 27:116-27:118, 27:121-27:124, 27:128-27:129, 27:132-27:133, 27:135-27:137, 27:139-27:141 and 27:143-27:145 alleles.

Primer mix 3: Specific PCR fragment of 200 bp in the B*37:13 allele. Specific PCR fragment of 235 bp in the B*37:03N allele.

Primer mix 4: Specific PCR fragment of 200 bp in the B*37:13 allele. Specific PCR fragment of 240 bp in the *37:04:01-37:04:02, 37:39 and 37:48 and the B*15:32:01-15:32:02, 15:299, 18:18, 35:48, 44:209 and 53:07 and in the C*03:90, 03:103, 03:178, 03:300, C*07:03, 07:186, 07:242, 07:265 alleles.

Primer mix 7: Specific PCR fragment of 130bp in the B *37:06:01-37:06:02 and 37:42N and the B*40:188, 47:01:01:01-47:01:02 and 47:05-47:10 alleles. Specific PCR fragment of 195 bp in the B *37:36 allele.

Primer mix 9: Specific PCR fragment of 80 bp in the B*37:08 and the B*07:27, 07:236, 08:126, 08:131, 08:139, 18:54, 27:92, 38:17 and 53:03 and in the A*68:112:02 allele. Specific PCR fragment of 125 bp in the B*37:42N allele. Specific PCR fragment of 155 bp in the B*37:30N allele. Specific PCR fragment of 215 bp in the B*37:22 allele. Specific PCR fragment of 495 bp in the B*15:209N allele.

Primer mix 10: Specific PCR fragment of 110 bp in the B*37:26 allele. Specific PCR fragment of 145 bp in the B*37:14 allele. Specific PCR fragment of 310 bp in the B*37:46 allele.

Primer mix 11: Specific PCR fragment of 95 bp in the B*37:18 allele. Specific PCR fragment of 215 bp in the B*37:09 and the B*07:04, 07:19, 07:25, 07:146, 08:01:01-08:05, 08:07-08:08N, 08:10-08:11, 08:14-08:15, 08:17-08:19N, 08:21-08:24, 08:26:01-08:39, 08:41-08:48, 08:50-08:54, 08:56:01-08:59:02, 08:61-08:69, 08:71-08:78, 08:80-08:83, 08:85-08:88, 08:90-08:106, 08:108-08:128, 08:130-08:142, 08:144-08:155, 35:87, 41:02:01-41:02:06, 41:04, 41:10-41:11, 41:13, 41:15, 41:18-41:19, 41:23-41:24, 41:27, 41:31, 41:36, 41:38-41:41, 42:01:01-42:02:01:02, 42:05:01-42:12, 42:14-42:18, 42:20-42:21, 44:106, 44:158, 44:166 and 44:213 alleles. Specific PCR fragment of 355 bp in the B*37:43 allele.

Primer mix 13: Specific PCR fragment of 170 bp in the B*37:15 allele. Specific PCR fragment of 195 bp in the B*37:33N and the 27:94N allele. Specific PCR fragment of 255 bp in the B*37:10 and the B*13:23, 13:55, 18:09, 27:01, 27:142, 40:47, 40:96, 40:157, 40:201, 44:15, 44:55, 44:103, 44:188, 44:212 and 49:02 alleles.

Primer mix 14: Specific PCR fragment of 110 bp in the B*37:47 and the B*27:04:01-27:04:05, 27:06, 27:24-27:25, 27:66N, 27:68-27:69, 27:100, 27:103, 27:105, 27:109, 27:112-27:115 and 27:138 alleles. Specific PCR fragment of 245 bp in the B*37:11 and 37:37 and the B*07:143, 08:62, 15:71, 15:175, 15:204, 15:210, 15:212, 15:225, 15:344, 18:01:01:01-18:01:12, 18:01:14-18:01:20, 18:01:22-18:03, 18:05-18:08, 18:10-18:15, 18:17N-18:24, 18:26-18:42, 18:44:01-18:48, 18:50-18:53, 18:55-18:66, 18:68-18:78, 18:80-18:106, 18:109-18:116, 27:08, 27:12, 27:18, 27:26, 27:33, 27:40, 27:44, 35:50, 35:63, 35:84, 35:162, 35:217, 35:231, 35:280, 39:31, 39:101, 40:01:01-40:01:04, 40:01:06-40:09, 40:11:01-40:11:02, 40:14:01-40:16, 40:18, 40:20, 40:22N-40:36, 40:38-40:39, 40:42-40:46, 40:48, 40:50-40:57, 40:59-40:75, 40:77-40:95, 40:97-40:108, 40:111-40:116, 40:118N-40:128, 40:130:01-40:134, 40:136, 40:138-40:156, 40:158-40:163, 40:165-40:187, 40:189-40:190, 40:192-40:200, 40:202-40:209, 40:211-40:218, 40:220-40:228, 40:230-40:255, 40:257-40:280, 40:282-40:291N, 40:293-40:305, 40:307-40:317, 41:01:01-41:04, 41:06-41:32, 41:34-41:41, 42:02:01:01-42:02:01:02, 42:09, 42:17-42:18, 44:131, 45:01:01-45:17, 46:14, 48:22, 48:30, 50:01:01-50:02, 50:04-50:12, 50:14-50:20, 50:31-50:42, 54:02, 55:16, 56:35 and 73:01-73:02 and in the C*03:278, C*16:86 alleles.

Primer mix 15: Specific PCR fragment of 135 bp in the B*37:11, 37:37 and the B*07:65, 07:134, 08:32, 08:133, 15:202, 15:239, 15:308, 18:01:01:01-18:01:20, 18:01:22-18:08, 18:10-18:15, 18:17N-18:36, 18:38-18:47, 18:50-18:53, 18:55-18:65, 18:68-18:78, 18:80-18:106, 18:109-18:116, 35:01:01:01-35:02:02, 35:02:04-35:03:02, 35:03:04-35:08:01, 35:08:03-35:08:04, 35:08:06-35:18, 35:20:01-35:24:02, 35:28-35:45, 35:48, 35:50-35:55, 35:57-35:62, 35:64:01-35:72, 35:74-35:153, 35:155-35:197, 35:199-35:202, 35:204-35:206, 35:208-35:209, 35:211-35:216N, 35:218-35:221, 35:223-35:253, 35:255-35:270, 35:272-35:273, 35:275-35:298, 35:300, 39:19:01-39:19:02, 40:185, 40:209, 40:308, 46:40:01-46:40:02, 56:06, 56:45, 58:64, 78:01:01-78:05 and 78:07-78:08 alleles. Specific PCR fragment of 315 bp in the B*37:45 allele.

Primer mix 16: Specific PCR fragment of 170 bp in the B*37:09, 37:12, 37:35, 37:56 and the B*07:04, 07:19, 07:25, 07:146, 08:01:01-08:01:22, 08:01:24-08:05, 08:07-08:12:03, 08:15-08:19N, 08:21-08:24, 08:26:01-08:27, 08:29-08:39, 08:41-08:48, 08:50-08:54, 08:56:01-08:69, 08:71-08:73, 08:75-08:78, 08:80-08:86N, 08:88-08:110, 08:112-08:142, 08:144-08:155, 15:51, 15:83, 15:179:01-15:179:02, 15:199, 15:218Q, 18:56, 18:113, 35:38, 35:87, 35:115, 35:169, 35:260, 38:30, 40:136, 40:231, 40:279, 41:01:01-41:09, 41:11-41:41, 42:01:01-42:02:01:02, 42:04-42:18, 42:21, 44:14, 44:17, 44:51, 44:62, 44:123, 44:166, 44:213, 44:230, 45:09, 45:14, 51:08:01-51:08:02, 51:20, 51:36, 51:44N, 51:97, 51:141, 51:153, 52:19, 53:22, 55:20, 55:56, 56:13, 57:09, 57:24 and 82:01-82:03 alleles. Specific PCR fragment of 195 bp in the B*37:20 allele.

Primer mix 17: Specific PCR fragment of 165 bp in the B*37:25 and the B*45:03 alleles. Specific PCR fragment of 195 bp in the B*37:20 allele.

Primer mix 19: Specific PCR fragment of 215 bp in the B*37:21 and B*08:62, 18:07:01-18:07:02, 18:29, 18:79, 27:23, 35:50, 35:84, 35:162, 35:217, 35:231, 35:280, 40:07-40:08, 40:13, 40:68, 40:106, 40:162, 40:220, 40:232, 51:103, 51:116 and 51:151 alleles. Specific PCR fragment of 315 bp in the B*37:45 allele.

Primer mix 20: Specific PCR fragment of 195 bp in the B*37:36 allele. Specific PCR fragment of 230 bp in the B*37:16Q and 37:39, 37:48 and the B*15:32:01-15:32:02, 15:299, 35:48, 44:209, 53:07 and in the C*03:90, 03:103, 03:178, 03:300, C*07:03, 07:186, 07:242, 07:265 alleles.

Primer mix 21: Specific PCR fragment of 165 bp in the B*37:27 and the B*27:57 alleles. Specific PCR fragment of 195 bp in the B*37:33N and the 27:94N alleles. Specific PCR fragment of 270 bp in the B*37:24 and the B*27:122 and 40:152 alleles.

Primer mix 23: Specific PCR fragment of 90 bp in the B*37:19:01 and the B*14:10, 35:59:01, 35:242, 38:20, 39:42, 39:106 and 44:117 and in the C*15:05:08 and C*17:01:10 alleles. Specific PCR fragment of 350 bp in the B*37:23 and 37:43 and the B*27:69, 40:213, 51:32 and 58:52 alleles.

‘w’, might be weakly amplified.

4The following HLA-B*37 alleles can be distinguished by the different sizes of the HLA-specific PCR product:

| Alleles | Primer mix |
| --- | --- |
| B*37:24, 37:27 | 21 |
| B*37:26, 37:46 | 10 |

The HLA-B*37 subtyping kit cannot distinguish the B*37:01:01 to B*37:01:11, the B*37:04:01 and 37:04:02 or the B*37:06:01-37:06:02 alleles.