HLA-B\*37 (101.541-06/06u) Lot No: 5N1 Expiry Date: 2025-10-01

Name:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

Sample ID:\_\_\_\_\_\_\_\_\_\_\_\_\_\_

DNA Conc.(ng/ul):\_\_\_\_\_\_\_\_\_

Test Date:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

Tested By:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

Review Date:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

Reviewed By:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

*Interpretation:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_*

*Failed lanes: \_\_\_\_\_\_\_\_\_*

*Comments: \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_*

**Gel Picture**

|  |
| --- |
| PHOTO DOCUMENT |





Abbrevations

ICB: Internal Control Band

AmpS: Amplicon Size

**Notes:**

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 10 and 21 may have tendency of unspecific amplification.

Primer mixes 23 has a tendency of giving rise to primer oligomer formation.

Primer mixes 17 and 19 may give rise to a lower yield of HLA-specific PCR product than the other B\*37 primer mixes.

Primer mix 31 contains a negative control, which will amplify the majority of HLA amplicons as well as the amplicons generated by control primer pairs. PCR product sizes range from 75 to 200 base pairs. The PCR product generated by the control primer pair is 200 base pairs.











**1**HLA-B\*37 alleles in bold lettering are listed as confirmed alleles on the IMGT/HLA web page [www.ebi.ac.uk/imgt/hla](http://www.ebi.ac.uk/imgt/hla), release 3.29.0, August 2017.

**2**Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page <http://hla.alleles.org/alleles/deleted.html>.

**3**The following HLA-B\*37 primer mixes have two or more product sizes:

|  |  |  |  |
| --- | --- | --- | --- |
| Primer Mix | Size of spec. PCR product | Amplified HLA-B\*37 alleles | Other amplified HLA-B alleles |
|  **2** | 95 bp 135 bp | \*37:02\*37:28 | \*07:197, 15:342, 18:165, 27:01-27:05:15, 27:05:17-27:05:37, 27:05:39-27:05:55, 27:08, 27:10, 27:12:01:01-27:13:02, 27:15-27:18, 27:23, 27:25-27:26, 27:28-27:29, 27:31, 27:36-27:40, 27:42, 27:44-27:45, 27:47-27:69, 27:71-27:75, 27:77, 27:79-27:80, 27:82-27:90:04, 27:92-27:101, 27:103-27:105, 27:108-27:124, 27:126, 27:128-27:129, 27:131-27:135, 27:137, 27:139-27:149, 27:151-27:152, 27:155-27:163, 27:165-27:167, 27:169-27:179, 27:181-27:186, 27:188-27:191, 27:193-27:207, 27:209-27:214, 27:216-27:217, 27:220-27:221, 27:223N-27:232, 27:234-27:248, 40:429, 47:04-47:05, 48:26, **A\*32:89**\*07:27, 07:236, 08:126, 08:131, 08:139, 27:03, 27:05:02:01-27:05:20, 27:05:22-27:05:55, 27:07:01-27:07:06, 27:09-27:10, 27:13:01-27:14, 27:16-27:17, 27:19:01:01-27:19:01:02, 27:27-27:29, 27:32, 27:34-27:35, 27:37-27:39, 27:41, 27:43, 27:45-27:48, 27:50:01-27:51, 27:54-27:56, 27:58-27:60, 27:64N, 27:70-27:74, 27:76, 27:78, 27:80-27:82, 27:84-27:85, 27:87-27:88, 27:90:01-27:90:04, 27:92-27:94N, 27:96:01-27:99, 27:101, 27:104, 27:111, 27:116-27:118, 27:121-27:124, 27:127-27:130, 27:132-27:133, 27:135-27:137, 27:139-27:141, 27:143-27:146, 27:148, 27:150-27:152, 27:155, 27:159-27:162, 27:164, 27:166-27:167, 27:169-27:170, 27:174-27:175, 27:177-27:180, 27:182, 27:184-27:187, 27:189-27:191, 27:193-27:196, 27:199-27:201, 27:205-27:207, 27:209, 27:212N, 27:216-27:221, 27:223N-27:225N, 27:227-27:228, 27:230-27:231, 27:233-27:235, 27:237-27:239, 27:241, 27:243N-27:244, 27:248, 38:17, 55:90 |
|  **3** | 200 bp 235 bp | \*37:13\*37:03N |  |
|  **4** | 200 bp 240 bp | \*37:13\*37:04:01-37:04:02, 37:39, 37:48, 37:97 | \*15:32:01-15:32:02, 15:299, 15:583, 18:18:01:01-18:18:01:02, 18:204, 35:48, 35:433, 44:209, 53:07, 53:65, 58:121, **C\*03:90, C\*03:103, C\*03:178, C\*03:300, C\*03:474, C\*07:03, C\*07:186, C\*07:242, C\*07:265, C\*07:516, C\*07:748** |
|  **7** | 130 bp 195 bp | \*37:06:01-37:06:02, 37:42N\*37:36 | \*40:188, 44:257, 47:01:01:02-47:01:05, 47:05-47:10, 47:12 |
|  **9** | 80 bp125 bp 155 bp 215 bp 495 bp | \*37:08 \*37:42N\*37:30N\*37:22:01-37:22:02 | \*07:27, 07:236, 08:126, 08:131, 08:139, 18:54, 27:92, 38:17, 53:03, 53:55, 55:90, **A\*68:112:02**\*27:05:18:01-27:05:18:02\*15:209N, 39:139N |
| **10** | 110 bp 145 bp310 bp | \*37:26\*37:14\*37:46 |  |
| **11** | 95 bp 215 bp355 bp | \*37:18\*37:09\*37:43 | \*07:04:01-07:04:02, 07:19, 07:25, 07:146, 07:349, 08:01:01:01-08:05, 08:07-08:08N, 08:10-08:11, 08:14-08:15, 08:17-08:19N, 08:21-08:24, 08:26:01-08:39, 08:41-08:48, 08:50-08:54:02, 08:56:01-08:59:02, 08:61-08:69, 08:71-08:78, 08:80-08:83, 08:85-08:88, 08:90-08:106, 08:108-08:128, 08:130-08:142, 08:144-08:155, 08:157-08:170, 08:172-08:180, 08:182-08:183, 08:185-08:204, 08:206-08:231, 08:233-08:253, 08:256-08:280, 35:87, 41:02:01:01-41:02:09, 41:04, 41:10-41:11, 41:13, 41:15, 41:18-41:19, 41:23-41:24, 41:27, 41:31, 41:36, 41:38-41:43, 41:45N-41:47, 41:49, 41:51-41:52, 41:55, 41:58, 41:65, 41:68-41:69, 41:71-41:72, 41:74, 42:01:01:01-42:02:02, 42:05:01-42:12, 42:14-42:18, 42:20-42:31, 44:106, 44:158, 44:166, 44:213 |
| **14** | 110 bp245 bp | \*37:47 \*37:11, 37:37, 37:67 | \*27:04:01-27:04:06, 27:06:01:01-27:06:01:02, 27:15, 27:20, 27:21:02, 27:24-27:25, 27:66N, 27:68-27:69, 27:91, 27:100, 27:103, 27:105, 27:109, 27:112-27:115, 27:138, 27:147, 27:149, 27:165, 27:173, 27:192, 27:208, 27:210-27:211, 27:214, 27:222, 27:226, 27:229, 27:240, 27:247\*07:143, 08:62, 08:226, 08:242, 15:71, 15:175, 15:204, 15:210, 15:212, 15:225, 15:344, 15:381, 15:426, 15:559, 15:571, 18:01:01:01-18:01:12, 18:01:14-18:01:20, 18:01:22-18:03:02, 18:05:01:01-18:08, 18:10-18:15, 18:17N-18:24, 18:26-18:42, 18:44:01-18:48, 18:50-18:53, 18:55-18:66, 18:68-18:78, 18:80-18:106Q, 18:109-18:124, 18:126-18:131:01:02, 18:133-18:135:01:02, 18:137-18:141, 18:143-18:146, 18:148-18:193, 18:196-18:197, 18:199-18:212, 27:08, 27:12:01:01-27:12:01:03, 27:18, 27:26, 27:33, 27:40, 27:44, 27:153, 27:204, 27:242, 35:50, 35:63, 35:84, 35:162, 35:217, 35:231, 35:280, 35:317, 35:323, 35:441, 39:31:01:01-39:31:01:02, 39:101, 39:128, 40:01:01-40:01:04, 40:01:06-40:06:14, 40:06:16-40:09, 40:11:01:01-40:11:02, 40:14:01-40:16:01:02, 40:18, 40:20:01:01-40:20:01:02, 40:22N-40:36, 40:38-40:39, 40:42-40:46, 40:48, 40:50:01:01-40:57, 40:59-40:75, 40:77-40:95, 40:97-40:108, 40:111-40:116, 40:118N-40:128, 40:130:01-40:134, 40:136, 40:138-40:156, 40:158-40:163, 40:165-40:187, 40:189-40:190, 40:192-40:200, 40:202-40:209, 40:211-40:218, 40:220-40:228, 40:230-40:255, 40:257-40:280, 40:282-40:291N, 40:293-40:305, 40:307-40:317, 40:319-40:324, 40:326-40:327, 40:329-40:334, 40:336-40:339, 40:341, 40:343, 40:345N-40:364, 40:366-40:393, 40:395-40:398, 40:400-40:422, 40:424-40:448, 40:450-40:456, 40:458-40:461, 40:463-40:495, 41:01:01:01-41:04, 41:06-41:32, 41:34-41:45N, 41:48-41:59, 41:61-41:74, 42:02:01:01-42:02:02, 42:09, 42:17-42:18, 42:26-42:27, 44:131, 45:01:01:01-45:25, 45:27, 46:14, 48:22, 48:30, 50:01:01:01-50:01:17, 50:01:19-50:02:01:02, 50:04-50:12, 50:14-50:20, 50:31-50:42, 50:44-50:75, 54:02, 55:16, 56:35, 73:01:01:01-73:03, **C\*03:278, C\*03:504, C\*07:714, C\*07:724, C\*16:86** |
| **15** | 135 bp315 bp | \*37:11, 37:37, 37:67\*37:45 | \*07:65, 07:134, 07:400, 08:32, 08:133, 08:180, 08:239, 15:202, 15:239, 15:308, 15:376, 18:01:01:01-18:01:20, 18:01:22-18:03:01:03, 18:04:01-18:08, 18:10-18:15, 18:17N-18:36, 18:38-18:47, 18:50-18:53, 18:55-18:65, 18:68-18:78, 18:80-18:106Q, 18:109-18:124, 18:126-18:131:01:02, 18:133-18:135:01:02, 18:137-18:141, 18:143-18:146, 18:148-18:177, 18:179-18:193, 18:196-18:197, 18:199-18:212, 35:01:01:01-35:01:51, 35:01:53-35:02:02, 35:02:04-35:02:08, 35:02:10-35:02:13, 35:02:15-35:03:02, 35:03:04-35:08:01:10, 35:08:03-35:08:04, 35:08:06-35:18, 35:20:01-35:24:02, 35:28:01:01-35:43:01, 35:43:03-35:45, 35:48, 35:50-35:55, 35:57-35:62, 35:64:01-35:72, 35:74-35:153, 35:155-35:197, 35:199-35:202, 35:204-35:206, 35:208:01-35:209, 35:211-35:216N, 35:218-35:221, 35:223-35:253, 35:255-35:270, 35:272-35:273, 35:275-35:298, 35:300-35:310, 35:312-35:316, 35:318-35:328, 35:330-35:354, 35:356-35:360, 35:362-35:449, 35:451-35:470, 35:472-35:524, 35:526-35:529, 39:19:01-39:19:02, 40:185, 40:209, 40:308, 40:350, 40:365, 45:22, 46:40:01-46:40:02, 56:06, 56:45, 58:64, 78:01:01:02-78:05, 78:07-78:08 |
| **16** | 170 bp195 bp | \*37:09, 37:12, 37:35, 37:56, 37:77, 37:84\*37:20 | \*07:04:01-07:04:02, 07:19, 07:25, 07:146, 07:349, 08:01:01:01-08:01:22, 08:01:24-08:05, 08:07-08:12:03, 08:15-08:19N, 08:21-08:24, 08:26:01-08:27, 08:29-08:39, 08:41-08:48, 08:50-08:54:02, 08:56:01-08:69, 08:71-08:73, 08:75-08:78, 08:80-08:86N, 08:88-08:110, 08:112-08:142, 08:144-08:155, 08:157-08:170, 08:172-08:199, 08:201-08:204, 08:206-08:226, 08:228-08:244, 08:246-08:250, 08:252N-08:254, 08:256-08:263, 08:265-08:280, 13:154, 15:51, 15:83:01:01-15:83:01:02, 15:179:01-15:179:02, 15:199, 15:218Q, 15:518, 18:56, 18:113, 35:38, 35:87, 35:115, 35:169, 35:260, 38:30, 39:156, 40:136, 40:231, 40:279, 41:01:01:01-41:09, 41:11-41:51, 41:53-41:61, 41:63-41:74, 42:01:01:01-42:02:02, 42:04-42:18, 42:21-42:31, 44:14, 44:17, 44:51, 44:62, 44:123, 44:166, 44:213, 44:230, 45:09, 45:14, 51:08:01:01-51:08:05, 51:20, 51:36, 51:44N, 51:97, 51:141, 51:153, 51:263, 51:275, 51:342, 52:19, 53:22, 55:20, 55:56, 56:13, 57:09, 57:24, 57:119, 82:01:01:01-82:04, **C\*07:626** |
| **17** | 165 bp195 bp | \*37:25\*37:20 | \*45:03 |
| **19** | 215 bp315 bp | \*37:21, 37:67\*37:45  | \*08:62, 08:242, 18:07:01-18:07:03:02, 18:29, 18:79, 18:136, 27:23, 35:50, 35:84, 35:162, 35:217, 35:231, 35:280, 35:323, 35:441, 40:07-40:08, 40:13, 40:68, 40:106, 40:162, 40:220, 40:232, 41:46, 42:26, 49:67, 51:103, 51:116, 51:151 |
| **20** | 195 bp 230 bp | \*37:36\*37:16Q, 37:39, 37:48, 37:97 | \*15:32:01-15:32:02, 15:299, 15:583, 35:48, 35:433, 44:209, 53:07, 53:65, 58:121, **C\*03:90, C\*03:103, C\*03:178, C\*03:300, C\*03:474, C\*07:03, C\*07:186, C\*07:242, C\*07:265, C\*07:494Q, C\*07:748** |
| **21** | 165 bp 195 bp 270 bp | \*37:27\*37:33N\*37:24  | \*27:57\*27:94N\*27:122, 40:152 |
| **23**  | 90 bp 350 bp | \*37:19:01 \*37:23, 37:43  | \*14:10, 15:460, 35:59:01, 35:242, 38:20:01, 39:42, 39:106, 44:117, 44:251, **C\*03:26:02, C\*15:140, C\*15:152:03, C\*17:01:10**\*07:329, 13:151, 15:227, 18:183, 27:69, 27:192, 40:213, 40:485, 42:28, 44:338, 51:32, 58:52 |

**4**The following HLA-B\*37 alleles can be distinguished by the different sizes of the HLA-specific PCR product:

|  |  |
| --- | --- |
| **Alleles** | **Primer mix** |
| B\*37:24, 37:27 | 21 |
| B\*37:26, 37:46 | 10 |

**5**The following alleles will give rise to identical amplification patterns with the HLA-B\*37 subtyping kit. These alleles can be distinguished by the HLA-B low resolution kit and/or the HLA-B\*18 subtyping kits:

|  |
| --- |
| **Alleles** |
| B\*37:40, 37:44, 37:66, 37:83, 37:94, B\*18:195 |

Abbreviations

w: may be weakly amplified.