PHOTO DOCUMENT

Olerup SSP AB

Olerup SSPTM HLA-B*39

Prod. No:	101.566-12/04 – including Taq pol.

101.566-12u/04u– without Taq pol.

Lot No:    	Y49

Expiry Date:    	2009-October-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ug/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 32 primer mixes used for HLA-B*39 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified HLA-B*39 alleles3 | Other amplified HLA-Class I alleles4 |
| --- | --- | --- | --- | --- |
| 15 | 230 bp | 800 bp | 39010101- 39010102L, 390103-3905, 3907-3920, 3922-3929, 3931, 3935, 3937-3941 | 1569, 380101-3807, 3808weakly, 3809, 3811-3817, 670101-6702 |
| 25 | 170 bp | 1070 bp | 39010101- 39010102L, 390103-390104, 3903-390602, 3907?, 3909, 3911, 3912, 3914, 3915, 3918, 3919, 3922, 3924-3938Q, 3940N-3943 | 1401-1407N, 1509, 1510, 1518, 1523, 1537, 1551, 1552, 1572, 1580, 1590, 1593, 1599, 9508, 9514, 9515, 9519, 9524, 9533, 9534, 380101-380202, 3804, 3805, 3808-3815, 3817 |
| 3 | 200 bp | 1070 bp | 3905, 3908, 3911, 391301-391302, 3920, 3937, 3943 | 0716, 0727, 0737, 0738, 0750, 0803, 0823, 1503, 1509, 1510, 1518, 1523, 1529, 1537, 1538, 1547, 1549, 1551, 1552, 1554, 1562, 1564, 1568, 1569, 1572, 1574, 1580, 1590, 1591, 1593, 1598, 1599, 9503, 9508, 9514, 9515, 9519, 9523, 9524, 9527, 9531-9534, 180101-1815, 1817N-1827, 3521, 3524, 3525, 370101-3713, 380101-380202, 3804-3812, 3815, 3817, 4012, 480101-4819, 5137, 5145, 5302, 5306, 5714, 5809, 5812 |
| 46 | 145 bp | 1070 bp | 3926 |  |
| 56 | 105 bp | 800 bp | 39010101-39010102L, 390103-390201, 3903-3909, 3911, 3912, 3914, 3915, 3918, 3922, 3924-3943 | 1401-1407N, 3526, 670101 |
| 65,7 | 170 bp | 800 bp | 390201-390202, 3908, 391301-391302, 3923, 3939 | 0727, 0750, 0804, 0817, 1503, 1547, 1554, 1561, 1562, 1564, 1568, 1569, 1591, 1598, 9503, 9523, 9527, 9531, 9532, 180101-1803, 1805, 1806, 1808-1815, 1817N-1827, 370101-3713, 3803, 4012, 480101-4805, 4807-4819 |
| 78 | 165, 220 bp | 1070 bp | 3903, 3914, 3917, 3924, 3929, 3937 |  |
| 8 | 180 bp | 1070 bp | 3904, 3907? | 3526 |
| 95 | 230 bp | 800 bp | 390601-390602 |  |
| 105,6,9 | 105, 165 bp | 800 bp | 3907, 3919 | 1569, 3535, 3582, 5122, 7803 |
| 115,6 | 160 bp | 1070 bp | 3908, 3911, 3918, 3933, 3936 | 1318, 1558, 1573, 5410, 5509, 5521 |
| 12 | 175 bp | 800 bp | 3907, 3915, 3929 | 1569, 3535 |
| 135,6,10 | 80, 215 bp | 1070 bp | 3909, 3923 |  |
| 146 | 80 bp | 1070 bp | 3910, 3916, 3917, 3920 | 070201-0707, 0709, 0711, 0712, 0714-0726, 0728-0731, 0733-0735, 0737, 0739-0749N, 0751-0760, 151101-151103, 1576, 9501, 3576, 4061, 4201, 4202, 4204-4209, 4506, 5401-5411, 5413, 550101-5505, 5507-5517, 5519- 5528, 5601-5606, 5609-5613, 5615-5620, 670101-670102, 8101-8103, 8201, 8202, 8301 |
| 15 | 190 bp | 800 bp | 3912, 3931 |  |
| 1611 | 165, 225 bp | 800 bp | 3914, 3916, 3929 |  |
| 17 | 180 bp | 800 bp | 3925N |  |
| 18 | 160 bp | 800 bp | 390201-390202, 3908, 391301-391302, 3922, 3923 | 0817, 1503, 1547, 1549, 1554, 1561, 1562, 1568, 1569, 1574, 1591, 1598, 9503, 9523, 9527, 9531, 9532, 1812, 370101-3707, 3709-3713, 3803, 3804, 4012, 480101-4805, 4807-4815, 4817-4819 |
| 1912 | 385 bp | 1070 bp | 39010101-39010102L, 390103-390104, 3903-3907, 3909-3912, 3914-3920, 3924-3942 | 070201-0712, 0714-0753, 0755-0760, 080101-0809, 0811-0816, 0818-0834, 1401-1407N, 380101-380202, 3805-3817, 4025?, 4068?, 4201, 4202, 4204-4209, 4806, 670101-670102, 7301, 8102 |
| 206 | 100 bp | 800 bp | 3918, 3927 | 1569 |
| 21 | 220 bp | 800 bp | 39010101-39010102L, 390103-390202, 3904, 3905, 3907-391302, 3915, 3917-3920, 3922, 3923, 3925N-3927, 3931, 3935, 3938Q-3941 | 1569, 380101-3807, 3808weakly, 3809, 3811-3817, 670101-6702 |
| 22 | 225 bp | 800 bp | 3928 |  |
| 236,7 | 85 bp | 800 bp | 39010101-39010102L, 390103, 390201, 3903, 3904, 390601-390602, 3909, 3912, 3914, 3915, 3918, 3922, 3924-3936, 3938Q-3943 | 1401-1407N, 670101 |
| 2412 | 185 bp | 1070 bp | 39010101-39010102L, 390103-390201, 3903, 3905-390602, 3907?, 3908, 3909, 3911, 3914, 3915, 3918, 3922, 3924-3943 | 1401, 1407N, 380101-3803, 3807- 3817, 670101 |
| 25 | 230 bp | 1070 bp | 39010102L |  |
| 26 | 265 bp | 800 bp |  | 1401-1404, 1407N |
| 276,13 | 80, 185 bp | 800 bp | 3932, 3938Q | 1315, 511301-511302, 5119 |
| 28 | 180 bp | 1070 bp | 3932, 3933, 3940N | 1303, 1558, 1573, 9533, 9537, 3503, 3506, 3513, 3536, 3538, 3555, 3556, 3559, 3570, 3574, 3575, 4071, 4410, 4611, 511301-511302, 5115, 5304, 5403, 5601, 5602, 5604, 5607, 5608, 5611, 5613, 5614, 5616, 5617, 5619N, 5620, 8201, 8202, Cw*1515 |
| 29 | 245 bp | 800 bp | 390601-390602, 3934 | 0809, 1542, 1583, 3560, 4101, 4105-4107, 4204, 5121, 5136, 5401, 5402, 5404, 5405N, 5407, 5408N, 5410-5413, 550101-5503, 5505, 5507, 5510-5513, 5515-5521, 5523, 5525-5528, 5901, 5902 |
| 306 | 60 bp | 800 bp | 3918, 3935 | 1569 |
| 3112 | 150 bp | 800 bp | 3930, 3934, 3943 | 0748, 0813, 0825, 1301-1304, 1306-1315, 1317, 1319, 1321, 1505, 1520, 1531, 1552, 1555, 1586, 1591, 9507, 9514, 9523, 9524, 180101-1812, 1814, 1817N, 1818, 1820, 1822-1827, 2701-2703, 270502-270510, 2707-2709, 2711-2714, 2716, 2717, 2719, 2723, 2726-2738, 350101-3507, 350901-3510, 3512, 3513, 3515-3517, 3519-352002, 3522-3534, 3536, 3537, 3539-3542, 3546-3557, 3559, 3560, 3563-3565Q, 3568-3572, 3574-3578, 3582, 3583, 3702, 400101-4004, 40060101-401403, 4018-4022N, 4024, 4025, 4027, 4029-4031, 4033-4040, 4042-4050, 4052-4066, 4068-4078, 440301-440303, 4407, 4410, 4413, 4426, 4430-4432, 4436-4440, 4443, 4445-4447, 4450, 4454, 4457, 4606, 4611, 47010101-4705, 480101-4804, 4806, 4807, 4809-4814, 4816-4818, 510901-5110, 5115, 5131, 5140, 5142, 5203, 5210, 530101-5305, 5307, 5309, 5310, 5312, 5313, 5401-5405N, 5407-5409, 5411-5413, 550201-550202, 5504, 5507, 5508, 5510, 5512, 5513, 5516, 5518, 5519, 5522, 5523, 5526, 5527, 5601, 5602, 5604, 5607-5612, 5614-5620, 570101-570103, 570301-570302, 5706-5708, 5710, 5711, 5714, 5801, 5802, 5804, 5807, 5809-5815, 5901, 5902, 7301, 7804, 8101-8103 |
| 32 | 230 bp | 800 bp | 3924, 3941, 3942 | 3535 |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-B*39 SSP subtypings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 bp or more. Size differences shorter than 20 bp are not given. For high resolution SSP kits the length of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control band may sometimes be observed. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most tubes, or a band of 800 base pairs, for some tubes.

Tube number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-B*39 subtyping.

In addition, tubes number 5, 6, 9, 10, 12, 15 to 18, 20 to 23, 26, 27, 29 to 32 contain the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3For many HLA-B alleles complete 4th exon nucleotide sequences are not available. In these instances it is not known whether some of the primers of the SSP sets are completely matched with the target sequences or not. We assume that unknown sequences in the 4th exon are conserved within allelic groups.

The B*3910 and B*670102 alleles will give rise to identical amplification patterns with the HLA-B*39 subtyping kit. These two alleles can be distinguished by the HLA-B low resolution and/or HLA-B*67 kits.

4Due to the sharing of sequence motifs between HLA-B alleles, non-HLA-B*39 alleles will be amplified by primer mixes 1 to 3, 5 to 8, 10 to 12, 14, 18 to 21, 23, 24, 26 to 32. In addition, the Cw*1515 allele will be amplified by primer mix 28.

5The primer pairs in primer mixes 1, 2, 6, 9, 10, 11 and 13 may give rise to a primer oligomer artifact.

6Short specific PCR fragments are less intense and not as sharp as longer specific bands.

7Primer solutions 6 and 23 may yield less specific PCR product than the other HLA-B*39 primer mixes.

8Primer solution 7: Specific PCR fragment of 165 bp in the B*3917 allele. Specific PCR fragment of 220 bp in the B*3903, B*3914, B*3924, B*3929 and B*3937 alleles.

9Primer solution 10: Specific PCR fragment of 105 bp in the B*3919, B*5122 and B*7803 alleles. Specific PCR fragment of 165 bp in the B*3907, B*1569, B*3535 and B*3582 alleles.

10Primer solution 13: Specific PCR fragment of 80 bp in the B*3923 allele. Specific PCR fragment of 215 bp in the B*3909 allele.

11Primer solution 16: Specific PCR fragment of 165 bp in the B*3914 and B*3929 alleles. Specific PCR fragment of 225 bp in the B*3916 allele.

12Primer mixes 19, 24 and 31 have a tendency of giving rise to non-specific amplifications.

13Primer mix 27: Specific PCR fragment of 80 bp in the B*3932 and B*5119 alleles. Specific PCR fragment of 185 bp in the B*3938Q  and B*1315 and B*511301-511302 alleles.

‘?’, nucleotide sequence information not available for the primer matching sequence.

1The nucleotide sequence of the B*390102 allele has been shown to be identical B*390101.

The nucleotide sequence of the B*3921 allele has been shown to be identical B*3924.

2The B*3910 and B*670102 alleles will give rise to identical amplification patterns with the HLA-B*39 subtyping kit. These two alleles can be distinguished by the HLA-B low resolution and/or HLA-B*67 kits.

3Primer solution 7: Specific PCR fragment of 165 bp in the B*3917 allele. Specific PCR fragment of 220 bp in the B*3903, B*3914, B*3924, B*3929 and B*3937 alleles.

Primer solution 10: Specific PCR fragment of 105 bp in the B*3919, B*5122 and B*7803 alleles. Specific PCR fragment of 165 bp in the B*3907, B*1569, B*3535 and B*3582 alleles.

Primer solution 13: Specific PCR fragment of 80 bp in the B*3923 allele. Specific PCR fragment of 215 bp in the B*3909 allele.

Primer solution 16: Specific PCR fragment of 165 bp in the B*3914 and B*3929 alleles. Specific PCR fragment of 225 bp in the B*3916 allele.

Primer mix 27: Specific PCR fragment of 80 bp in the B*3932 and B*5119 alleles. Specific PCR fragment of 185 bp in the B*3938Q  and B*1315 and B*511301-511302 alleles.

‘w’, might be weakly amplified.

‘?’, nucleotide sequence information not available for the primer matching sequence.