Olerup SSP®

Olerup SSP® HLA-B*39

Prod. No:	101.566-12/12u

101.566-04/04u

Lot No:    	67G

Expiry Date:    	2012-February-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

PHOTO DOCUMENT

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 32 primer mixes used for HLA-B*39 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified HLA-B*39 alleles3 | Other amplified HLA-Class I alleles4 |
| --- | --- | --- | --- | --- |
| 1 | 230 bp | 800 bp | *39010101, 39010102L, 390103-3905, 3907-3920, 3922-3929, 3931, 3935, 3937-3941, 3944-3946, 3948, 3949, 3951-3954 | *1569, 380101-3807, 3809, 3811-3814, 3816-3819, 5820, 670101-6702 |
| 2 | 170 bp | 1070 bp | *39010101, 39010102L, 390103-390110, 3903-3907, 3909, 3911, 3912, 3914, 3915, 3918-391902, 3922, 3924-3938Q, 3940N-3948, 3950, 3951, 3953 | *0769, 0785, 1401-1416, 1509, 1510, 151801-151803, 1523, 1537, 1551, 1552, 1572, 1580, 1590, 1593, 1599, 9508, 9514, 9515, 9519, 9524, 9533, 9534, 9553, 9561, 9576, 1833, 380101-380202, 3804, 3805, 3808-3815, 3817-3820 |
| 3 | 200 bp | 800 bp | *3905, 3907, 3908, 3911, 391301, 391302, 3920, 3937, 3943, 3949 | *0716, 0727, 0737, 0738, 0750, 0775, 0791, 0803, 0823, 1414, 150301-150303, 1509, 1510, 151801-151803, 1523, 1529, 1537-153802, 1547, 1549, 1551, 1552, 1554, 1562, 1564, 1568, 1569, 1572, 1574, 1580, 1590, 1591, 1593, 1598, 1599, 9503, 9508, 9514, 9515, 9519, 9523, 9524, 9527, 9531-9534, 9551, 9553, 9556, 9558, 9561, 9573, 9576, 180101-1815, 1817N-1845, 3521, 3524, 3525, 370101-3719, 380101-380202, 3804-3812, 3815, 3817-3820, 4012, 4211, 480101-4820, 5137, 5145, 5163, 5216, 5302, 5306, 5714, 5809, 5812 |
| 45,9 | 145, 385 bp | 1070 bp | *3926, 3951 |  |
| 55 | 105 bp | 800 bp | *39010101, 39010102L, 390103-390108, 390110, 390201, 3903-3909, 3911, 3912, 3914, 3915, 3918, 391902, 3922, 3924-3948, 3950-3954 | *1401-140202, 1403-1416, 3526, 670101 |
| 6 | 170 bp | 800 bp | *390201, 390202, 3908, 391301, 391302, 3923, 3939, 3949 | *0727, 0750, 0804, 0817, 150301-150303, 1547, 1554, 1561, 1562, 1564, 1568, 1569, 1591, 1598, 9503, 9523, 9527, 9531, 9532, 9551, 9556, 9558, 9573, 180101-1803, 1805, 1806, 1808-1815, 1817N-1828, 1830-1832, 1834-1845, 370101-3719, 3803, 4012, 4211, 480101-4805, 4807-4820, 5216 |
| 76,10 | 165, 220 bp | 1070 bp | *3903, 3914, 3917, 3924, 3929, 3937 | *3819 |
| 811 | 180, 405 bp | 1070 bp | *3904, 3946, 3947, 3950 | *1411, 3526, 4036, 4077, 4086, 4807 |
| 96 | 230 bp | 1070 bp | *390601, 390602 |  |
| 105,12 | 105, 165 bp | 800 bp | *3907, 391901, 391902 | *1569, 1833, 3535, 3582, 3585, 5122, 7803 |
| 118 | 170 bp | 1070 bp | *3908, 3911, 3918, 3933, 3936, 3940N, 3953 | *0778, 1318, 1558, 1573, 5410, 5509, 5521, 5537, 5619N |
| 12 | 175 bp | 800 bp | *3907, 3915, 3929, 3949 | *1569, 3535 |
| 135,6,13 | 85, 385 bp | 800 bp | *3923, 3927, 3944, 3951 |  |
| 145 | 80 bp | 1070 bp | *3910, 3916, 3917, 3920 | *070201-0707, 0709, 0711, 0712, 0714-0726, 0728-0731, 0733-0735, 0737, 0739-0749N, 0751-0768, 0770-0780, 0782, 0784, 0786-0799, 151101-151104, 1576, 9501, 9543, 9548, 9552, 3576, 4061, 420101-4202, 4204-4210, 4212, 4213, 4506, 5401-5411, 5413-5419, 550101-5505, 5507-5517, 5519-5533, 5535-5538, 560101-5606, 5609-5613, 5615-5620, 5622-5626, 670101, 670102, 8101-8104N, 8201, 8202, 8301 |
| 155,7,14 | 145, 190 bp | 800 bp | *3912, 3931, 3938Q, 3954 |  |
| 168,12,15 | 165, 220 bp | 800 bp | *3909, 3914, 3916, 3929 |  |
| 17 | 180 bp | 800 bp | *3925N |  |
| 18 | 160 bp | 800 bp | *390201, 390202, 3908, 391301, 391302, 3922, 3923, 3949 | *0817, 0838, 150301-150303, 1547, 1549, 1554, 1561, 1562, 1568, 1569, 1574, 1591, 1598, 9503, 9523, 9527, 9531, 9532, 9551, 9556, 9558, 9573, 1812, 370101-3707, 3709-3719, 3803, 3804, 4012, 4211, 480101-4805, 4807-4815, 4817-4820, 5216 |
| 19 | 385 bp | 1070 bp | *39010101, 39010102L, 390103-390110, 3903-3907, 3909-3912, 3914-3920, 3924-3948, 3950-3954 | *070201-0712, 0714-0753, 0755-0764, 0766-0799, 080101-0809, 0811-0816, 0818-0837, 0839-0847, 1401-1403, 1405-1416, 380101-380202, 3805-3820, 4025?, 4068?, 420101-4202, 4204-4210, 4212, 4213, 4806, 670101, 670102, 7301, 8102 |
| 205,7,16 | 100, 215 bp | 800 bp | *390105, 3909, 3918, 3927 | *1569 |
| 21 | 220 bp | 800 bp | *39010101, 39010102L, 390103-390202, 3904, 3905, 3907-391302, 3915, 3917-3920, 3922, 3923, 3925N-3927, 3931, 3935, 3938Q-3941, 3942w, 3944-3946, 3948, 3949, 3951-3954 | *1569, 380101-3807, 3808w, 3809, 3811-3818, 5820, 670101-6702 |
| 225,6,17 | 150, 225 bp | 800 bp | *3928, 3945 |  |
| 235,7 | 85 bp | 800 bp | *39010101, 39010102L, 390103, 390105-390108, 390110, 390201, 3903, 3904, 390601, 390602, 3909, 3912, 3914, 3915, 3918, 391902, 3922, 3924-3936, 3938Q-3948, 3950-3954 | *1401-140202, 1403-1413, 1415, 1416, 3814w, 670101 |
| 247 | 185 bp | 1070 bp | *39010101, 39010102L, 390103-390108, 390110, 390201, 3903, 3905-3909, 3911, 3914, 3915, 3918, 391902, 3922, 3924-3930, 3931w, 3932-3944, 3946-3948, 3950-3954 | *1401, 1407N, 1408, 1410, 1412, 1414, 380101-3803, 3807-3820, 670101 |
| 25 | 235 bp | 1070 bp | *39010102L |  |
| 268 | 265 bp | 800 bp |  | *1401-1404, 1407N, 1409, 1411, 1412, 1414-1416, 1844, Cw*0130 |
| 275,18 | 70, 195 bp | 800 bp | *3932, 3938Q, 3948 | *1315, 3596, 511301, 511302, 5119, 5162, 5214 |
| 288 | 180 bp | 1070 bp | *3932, 3933, 3940N | *1303, 1558, 1573, 9533, 350301, 350302, 3506, 3513, 3536, 3538, 3555, 3556, 3559, 3570, 3574, 3575, 3584, 3585, 3598, 4071, 4410, 4611, 511301, 511302, 5115, 5162, 5214, 5304, 5403, 560101-5602, 5604, 5607, 5608, 5611, 5613, 5614, 5616, 5617, 5619N, 5620, 5622, 5624-5626, 5904, 8201, 8202, Cw*1515 |
| 29 | 245 bp | 800 bp | *390601, 390602, 3934, 3950 | *0778, 0784, 0809, 130201-130206, 1308Q, 1309, 1314-1316, 1318, 1319, 1327, 1330, 1542, 1583, 2714, 3560, 40060101-400602, 4044, 4053, 4070, 4075, 4083, 4086, 4093, 4095, 4096, 4101, 4105-4107, 4109, 4112, 4204, 4508, 4510, 4618, 4907, 5110, 5116, 5121, 5131, 5134, 5136, 5182, 5208, 5401, 5402, 5404, 5405N, 5407, 5408N, 5410-5419, 550101-5503, 5505, 5507, 5509-5513, 5515-5538, 5623, 5901-5903, 5905, 7301 |
| 305,8,19 | 60, 180 bp | 800 bp | *3918, 3935, 3952 | *1569 |
| 318 | 155 bp | 800 bp | *3930, 3934, 3943 | *0748, 0813, 0825, 130101-1304, 1306-1315, 1317, 1319, 1321-1323, 1325-1330, 1505, 1520, 1531, 1552, 1555, 1586, 1591, 9507, 9514, 9523, 9524, 9551, 9555, 180101-1812, 1814, 1817N, 1818, 1820, 1822-1829, 1831-1845, 2701-2703, 270502-270512, 2707-2709, 2711-2714, 2716, 2717, 2719, 2723, 2726-2739, 2741-2744, 2746, 2748-2753, 2755-2759N, 350101-3507, 350901-3510, 351201-3513, 3515-3517, 3519-352002, 3522-3534, 3536, 3537, 3539-3542, 3546-3557, 3559, 3560, 3563-3565Q, 356801-3572, 3574-3578, 3581-3585, 3588-3595, 3597, 3598, 3702, 400101-4004, 40060101-401403, 4018-4022N, 4024, 4025, 4027, 4029-4031, 4033-4040, 4042-4050, 4052-4094, 4096, 4097, 440301-440308, 4407, 4410, 4413, 4426, 4430-4432, 4436-4440, 4443, 4445-4447, 4450, 4454, 4457, 4460, 4461N, 446401-4465, 4469, 4477, 4481, 4482, 4485, 4606, 4611, 461301, 461302, 4618, 4621, 47010101-4705, 480101-4804, 4806, 4807, 4809-4814, 4816-4818, 4820, 510901-5110, 5115, 5131, 5140, 5142, 5173, 5181, 5203, 5210, 530101-5305, 5307, 5309, 5310, 5312-5315, 5317-5319, 5401-5405N, 5407-5409, 5411-5419, 550201-550205, 5504, 5507, 5508, 5510, 5512, 5513, 5516, 5518, 5519, 5522, 5523, 5526, 5527, 5530, 5532, 5534, 5535, 560101-5602, 5604, 5607-5612, 5614-5620, 5622-5624, 5626, 570101-570105, 570301, 570302, 5706-5708, 5710, 5711, 5714-5718, 5720, 5721, 5723, 5725, 5726, 580101-5802, 5804, 5807, 5809-5818, 5821, 5823, 5824, 5901-5905, 7804, 8101-8104N |
| 32 | 230 bp | 800 bp | *3903w, 3914w, 3924, 3929w, 3937w, 3941, 3942 | *3535, 3819w, 3820 |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-B*39 SSP subtypings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 bp or more. Size differences shorter than 20 bp are not given. For high resolution SSP kits the length of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control band may sometimes be observed. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most wells, or a band of 800 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-B*39 subtyping.

In addition, wells number 3, 5, 6, 10, 12, 13, 15 to 18, 20 to 23, 26, 27, 29 to 32 contain the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3For many HLA-B alleles complete 4th exon nucleotide sequences are not available. In these instances it is not known whether some of the primers of the SSP sets are completely matched with the target sequences or not. We assume that unknown sequences in the 4th exon are conserved within allelic groups.

The B*3910 and B*670102 alleles will give rise to identical amplification patterns with the HLA-B*39 subtyping kit. These two alleles can be distinguished by the HLA-B low resolution and/or HLA-B*67 kits.

## 4Due to the sharing of sequence motifs between HLA-B alleles, non-HLA-B*39 alleles will be amplified by primer mixes 1 to 3, 5, 6,  8, 10 to 12, 14, 18 to 21, 23, 24, 26 to 32. In addition, the Cw*0130 allel will be amplified by primer mix 26 and the Cw*1515 allele will be amplified by primer mix 28.

5Short specific PCR fragments are less intense and not as sharp as longer specific bands.

6The primer pairs in primer mix 7, 9, 13 and 22 may give rise to a primer oligomer artifact.

7Primer solutions 15, 20, 23 and 24 may yield less specific PCR product than the other HLA-B*39 primer mixes, in primer mix 15 notably for the B*3931 allele.

8Primer mixes 11, 16, 26, 28, 30 and 31 have a tendency of giving rise to non-specific amplifications.

9Primer mix 4: Specific PCR fragment of 145 bp in the B*3926 allele. Specific PCR fragment of 385 bp in the B*3951 allele.

10Primer solution 7: Specific PCR fragment of 165 bp in the B*3917 allele. Specific PCR fragment of 220 bp in the B*3903, B*3914, B*3924, B*3929 and B*3937 and in the B*3817 alleles.

11Primer solution 8: Specific PCR fragment of 180 bp in the B*3904, B*3947 and B*3950 and in the B*1411, B*3526, B*4036, B*4077, B*4086 and B*4807alleles. Specific PCR fragment of 405 bp in the B*3946 allele.

12Primer solution 10: Specific PCR fragment of 105 bp in the B*391901-391902 and in the B*1833, B*3582, B*3585, B*5122 and B*7803 alleles. Specific PCR fragment of 165 bp in the B*3907 and in the B*1569 and B*3535, alleles.

13Primer mix 13: Specific PCR fragment of 85 bp in the B*3923, B*3927 and B*3944 alleles.  Specific PCR fragment of 385 bp in the B*3951 allele.

14Primer mix 15: Specific PCR fragment of 145 bp in the B*3954 allele. Specific PCR fragments of 190 bp in the B*3912, B*3931 and B*3938Q alleles.

15Primer solution 16: Specific PCR fragment of 165 bp in the B*3914 and B*3929 alleles. Specific PCR fragment of 220 bp in the B*3909 and B*3916 alleles.

16Primer solution 20: Specific PCR fragment of 100 bp in the B*390105, B*3918 and B*3927 and in the B*1569 alleles. Specific PCR fragment of 215 bp in the B*3909 allele.

17Primer solution 22: Specific PCR fragment of 150 bp in the B*3945 allele. Specific PCR fragment of 225 bp in the B*3928 allele.

18Primer mix 27: Specific PCR fragment of 70 bp in the B*5119 alleles. Specific PCR fragment of 195 bp in the B*1315, B*3596, B*511301-511302, B*5162 and B*5214 alleles. Specific PCR fragments of 70 and 195 bp in the B*3932, B*3938Q and B*3948 alleles (please observe that both specific PCR fragments may not be obtained).

19Primer mix 30: Specific PCR fragment of 60 bp in the B*3918 and B*3935 and in the B*1569 alleles. Specific PCR fragment of 180 bp in the B*3952 allele.

‘w’, might be weakly amplified.

‘?’, nucleotide sequence information not available for the primer matching sequence.

1The nucleotide sequence of the B*390102 allele has been shown to be identical B*390101.

The nucleotide sequence of the B*3921 allele has been shown to be identical B*3924.

2The B*3910 and B*670102 alleles will give rise to identical amplification patterns with the HLA-B*39 subtyping kit. These two alleles can be distinguished by the HLA-B low resolution and/or HLA-B*67 kits.

3Primer mix 4: Specific PCR fragment of 145 bp in the B*3926 allele. Specific PCR fragment of 385 bp in the B*3951 allele.

Primer solution 7: Specific PCR fragment of 165 bp in the B*3917 allele. Specific PCR fragment of 220 bp in the B*3903, B*3914, B*3924, B*3929 and B*3937 and in the B*3817 alleles.

Primer solution 8: Specific PCR fragment of 180 bp in the B*3904, B*3947 and B*3950 and in the B*1411, B*3526, B*4036, B*4077, B*4086 and B*4807alleles. Specific PCR fragment of 405 bp in the B*3946 allele.

Primer solution 10: Specific PCR fragment of 105 bp in the B*391901-391902 and in the B*1833, B*3582, B*3585, B*5122 and B*7803 alleles. Specific PCR fragment of 165 bp in the B*3907 and in the B*1569 and B*3535, alleles.

Primer mix 13: Specific PCR fragment of 85 bp in the B*3923, B*3927 and B*3944 alleles.  Specific PCR fragment of 385 bp in the B*3951 allele.

Primer mix 15: Specific PCR fragment of 145 bp in the B*3954 allele. Specific PCR fragments of 190 bp in the B*3912, B*3931 and B*3938Q alleles.

Primer solution 16: Specific PCR fragment of 165 bp in the B*3914 and B*3929 alleles. Specific PCR fragment of 220 bp in the B*3909 and B*3916 alleles.

Primer solution 20: Specific PCR fragment of 100 bp in the B*390105, B*3918 and B*3927 and in the B*1569 alleles. Specific PCR fragment of 215 bp in the B*3909 allele.

Primer solution 22: Specific PCR fragment of 150 bp in the B*3945 allele. Specific PCR fragment of 225 bp in the B*3928 allele.

Primer mix 27: Specific PCR fragment of 70 bp in the B*5119 alleles. Specific PCR fragment of 195 bp in the B*1315, B*3596, B*511301-511302, B*5162 and B*5214 alleles. Specific PCR fragments of 70 and 195 bp in the B*3932, B*3938Q and B*3948 alleles (please observe that both specific PCR fragments may not be obtained).

Primer mix 30: Specific PCR fragment of 60 bp in the B*3918 and B*3935 and in the B*1569 alleles. Specific PCR fragment of 180 bp in the B*3952 allele.

‘w’, might be weakly amplified.

‘?’, nucleotide sequence information not available for the primer matching sequence.

Changes in revision R01 compared to R00:

- The B*3903, 3914, 3929 and 3937 and the B*3819 alleles are weakly amplified by primer mix 32. This has been corrected in the Specificity and Interpretation Tables.

Change in revision R02 compared to R01:

- The B*3931 allele is weakly amplified by primer mix 24.

- Primer mix 15 gives a low yield of specific product for the B*3931 allele.