Gel Documentation Form and Worksheet

HLA-B*39 (101.566-12/04 -12u/04u)	Lot No: 6D6	Expiry Date: 2018-09-01

Name:__________________

Sample ID:______________

DNA Conc.(ng/ul):_________

Test  Date:_______________

Tested By:_______________

Review Date:_______________

Reviewed By:_______________

Interpretation:__________	Failed lanes: ___________ 	   Comments:_______________

Gel Picture

PHOTO DOCUMENT

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 3, 22 and 23 have a tendency to giving rise to primer oligomer formation.

Primer mixes 3, 7, 8, 11, 12, 15, 16, 20, 24, 26 and 30 may have tendencies of unspecific amplifications.

Primer mixes 25 and 28 may give rise to a long unspecific amplification product of approximately 600

bp. This should be disregarded when interpreting the B*39 typings.

Primer mix 24 may give rise to a lower yield of HLA-specific PCR product than the other B*39 primer mixes.

Primer mix 39 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 430 base pairs.

1HLA-B*39 alleles in bold lettering are listed as confirmed alleles on the on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.22.0, October 2015.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3Primer mix 4: Specific PCR fragment of 65 bp in the B*39:58 and the B*08:01:24 and 35:03:17 allele. Specific PCR fragment of 145 bp in the B*39:26 allele. Specific PCR fragment of 385 bp in the B*39:51 allele.

Primer mix 7: Specific PCR fragment of 120 bp in the B*39:68 and the B*67:03 and in the A*02:403, A*30:63, A*68:108, C*05:70, C*06:40 and C*07:137:01-07:137:02 alleles. Specific PCR fragment of 165 bp in the B*39:17 allele. Specific PCR fragment of 220 bp in the B*39:03, 39:14, 39:24:01-39:24:02, 39:29, 39:37, 39:76 and the B*38:19 alleles. Specific PCR fragment of 120 and 220 bp in the B*08:55 allele. Specific PCR fragment of 120 and 165 bp in the C*15:51 allele.

Primer mix 8: Specific PCR fragment of 180 bp in the B*39:04, 39:47 and 39:50 and the B*07:214, 08:01:24, 14:28, 14:46, 15:189, 35:03:17, 35:26, 38:33, 40:36, 40:77, 40:86, 40:314, 48:07 and 49:26 alleles. Specific PCR fragment of 405 bp in the B*39:46 allele.

Primer mix 9: Specific PCR fragment of 100 bp in the B*39:01:05 and 39:18 and the B*15:69, 15:186, 15:265 and 38:01:03 and in the C*04:188, C*06:40, C*14:63 and C*15:51 alleles. Specific PCR fragment of 230 bp in the B*39:06:01-39:06:05, 39:57, 39:62, 39:64, 39:83 and 39:90 and the B*35:276 and 51:101 alleles. Specific PCR fragment of 310 bp in the B*39:59 allele.

Primer mix 10: Specific PCR fragment of 105 bp in the B*39:19:01-39:19:02 and the B*35:82, 35:85, 35:135, 35:294, 37:57, 51:22 and 78:03 alleles. Specific PCR fragment of 165 bp in the B*39:07 and 39:91 and the B*15:69, 15:186, 15:265 and 35:35 and in the C*14:63 alleles. Specific PCR fragment of 210 bp in the B*39:09:01-39:09:03, 39:43 and 39:88 and the B*15:38:01-15:38:02, 15:185, 15:335, 15:364, 15:368, 18:01:01:01-18:09, 18:12:01-18:15, 18:17N, 18:19-18:20, 18:22-18:25, 18:27-18:32, 18:34, 18:36-18:40, 18:42-18:55, 18:57, 18:59-18:60, 18:62-18:100, 18:102-18:109, 18:111-18:116, 35:21, 35:24:01-35:24:02, 35:188, 35:190, 35:287, 51:37, 51:45, 51:63, 51:97, 52:39, 53:02, 53:06, 53:28, 56:31, 57:14:01-57:14:02 and 58:09 and in the C*12:87alleles. Specific PCR fragment of 105 bp and 210 bp in the B*18:33 allele. Specific PCR fragment of 210 bp and 255 bp in the B*18:18 allele.

Primer mix 11: Specific PCR fragment of 160 bp in the B*39:08, 39:11, 39:18, 39:33 and 39:36 and the B*07:78, 07:214, 07:246, 13:18, 13:31, 13:41, 13:73, 15:58, 15:73, 15:303, 46:61, 54:10, 54:20, 54:33, 55:09, 55:21, 55:37, 55:52 and 56:43 and in the C*06:34:02 and C*07:447 alleles. Specific PCR fragment of 185 bp in the B*39:40:01N-39:40:02N, 39:53 and 39:57 and the B*56:19N alleles.

Primer mix 12: Specific PCR fragment of 175 bp in the B*39:07, 39:15, 39:29, 39:49, 39:55 and 39:91 and the B*15:69, 15:186, 35:35 and 38:29 and in the C*15:51 alleles.  Specific PCR fragment of 205 bp in the B*39:61 and 39:99 and the 40:283? allele.

Primer mix 13: Specific PCR fragment of 80 bp in the B*39:77 allele.  Specific PCR fragment of 230 bp in the B*39:24:01-39:24:02, 39:28, 39:41-39:42 and 39:76 and the B*15:265, 35:35, 37:40 and 38:20 and in the C*15:51 alleles.

Primer mix 15: Specific PCR fragment of 145 bp in the B*39:54 allele. Specific PCR fragments of 190 bp in the B*39:12, 39:31 and 39:38Q and the B*08:01:24, 14:28 and 49:26 alleles.

Primer mix 16: Specific PCR fragment of 165 bp in the B*39:14, 39:29 and 39:70 and the B*08:55 and 51:101 alleles. Specific PCR fragment of 220 bp in the B*39:09:01-39:09:03, 39:16 and 39:88 and in the C*07:137:01-07:137:02 alleles.

Primer mix 17: Specific PCR fragment of 80 bp in the B*39:23 and 39:44 alleles. Specific PCR fragment of 180 bp in the B*39:25N allele.  Specific PCR fragment of 205 bp in the B*39:61 allele.

Primer solution 25: Specific PCR fragment of 150 bp in the B*39:60 and 39:82 and in the C*08:83 and C*15:31 alleles. Specific PCR fragment of 235 bp in the B*39:01:01:02L allele.

Primer mix 27: Specific PCR fragment of 180 bp in the B*39:38Q and 39:63 alleles. Specific PCR fragment of 205 bp in the B*39:32 and 39:48 and the B*13:15, 35:96, 35:109, 51:13:01-51:13:02, 51:62, 51:92:01-51:92:02, 51:106, 52:14, 52:25 and 56:34 and in the C*15:39 alleles.

Primer mix 28: Specific PCR fragment of 140 bp in the B*39:56 allele. Specific PCR fragment of 180 bp in the B*39:32-39:33, 39:40:01N-39:40:02N and 39:74 and the B*13:03, 13:48, 15:58, 15:73, 15:133, 15:229, 15:253, 15:303, 35:03:01-35:03:05, 35:03:07-35:03:21, 35:06, 35:13, 35:36, 35:38, 35:55-35:56, 35:59:01-35:59:02, 35:70, 35:74-35:75, 35:84-35:85, 35:98, 35:106, 35:109, 35:127-35:128, 35:136, 35:150:01-35:153, 35:155-35:156, 35:160, 35:163, 35:167, 35:169, 35:179, 35:181, 35:193, 35:195, 35:198, 35:204-35:205, 35:221, 35:223, 35:231, 35:235-35:237, 35:242-35:243, 35:246, 35:256-35:257, 35:267, 35:274, 35:278-35:279, 35:281-35:282, 35:290-35:291, 35:296, 35:298, 38:30, 40:71, 44:10, 44:226, 45:14, 46:11, 46:61, 50:39-50:40, 51:13:01-51:13:02, 51:15, 51:62, 51:92:01-51:92:02, 51:106, 52:14, 52:25, 53:04, 53:31, 53:33, 54:03, 54:33, 56:01:01:01-56:01:04, 56:01:06-56:02, 56:04, 56:07-56:08, 56:11, 56:13-56:14, 56:16-56:17, 56:19N-56:20:02, 56:22, 56:24-56:30, 56:33-56:49, 57:76, 58:34, 58:73, 59:04 and 82:01-82:03 and in the C*15:15 alleles.

Primer mix 30: Specific PCR fragment of 65 bp in the B*39:18, 39:35 and 39:64 and the B*15:69, 15:265, 38:32 and 51:101 and in the C*04:188, C*05:70, C*06:40, C*14:63 and C*15:51 alleles. Specific PCR fragment of 135 bp in the B*39:75 and in the C*07:175 alleles. Specific PCR fragment of 190 bp in the B*39:52 and 39:62 and the B*38:54 and in the A*02:403, A*30:63 and A*68:108 alleles.

Primer mix 32: Specific PCR fragment of 190 bp in the B*39:89 allele. Specific PCR fragment of 225 bp in the B*39:79 and in the C*06:40 and C*14:63 alleles.

Primer mix 33: Specific PCR fragment of 120 bp in the B*39:93 and the B*38:58 alleles. Specific PCR fragment of 230 bp in the B*39:87N and the B*15:226N and in the C*01:56N alleles.

4The B*39:10:01 and 39:96 and the 67:01:02-67:01:03 and 67:05 alleles will give rise to identical amplification patterns with the HLA-B*39 subtyping kit. These two alleles can be distinguished by the HLA-B low resolution and/or HLA-B*67 kits.

The B*39:47 and the B*14:46 alleles will give rise to identical amplification patterns with the HLA-B*39 subtyping kit. These two alleles can be distinguished by the HLA-B low resolution kit.

The B*39:104 and the B*38:41 alleles will give rise to identical amplification patterns with the HLA-B*39 subtyping kit. These two alleles can be distinguished by the HLA-B low resolution and/or HLA-B*38 kits.

The B*39:106 and the B*14:08:01-14:08:02 and 14:10 alleles will give rise to identical amplification patterns with the HLA-B*39 subtyping kit. These two alleles can be distinguished by the HLA-B low resolution and/or HLA-B*14 kits.

5The following HLA-B*39 alleles can be distinguished by the different sizes of the HLA-specific PCR product:

| Alleles | Primer mix | Alleles | Primer mix |
| --- | --- | --- | --- |
| B*39:01:01:02L, 39:60 | 25 | B*39:35, 39:75 | 30 |
| B*39:01:05, 39:59 | 9 | B*39:41, 39:77 | 13 |
| B*39:25N, 39:44 | 17 | B*39:62, 39:64 | 30 |
| B*39:26, 39:51 | 4 | B*39:87N, 39:93 | 33 |

The HLA-B*39 subtyping kit cannot distinguish the silent mutations in the B*39:01:01:01, 39:01:01:03-39:01:01:04, 39:01:03, 39:01:06-39:01:08 and 39:01:10-39:01:22 alleles, the B*39:05:01-39:05:02 alleles, the B*39:06:01-39:06:02 and 39:06:04-39:06:05 alleles, the B*39:13:01-39:13:02 alleles, the B*39:24:01-39:24:02 alleles or the B*39:40:01N-39:40:02N alleles.

‘w’, might be weakly amplified.

‘?’, nucleotide sequence information not available for the primer matching sequence.

Change in revision R01 compared to R00:

- In primer mix 7 the C*15:51 alleles gives rise to PCR products of 120 and 165 bp. This has been corrected in the Specificity Table and in the footnotes to the Interpretation Tables.