PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® HLA-B*40

Prod. No:	101.523-24/24u

101.523-03/03u

Lot No:    	22F

Expiry Date:    	2010-November-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 72 primer mixes used for HLA-B*40 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified HLA-B*40 alleles3 | Other amplified HLA class I alleles4 |
| --- | --- | --- | --- | --- |
| 1 | 175 bp | 800 bp | 400101-400106, 4007, 4010, 4012, 4021-4023, 4025, 4033, 4036, 4038, 4042, 4043, 4046-4049, 4051, 4052, 4054, 4055, 4060-4063, 4065-4067, 4069, 4073, 4074, 4076, 4079, 4081, 4084, 4087, 4088, 4092 | 9516, 9524, 4431, 4606, 480301-480302 |
| 2 | 215 bp | 1070 bp | 400101-400602, 4009, 4011, 401401-4016, 4018-4020, 4022N-4039, 4042-4048, 4050-4057, 4059, 4060, 4062-4067, 4069-4074, 4076-4093 | 1323, 1571, 180101-1803, 1805, 1806, 1808-1815, 1817N-1824, 1826-1828, 1830-1832, 2718, 2729, 3563, 370101-3714, 4101-4108, 4415, 4418, 4455, 4501-4505, 4507-4509, 490101-4905, 5001, 5002, 5004 |
| 35 | 130 bp | 1070 bp | 400201-400602, 4008, 4009, 4011, 4013, 4018-4020, 4024, 4026-4029, 4035, 4037, 4039, 4040, 4044, 4050, 4056, 4057, 4064, 4068, 4070, 4071, 4074, 4075, 4078, 4082, 4085, 4086, 4089-4091 |  |
| 45 | 95 bp | 800 bp | 400201-400203, 4005, 4008, 4009, 4013, 4015, 4016, 4018, 4019, 4024, 4027, 4029, 4031, 4032, 4035, 4037, 4039, 4040, 4045, 4050, 4056-4058, 4078, 4080, 4082, 4085, 4089-4091 | 070201-0706, 0708, 0710, 0713, 0715, 0716, 0719-0768, 080101-0808N, 0810, 0811, 0813-0815, 0817-0827, 0829-0834, 0836, 0838, 1545, 2707, 2711, 2724, 2732-2734, 2743, 3522, 3531, 3587, 3707, 3709, 3914, 3929, 4102, 4104, 4201, 4202, 420501-420502, 4207, 4208, 4454, 480101-480102, 4804-4806, 4808-4813, 4815, 4816, 4818, 4819, 5504, 8101-8104N, Cw*0305, Cw*0325, Cw*0327 |
| 5 | 190 bp | 800 bp | 4003, 4018, 4020, 4024, 4031-4033, 4044, 4052, 4059 | 070201-0704, 0707-0713, 0715, 0716, 071801-071802, 0721-0723, 0725-0727, 0729, 0730, 0733, 0735-0739, 0741, 0742, 0744-0752, 0754-0759, 0761-0763, 0765-0768, 1540, 1547, 1549, 1552, 9514, 9517, 9524, 9538, 1810, 1821, 2734, 2741, 3515, 3533, 3566, 3586, 3707, 4606, 4808, 5713, Cw*0206 |
| 65 | 95 bp | 1070 bp | 4003, 4027 | 0709, 0711, 0717, 0828, 0835, 0837, 1405, 1507, 1555, 1568, 9526, 1814, 3505, 3551, 3558, 3566, 3572, 3589, 3903, 3924, 3937, 4206, 4209, 4612, 4807, 4814, 5314, 5818, Cw*0808 |
| 75,7 | 105 bp | 800 bp | 4012, 4046, 4093 | 0727, 0750, 0804, 0817, 1503, 1547, 1554, 1561, 1562, 1564, 1568, 1569, 1591, 1598, 9503, 9523, 9527, 9531, 9532, 9551, 2718, 2729, 3803, 390201-390202, 3908, 391301-391302, 3923, 3939, 4440, 4444, 480101-4805, 4807-4819, 5518 |
| 8 | 235 bp | 1070 bp | 4005, 4071 | 0720, 0724, 0760, 0821, 0825, 1507, 1545, 1555, 1568, 9526, 3505, 3516, 3517, 3522, 3530, 3551, 3558, 3572, 3589, 4454, 4612, 5314, 5818, Cw*0305, Cw*0325, Cw*0327 |
| 9 | 160 bp | 1070 bp | 40060101-400602, 4044, 4053, 4070, 4075, 4083, 4086, 4093 | 2714, 510109, 5110, 512402, 5411, 7301 |
| 10 | 215 bp | 1070 bp | 4007, 4008, 4013, 4068 | 1807, 1829, 2723, 3550, 3584 |
| 115 | 85 bp | 1070 bp | 4077, 4087 | 9524, 4814, 5805weakly |
| 125 | 95 bp | 1070 bp | 4008, 4025, 4068 | 9543, 3519, 3547, 4412 |
| 135 | 90 bp | 1070 bp | 4003, 4009, 4018, 4024, 4031, 4032 | 070201-0704, 0708-0711, 0713, 0715-0717, 0719-0731, 0733, 0735-0739, 0741-0752, 0754-0768, 0807, 0828, 0837, 1507, 1545, 1555, 1568, 1814, 1822, 2734, 3505, 3516, 3517, 3530, 3551, 3558, 3566, 3572, 3589, 3707, 3709, 3929, 4104, 4209, 4454, 4612, 4808, 5314, 5818, Cw*0305, Cw*0325, Cw*0327 |
| 145 | 90 bp | 800 bp | 4010 | 1546, 9506, 3547 |
| 15 | 150 bp | 800 bp | 4013, 4019, 4047 | 3706, 44020101-4405, 4407, 4410-4439, 4441-4443, 4445, 4447-4456N, 4458N, 4459, 4461N-446402, 47010101-47010102, 4703-4705, 490101-4902, 4904, 4905 |
| 165,10 | 75, 160 bp | 800 bp | 401401, 4015, 407201, 4077 | 3566, 420501-420502, 4812, Cw*0738 |
| 175,6 | 120 bp | 800 bp | 400101-400106, 4007, 4010, 4012, 401401-401403, 4021, 4022N, 4025, 4030, 4031, 4033, 4034, 4036, 4038, 4042, 4043, 4045-4049, 4052, 4054, 4055, 4059-4063, 4065, 4066, 4069, 407201-4074, 4076, 4077, 4079-4081, 4084, 4087, 4088 | 0725, 0748, 2736, 3566, 4108, 420501-420502, 4431, 4606, 480101-480102, 480301-480302, 4806, 4807, 4809-4812, 4814, 4816-4819, 8101-8104N |
| 185 | 105 bp | 800 bp | 400106, 4026 | 070503weakly, 150102, 1509, 5106, 5159 |
| 19 | 310 bp | 1070 bp | 4004, 4024, 4028, 4059, 4060, 4068, 4070 | 0712, 1301-1304, 1306-1309, 1311-1320, 1322, 1323, 1516, 1520, 1534, 1542, 1562, 1567, 1585, 1586, 1595, 9537, 1822, 2714, 350101-3511, 3513-3515, 3518-3521, 3523-3530, 3533-3538, 3540N-3542, 3545-3550, 3552-3558, 3560, 3562-3566, 356801-3572, 3574-3578, 3580-3585, 3588-3594, 4104, 4107, 4410, 4415, 4418, 4448, 4501-4509, 4611, 4618, 4802, 4817, 490101-4905, 5001, 5002, 5115, 5156, 530101-5306, 5308-5316, 5401-5405N, 5407, 5408N, 5410-5417, 550101-5503, 5505, 5507, 5509-5513, 5515, 5516, 5518-5534, 5601, 5602, 5607-5611, 5613-5617, 5619N, 5620, 5622-5624, 5705, 580101-5802, 5804-5807, 5809-5819, 5901-5903, 8201, 8202 |
| 205,11 | 135, 285 bp | 1070 bp | 401401-4016, 4023, 4031, 4032, 4045, 4080 | 070201-071802, 0720-0732, 0734-0739, 0741-0752, 0754-0759, 0761-0768, 9538, 3510, 3566, 3707, 3907, 4108, 420501-420502, 480101-480102, 4805-4812, 4814-4816, 4818, 4819, 8101-8104N |
| 215,6 | 145 bp | 800 bp | 4021 | 1309, 15010101-150104, 150106-1502, 1504-1508, 151101-1512, 1514, 1515, 1519-1521, 1525-1528, 1530-1535, 1538-1540, 1542, 1544, 1545, 1548, 1550, 1555-1558, 1560, 1563, 1565, 1566, 1570, 1571, 1573, 1575-1579N, 1581-1586, 1588, 1592, 1594N, 1596, 1597, 9501, 9502, 9504, 9505, 9507, 9509- 9513, 9516-9518, 9520-9522, 9525, 9526, 9528, 9529, 9535-9542, 9544-9550, , 3546, 460101-4603, 4605-4618, 5712 |
| 22 | 235 bp | 800 bp | 400201-4003, 4008, 4009, 4013, 4015, 4016, 4018, 4019, 4024, 4027, 4029, 4031, 4032, 4035, 4037, 4040, 4045, 4050, 4056-4058, 4078, 4080, 4082, 4085, 4089-4091 | 070201-0706, 0708-0711, 0713, 0715-0717, 0721-0723, 0725-0730, 0732, 0733, 0735-0742, 0744-0759, 0761-0763, 0765-0768, 2707, 2711, 2724, 2732-2734, 2743, 3531, 3566, 3707, 480101-480102, 4804-4816, 4818, 4819, 8101-8104N |
| 235 | 95 bp | 800 bp | 4029 |  |
| 24 | 165 bp | 1070 bp | 4034 |  |
| 255 | 125 bp | 1070 bp | 4043, 4050 | 4506 |
| 265 | 120 bp | 1070 bp | 4015, 4016, 4023, 4032 | 070201-071802, 0720-0724, 0726-0732, 0734-0739, 0741-0747, 0749N, 0750, 0752, 0754-0759, 0761-0768, 9538, 3707, 4805, 4808, 4815 |
| 27 | 180 bp | 800 bp | 4022N |  |
| 28 | 325 bp | 1070 bp | 400101-400106, 4004, 4007, 4010-4012, 401401-401403, 4020-4023, 4025, 4026, 4028, 4033, 4036, 4038, 4042, 4043, 4046-4049, 4051, 4052, 4054, 4055, 4059-4069, 407201-4074, 4076, 4077, 4079, 4081, 4084, 4087, 4088, 4092 | 0707, 0712, 0714, 071801-071802, 0812, 0816, 1301, 1306, 1307N, 1310-1313, 1317, 1320-1323, 1401-1404, 140601-1409, 15010101-150104, 150106-1503, 1505, 1506, 1508-1515, 15170101-1521, 1523-1540, 1543, 1544, 1546-1554, 1556-1558, 1560-1566, 1569-1582, 1584-1594N, 1596-1599, 9501-9511N, 9513-9517, 9519-9525, 9527, 9528, 9531-9536, 9538-9540, 9542-9551, 180101-180103, 1803-1813, 1815, 1817N-1821, 1823N-1826, 1828-1832, 2719, 2721, 2730, 350101-350403, 3506-3515, 3518-3521, 3523-3529, 3532-3536, 3538-3550, 3552-3557, 3559, 3561-3565Q, 3567-3571, 3574-3586, 3588, 3590-3594, 370101-370104, 3703N-3706, 3708, 3710-3714, 380101-3817, 39010101-39010102L, 390103-390202, 3904, 3905, 3907-391302, 3915, 3917-3920, 3922, 3923, 3925N-3928, 3930-3932, 3935, 3936, 3938Q-3945, 410301-410302, 4108, 44020101-4414, 4416, 4417, 4419N, 4421-4446, 4448-4453, 4455-446402, 460101-4610, 4613-4617, 47010101-4703, 4802-480302, 4817, 5104, 5106, 5142, 5145, 5156, 5159, 5162, 530101-5313, 5315, 5316, 5406, 5409, 5508, 5602-5604, 5609-5612, 5618, 5705, 5711, 580101-5802, 5804-5807, 5809-5813, 5815-5817N, 5819, 5820, 670101-6702, 8201, 8202, 8301 |
| 29 | 235 bp | 800 bp | 4039 | 0719, 0731, 0734, 0743, 080101-0808N, 0810, 0813-0815, 0817-0820, 0822-0824, 0826-0838, 1405, 1814, 3587, 3709, 3903, 3914, 3924, 3929, 3937, 4102, 4104, 4201, 4202, 420501-4209, 5504, Cw*0808 |
| 30 | 190 bp | 1070 bp | 400101-400203, 4004, 40060101-4008, 4010-4016, 4019, 4021-4023, 4025, 4027, 4029, 4030, 4034-4038, 4040, 4043, 4045-4050, 4053-4058, 4061, 4062, 4064-4070, 407201-4088, 4090, 4091, 4093 | 070501-0706, 0732, 0740, 0753, 1301-130204, 1307N-1309, 1311, 1313-1323, 2707, 2711, 2720, 2724, 2727, 2733, 2743, 3531, 3588, 4431, 480101-480102, 480301-4807, 4809-4816, 4818, 4819, 5110, 5116, 5131, 5134, 5208, 5509, 5522, 5524, 8101-8104N |
| 316 | 225 bp | 1070 bp | 4038, 4048, 4052, 4059, 4060 | 0709, 0711, 0717, 9538, 3566, 4606, 4814, Cw*0738 |
| 32 | 255 bp | 1070 bp | 400101-4009, 4011, 401401-4016, 4018, 4020, 4022N-4039, 4042-4046, 4048, 4050-4057, 4059-4075, 4077-4093 | 1571, 180101-1803, 1805-1808, 1810-1815, 1817N-1824, 1826-1832, 2708, 2712, 2718, 2726, 2733, 2740, 2742, 3550, 3563, 3584, 3705, 3711, 3931, 4101-4104, 4106-4108, 4202, 4209, 4501-4509, 4614, 5001, 5002, 5004, 5402, 5516, 7301 |
| 33 | 245 bp | 1070 bp | 4004, 4030, 4034, 4059, 4064, 4068 | 0712, 0714, 071801-071802, 1301, 1307N, 1313, 1317, 1320-1323, 2719, 2730, 3515, 3533, 3588, 4416, 4421, 446401-446402, 4817, 5713, Cw*0334 |
| 34 | 235 bp | 1070 bp | 4042, 4060 | 0714, 0717, 0728, 2736, 4817, Cw*0738 |
| 355,8 | 140 bp | 800 bp | 4035, 4071 | 0768, 1405, 1507, 1545, 1555, 1568, 9526, 1814, 2732, 3505, 3522, 3551, 3558, 3566, 3572, 3589, 3903, 3914, 3924, 3929, 3937, 4454, 4612, 4812, 4814, 5314, 5818 |
| 365 | 135 bp | 1070 bp | 4010, 4040, 4049, 4058 | 1546, 1553, 9506, 3547, 44020101-4405, 4409-4414, 4416, 4417, 4419N-4439, 4441-4443, 4445-4448, 4450-4454, 4456N, 4458N, 4459, 4462-446402, 47010101-4705 |
| 375,12 | 125, 270 bp | 1070 bp | 4028, 4051 | 1822, 3521, 3524, 3581, 3704, 5104, 5146, 5156, 5302, 5306, 5714, 5809 |
| 385 | 120 bp | 1070 bp | 4028 | 350402, 4462, 5104, 5146, 5156 |
| 39 | 330 bp | 800 bp | 4012 | 1503, 1509, 1510, 1518, 1523, 1529, 1537, 1546, 1547, 1549, 1551-1554, 1561, 1562, 1564, 1569, 1572, 1574, 1580, 1590, 1591, 1593, 1598, 1599, 9503, 9506, 9508, 9514, 9515, 9519, 9523, 9524, 9527, 9531-9534, 9543, 9551, Cw*0746 |
| 405 | 120 bp | 1070 bp | 4038, 4052, 4059, 4060 | 9516, 9524, 4606 |
| 41 | 205 bp | 1070 bp | 4008, 4025, 4068 | 180101-1803, 1805-1811, 1813-1815, 1817N-1824, 1826-1832, 2723, 3550, 3584, 3708, 3931, 4202weakly, 4209 weakly, 5402weakly, 5516weakly, 7301weakly |
| 425 | 95 bp | 1070 bp | 4058 | 4434 |
| 43 | 185 bp | 1070 bp | 4030, 4034, 4059 | 4817 |
| 445,13 | 110, 205 bp | 800 bp | 4054, 4078 |  |
| 4514 | 175, 235 bp | 1070 bp | 4055, 4056 |  |
| 469 | 545 bp | 1070 bp | 4011, 4020, 4026, 4074 | 1571, 180101-180103, 1803, 1805-1807, 1809-1813, 1815, 1817N, 1819-1821, 1823N, 1824, 1826, 1828-1832, 2719, 2721, 2730, 3712, 3931 |
| 475 | 145 bp | 800 bp | 4057 | 0720, 0724 |
| 4815 | 215, 245 bp | 1070 bp | 4061, 4073 | 1806, 4614 |
| 495 | 120 bp | 1070 bp | 4036, 4048 | 4807 |
| 50 | 250 bp | 1070 bp | 4037, 4076 | 2742, 3705 |
| 51 | 195 bp | 1070 bp | 4062 |  |
| 525 | 90 bp | 1070 bp | 4063, 4092 | 9516 |
| 53 | 165 bp | 1070 bp | 4065 |  |
| 545 | 130 bp | 800 bp | 4066 |  |
| 555 | 125 bp | 800 bp | 4067 |  |
| 565 | 100 bp | 1070 bp | 4069 |  |
| 57 | 380 bp | 1070 bp | 4079 | 0723 |
| 58 | 190 bp | 1070 bp | 4031, 4080 | 480101-480102, 4804, 4806, 4807, 4809, 4811, 4814-4816, 4818, 4819, 8101-8104N |
| 59 | 190 bp | 1070 bp | 4081 |  |
| 605,16 | 135, 250 bp | 800 bp | 4082, 4088 |  |
| 615 | 80 bp | 1070 bp | 4083 | 1546, 9506, 9543, 3519, 3547 |
| 62 | 195 bp | 800 bp | 4084 |  |
| 63 | 180 bp | 1070 bp | 4085 |  |
| 645 | 105 bp | 1070 bp | 4086 | 9537, 390601-390602, 3933, 3934, 4502, 511301-511302, 5214, 5513, 5622, 5903, 7301, |
| 655,17 | 130, 170 bp | 1070 bp | 400201-400203, 4004-400602, 4008, 4011, 4013-4016, 4019, 4023, 4026-4029, 4035, 4037, 4039, 4040, 4050, 4053, 4056-4058, 4064, 4068, 4070-407202, 4075, 4077, 4078, 4082, 4083, 4086, 4090, 4091, 4093 | 070501-0706, 0732, 0734, 0740, 0753, 080101-0806, 0808N-0813, 0815-0827, 0829-0836, 0838, 1303, 1318, 1322, 1509, 1530, 1542, 1548, 1558, 1573, 1583, 1586, 9550, 1826, 2707, 2711, 2720, 2727, 2733, 2743, 350201-350203, 350401-350403, 3506, 350901-350902, 3512, 3518, 3522, 3531, 3544, 3559, 3560, 3581, 3583, 3587, 3588, 3712, 380101, 380201-3817, 39010101-39010102L, 390103-390602, 3908-3914, 3916, 3918-3920, 3922, 3924, 3926-3928, 3930-3942, 3945, 4101-410302, 4105-4108, 4201, 4202, 4204-4208, 4410, 4414, 4462, 4611, 4618, 4805, 4810, 4812-4814, 4819, 510101-510202, 510204-5116, 5119-512402, 5126, 5128-5136, 5138-5141N, 5143, 5144N, 5146, 5148-5162, 520101-5211, 5213, 5214, 5401-5405N, 5407, 5408N, 5410, 5412, 5413, 5417, 550101-5505, 5507-5513, 5515-5522, 5524-5526, 5528-5534, 5601, 5602, 5604-5608, 5610, 5612-5617, 5619N-5624, 5702-570302, 5707, 5709, 5712, 5717, 5808, 5820, 5901-5903, 670101-6702, 7801-7806, 8201, 8202 |
| 665 | 145 bp | 1070 bp | 401401-4016, 4032, 407201-407202, 4077 | 070201-071802, 0720-0732, 0734-0739, 0741-0752, 0754-0759, 0761-0768, 9538, 3566, 3707, 4108, 420501-420502, 4805, 4808, 4810, 4812, 4814, Cw*0738 |
| 675,6 | 85 bp | 1070 bp | 4089 |  |
| 68 | 225 bp | 1070 bp | 4090 |  |
| 695 | 70 bp | 1070 bp | 4091 | 1321 |
| 70 | 500 bp | 1070 bp | 4003, 4009, 4018, 4020, 4024, 4031-4033, 4042, 4044, 4052, 4059, 4060 | 1304, 1310, 1546, 1553, 9506, 3546, 3547, 3563, 4104, 44020101-4405, 4409, 4411-4413, 4416, 4417, 4419N-4430, 4432-4456N, 4458N-4461N, 4463-446402, 4509, 4904, 4905 |
| 715 | 120 bp | 800 bp | 4075 | 3812 |
| 726 | 155 bp | 1070 bp | 4078 |  |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-B*40 SSP typings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective lengths of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may  be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most wells, or a band of 800 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-B*40 subtyping.

In addition, wells number 4, 5, 7, 14, 15, 16, 17, 21, 22, 23, 27, 29, 35, 39, 44, 47, 54, 55, 60, 62 and 71 contain the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

## 3For several HLA-B alleles first exon nucleotide sequences are not available. In these instances it is not known whether some of the primers of the SSP sets are completely matched with the target sequences or not. We assume that unknown sequences in the first exon are conserved within allelic groups.

## 4Due to the sharing of sequence motifs between HLA-B alleles non-HLA-B*40 alleles will be amplified by primer mixes 1, 2, 4 to 22, 25, 26, 28 to 43, 46 to 50, 52, 57, 58, 61, 64 to 66 and 69 to 71.

In addition, the Cw*0206 allele will be amplified by primer mix 5, the HLA-Cw*0305, 0325 and 0327 alleles will be amplified by primer mixes 4, 8 and 13, the Cw*0334 allele will be amplified by primer mix 33, the Cw*0738 allele will be amplified by primer mixes 16, 31, 34 and 66, the Cw*0746 allele will be amplified by primer mix 39, the Cw*0808 alleles by primer mixes 6 and 29 and the Cw*17 alleles faintly by primer mix 7, 23 and 33

5Short specific PCR fragments are less intense and not as sharp as longer specific bands.

6Primer mixes 17, 21, 31, 67 and 72 has a tendency of giving rise to nonspecific amplifications.

7Primer mixes 7 may faintly amplify the Cw*17 alleles.

8Primer mix 35 has a tendency to give rise to a primer dimer artifact.

9Primer mix 46 may give rise to less specific PCR product than the other B*40 primer mixes.

10Primer mix 16: Specific PCR fragment of 75 bp in the B*4015 allele. Specific PCR fragment of 160 bp in the B*401401, B*407201, B*4077 and B*3566, B*420501-420502 and B*4812 and Cw*0738 alleles.

11Primer mix 20: Specific PCR fragment of 135 bp in the B*401401-401403, B*4023 and the B*0707, B*0712, B*0714, B*071801-071802, 9538 and 4108 alleles. Specific PCR fragment of 285 bp in the B*4031 and B*4080 and B*480101-480102, B*4806, B*4807, B*4809, B*4811 B*4814 to 4816, B*4818, B*4819 alleles, B*8101 to 8104N alleles. Specific PCR fragments of 135 and 285 bp in the B*4015, B*4016, B*4032, B*4045 and B*070201 to 0706, B*0708 to 0711, B*0713, B*0715 to 0717, B*0720 to 0732, B*0734 to 0739, B*0741 to 0752, B*0754-0759,  B*0761 to 0768, B*3566, B*3707, B*420501-420502, B*4805, B*4808, B*4810 and B*4812 alleles. (Both specific fragments may not always be obtained.)

12Primer mix 37: Specific PCR fragment of 125 bp in the B*4051 allele. Specific PCR fragment of 270 bp in the B*4028 and B*1822, 3521, 3524, 3581, 3704, 5104, 5146, 5156, 5302, 5306, 5714, 5809 alleles.

13Primer mix 44: Specific PCR fragment of 110 bp in the B*4078 allele. Specific PCR fragment of 205 bp in the B*4054 allele.

14Primer mix 45: Specific PCR fragment of 175 bp in the B*4056 allele. Specific PCR fragment of 235 bp in the B*4055 alleles.

15Primer mix 48: Specific PCR fragment of 215 bp in the B*4061 allele. Specific PCR fragment of 245 bp in the B*4073 and B*1806 and 4614 alleles.

16Primer mix 60: Specific PCR fragment of 135 bp in the B*4082 allele. Specific PCR fragment of 250 bp in the B*4088 allele.

17Primer mix 65: Specific PCR fragment of 130 bp in the B*400201-400203, 4004-400602, 4008, 4011, 4013-401403, 4019, 4026-4029, 4035, 4037, 4039, 4040, 4050, 4053, 4056-4058, 4064, 4068, 4070-407202, 4075, 4077, 4078, 4082, 4083, 4086, 4090, 4093 and  B*0734, 080101-0805, 0808N-0813, 0815-0819, 0821-0827, 0829-0836, 0838, 1303, 1322, 1509, 1542, 1548, 1583, 1586, 1826, 2707, 2711, 2720, 2727, 2733, 2743, 350201-350203, 350401-350403, 3506, 350901-350902, 3512, 3522, 3531, 3559, 3560, 3581, 3583, 3587, 3588, 3712, 380101, 380201-3817, 39010101-39010102L, 390103-390602, 3909, 3910, 3912-3914, 3916, 3919, 3920, 3922, 3924, 3926-3928, 3930-3932, 3934, 3935, 3937-3942, 3945, 4101-410302, 4105-4108, 4201, 4202, 4204-4208, 4410, 4414, 4462, 4611, 4618, 4810, 4812-4814, 510101-510202, 510204-5104, 5106-5116, 5119-512402, 5126, 5128, 5130-5136, 5138-5141N, 5143, 5144N, 5146, 5148-5153, 5155-5160, 5162, 520101-5211, 5213, 5214, 5401-5405N, 5407, 5408N, 5412, 5413, 5417, 550101-5505, 5507, 5508, 5510-5513, 5515-5520, 5522, 5524-5526, 5528-5534, 5601, 5602, 5604-5608, 5610, 5612-5617, 5619N-5624, 570301-570302, 5707, 5709, 5717, 5808, 5820, 5901-5903, 670101-6702, 7801-7806, 8201, 8202 alleles. Specific PCR fragment of 170 bp in the B*4023 and B*1318 and 4819 alleles. Specific PCR fragment of 130 and 170 bp in the B*4015, 4016 and in the B*070501-0706, 0732, 0740, 0753, 0806, 0820, 1530, 1558, 1573, 9550, 3518, 3544, 3908, 3911, 3918, 3933, 3936, 4805, 5105, 5129, 5154, 5161, 5410, 5509, 5521, 5702 and 5712 alleles. (Both specific fragments may not always be obtained.)

1The nucleotide sequence of the B*4017 allele has been shown to be identical B*4016. The nucleotide sequence of the B*4041 allele has been shown to be identical B*4040.

2Primer mix 16: Specific PCR fragment of 75 bp in the B*4015 allele. Specific PCR fragment of 160 bp in the B*401401, B*407201, B*4077 and B*3566, B*420501-420502 and B*4812 and Cw*0738 alleles.

Primer mix 20: Specific PCR fragment of 135 bp in the B*401401-401403, B*4023 and the B*0707, B*0712, B*0714, B*071801-071802, 9538 and 4108 alleles. Specific PCR fragment of 285 bp in the B*4031 and B*4080 and B*480101-480102, B*4806, B*4807, B*4809, B*4811 B*4814 to 4816, B*4818, B*4819 alleles, B*8101 to 8104N alleles. Specific PCR fragments of 135 and 285 bp in the B*4015, B*4016, B*4032, B*4045 and B*070201 to 0706, B*0708 to 0711, B*0713, B*0715 to 0717, B*0720 to 0732, B*0734 to 0739, B*0741 to 0752, B*0754-0759,  B*0761 to 0768, B*3566, B*3707, B*420501-420502, B*4805, B*4808, B*4810 and B*4812 alleles. (Both specific fragments may not always be obtained.)Primer mix 37: Specific PCR fragment of 125 bp in the B*4051 allele. Specific PCR fragment of 270 bp in the B*4028 and B*1822, 3521, 3524, 3581, 3704, 5104, 5146, 5156, 5302, 5306, 5714, 5809 alleles .Primer mix 44: Specific PCR fragment of 110 bp in the B*4078 allele. Specific PCR fragment of 205 bp in the B*4054 allele.

Primer mix 45: Specific PCR fragment of 175 bp in the B*4056 allele. Specific PCR fragment of 235 bp in the B*4055 alleles.

Primer mix 48: Specific PCR fragment of 215 bp in the B*4061 allele. Specific PCR fragment of 245 bp in the B*4073 and B*1806 and 4614 alleles.

Primer mix 60: Specific PCR fragment of 135 bp in the B*4082 allele. Specific PCR fragment of 250 bp in the B*4088 allele.

Primer mix 65: Specific PCR fragment of 130 bp in the B*400201-400203, 4004-400602, 4008, 4011, 4013-401403, 4019, 4026-4029, 4035, 4037, 4039, 4040, 4050, 4053, 4056-4058, 4064, 4068, 4070-407202, 4075, 4077, 4078, 4082, 4083, 4086, 4090, 4093 and  B*0734, 080101-0805, 0808N-0813, 0815-0819, 0821-0827, 0829-0836, 0838, 1303, 1322, 1509, 1542, 1548, 1583, 1586, 1826, 2707, 2711, 2720, 2727, 2733, 2743, 350201-350203, 350401-350403, 3506, 350901-350902, 3512, 3522, 3531, 3559, 3560, 3581, 3583, 3587, 3588, 3712, 380101, 380201-3817, 39010101-39010102L, 390103-390602, 3909, 3910, 3912-3914, 3916, 3919, 3920, 3922, 3924, 3926-3928, 3930-3932, 3934, 3935, 3937-3942, 3945, 4101-410302, 4105-4108, 4201, 4202, 4204-4208, 4410, 4414, 4462, 4611, 4618, 4810, 4812-4814, 510101-510202, 510204-5104, 5106-5116, 5119-512402, 5126, 5128, 5130-5136, 5138-5141N, 5143, 5144N, 5146, 5148-5153, 5155-5160, 5162, 520101-5211, 5213, 5214, 5401-5405N, 5407, 5408N, 5412, 5413, 5417, 550101-5505, 5507, 5508, 5510-5513, 5515-5520, 5522, 5524-5526, 5528-5534, 5601, 5602, 5604-5608, 5610, 5612-5617, 5619N-5624, 570301-570302, 5707, 5709, 5717, 5808, 5820, 5901-5903, 670101-6702, 7801-7806, 8201, 8202 alleles. Specific PCR fragment of 170 bp in the B*4023 and B*1318 and 4819 alleles. Specific PCR fragment of 130 and 170 bp in the B*4015, 4016 and in the B*070501-0706, 0732, 0740, 0753, 0806, 0820, 1530, 1558, 1573, 9550, 3518, 3544, 3908, 3911, 3918, 3933, 3936, 4805, 5105, 5129, 5154, 5161, 5410, 5509, 5521, 5702 and 5712 alleles. (Both specific fragments may not always be obtained.)

‘w’, might be weakly amplified.

Changes in revision R01 compared to R00: 1. Fragment sizes corrected in the interpretation table. Changes in revision R03 compared to R02: 1. Fragment sizes corrected in the footnotes to primer mix 45.