Olerup SSP®

Olerup SSP® HLA-B*41

Prod. No:	101.542-06/06u

Lot No:    	68K

Expiry Date:    	2013-May-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ug/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

PHOTO DOCUMENT

Specificities and sizes of the PCR products of the 16 primer mixes used for HLA-B*41 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified HLA-B*41 alleles3 | Other amplified HLA Class I alleles4 |
| --- | --- | --- | --- | --- |
| 1 | 285 bp | 800 bp | *41:01, 41:05, 41:07, 41:12, 41:14, 41:16 | *08:09, 42:04, 55:20 |
| 2 | 215 bp | 1070 bp | *41:02:01-41:02:03, 41:04, 41:10-41:11, 41:13, 41:15 | *07:04, 07:19, 07:25, 08:01:01-08:05, 08:07-08:08N, 08:10-08:11, 08:14-08:15, 08:17-08:19N, 08:21-08:24, 08:26-08:39, 08:41-08:48, 08:50-08:54, 08:56-08:59, 08:61-08:66, 35:87, 37:09, 42:01:01-42:02, 42:05:01-42:12, 42:14, 44:106 |
| 35,6 | 80 bp | 1070 bp | *41:03:01, 41:10 |  |
| 45,8 | 90 bp, 265 bp | 800 bp | *41:04, 41:07, 41:12 | *15:46, 15:106, 35:47, 40:10:01-40:10:02, 53:15, 55:20 |
| 5 | 255 bp | 1070 bp | *41:05 |  |
| 6 | 220 bp | 1070 bp | *41:01, 41:05-41:07, 41:09, 41:12, 41:14, 41:16 | *08:09, 15:83, 42:04, 44:15, 44:18, 44:20, 44:100, 45:01-45:13, 51:08, 51:20, 51:36, 51:44N, 51:97, 52:19, 55:20, 56:13 |
| 77 | 295 bp | 800 bp | *41:06, 41:09 | *13:02:01-13:04, 13:08Q-13:09, 13:14-13:16, 13:18-13:19, 13:27, 13:30-13:35, 13:37-13:38, 13:40, 15:04, 15:16, 15:42, 15:67, 15:83, 15:95, 15:137, 15:155, 27:14, 35:37, 35:60, 39:06:01-39:06:02, 39:33-39:34, 39:50, 39:57, 39:62, 40:06:01:01-40:06:03, 40:44, 40:53, 40:70, 40:75, 40:83, 40:86, 40:93, 40:95-40:96, 40:103, 40:109-40:110, 40:127, 40:131, 40:148, 44:15, 44:18, 44:20, 44:47, 44:100, 45:01-45:13, 46:11, 46:18, 49:01:01-49:15, 50:01:01-50:02, 50:04-50:11, 51:01:01-51:03, 51:05, 51:07:01-51:24:04, 51:26-51:41N, 51:43-51:44N, 51:48-51:55, 51:57-51:58, 51:60-51:61, 51:63, 51:65-51:80, 51:82-51:95, 51:97-51:106, 52:01:01-52:22, 54:01-54:05N, 54:07-54:08N, 54:10-54:23, 55:01:01-55:03, 55:05, 55:07, 55:09-55:13, 55:15-55:19, 55:21-55:48, 56:01:01-56:01:04, 56:05:01-56:08, 56:13-56:17, 56:19N-56:30, 58:08, 59:01-59:05, 73:01-73:02, 78:01-78:07 |
| 86 | 215 bp | 1070 bp | *41:01-41:16 | *13:23, 15:71, 15:175, 15:204, 18:01:01-18:03, 18:05-18:06, 18:08-18:15, 18:17N-18:24, 18:26-18:28, 18:30-18:32, 18:34-18:42, 18:44-18:52, 27:18, 27:29, 35:63, 37:01:01-37:20, 37:22-37:25, 40:01:01-40:02:02, 40:02:04-40:06:03, 40:09, 40:11:01-40:11:02, 40:14:01-40:16, 40:18-40:20, 40:22N-40:39, 40:42-40:48, 40:50-40:57, 40:59-40:60, 40:62-40:67, 40:69-40:74, 40:76-40:102, 40:104-40:105, 40:107-40:109, 40:111-40:128, 40:130-40:136, 40:138-40:158, 44:15, 44:18, 44:55, 44:103, 45:01-45:05, 45:07-45:13, 48:22, 49:01:01-49:09, 49:11-49:15, 50:01:01-50:02, 50:04-50:12 |
| 95,6 | 105 bp | 1070 bp | *41:08 | *07:04, 07:25, 40:136, 42:05:01-42:05:02 |
| 105 | 105 bp | 800 bp | *41:01-41:05, 41:07, 41:10-41:14, 41:16 | *08:01:01-08:05, 08:07-08:12:03, 08:14-08:19N, 08:21-08:24, 08:26-08:30N, 08:32-08:39, 08:41-08:42, 08:44-08:54, 08:56-08:66, 35:87, 42:01:01-42:02, 42:04, 42:06-42:07, 42:09-42:13, 55:20 |
| 11 | 245 bp | 800 bp | *41:01, 41:05-41:07, 41:09, 41:12, 41:14, 41:16 | *08:09, 15:42, 15:83, 35:60, 39:06:01-39:06:02, 39:34, 39:57, 39:62, 42:04, 45:08, 46:18, 51:21, 51:36, 51:101, 54:01-54:02, 54:04-54:05N, 54:07-54:08N, 54:10-54:23, 55:01:01-55:03, 55:05, 55:07, 55:10-55:13, 55:15-55:21, 55:23, 55:25-55:48, 56:23, 59:01-59:03, 59:05, 78:07 |
| 12 | 210 bp | 1070 bp | *41:02:01-41:04, 41:08, 41:10-41:11, 41:13, 41:15 | *07:04, 07:19, 07:25, 08:01:01-08:05, 08:07-08:08N, 08:10-08:12:02, 08:14-08:19N, 08:21-08:24, 08:26, 08:28-08:32, 08:34-08:39, 08:41-08:48, 08:50-08:54, 08:56-08:59, 08:61-08:66, 15:51, 15:179, 15:199, 35:87, 37:09, 37:12, 40:136, 42:01:01-42:02, 42:05:01-42:12, 42:14, 44:41, 44:106, C*07:04:03 |
| 13 | 285 bp | 1070 bp | *41:09 | *13:02:01-13:04, 13:08Q-13:09, 13:14-13:16, 13:18-13:19, 13:27, 13:30-13:35, 13:37-13:38, 13:40, 15:04, 15:16, 15:42, 15:67, 15:83, 15:95, 15:137, 15:155, 27:14, 35:37, 35:60, 39:06:01-39:06:02, 39:33-39:34, 39:50, 39:57, 39:62, 40:06:01:01-40:06:03, 40:44, 40:53, 40:70, 40:75, 40:83, 40:86, 40:93, 40:95-40:96, 40:103, 40:109-40:110, 40:127, 40:131, 40:148, 44:15, 44:18, 44:20, 44:47, 44:100, 45:01-45:13, 46:11, 46:18, 49:01:01-49:13, 49:15, 50:01:01-50:02, 50:04-50:11, 51:01:01-51:03, 51:05, 51:07:01-51:24:04, 51:26-51:41N, 51:43-51:44N, 51:48-51:55, 51:57-51:58, 51:60-51:61, 51:63, 51:65-51:80, 51:82-51:95, 51:97-51:106, 52:01:01-52:22, 54:01-54:05N, 54:07-54:08N, 54:10-54:23, 55:01:01-55:03, 55:05, 55:07, 55:09-55:13, 55:15-55:19, 55:21-55:48, 56:01:01-56:01:04, 56:05:01-56:08, 56:13-56:17, 56:19N-56:30, 58:08, 59:01-59:05, 73:01-73:02, 78:01-78:07 |
| 145,6,9 | 115 bp, 215 bp | 1070 bp | *41:13, 41:16 |  |
| 1510 | 155 bp, 375 bp | 800 bp | *41:11, 41:14 | *13:30, 40:79 |
| 16 | 290 bp | 1070 bp | *41:15 | *07:33, 07:53, 07:60, 07:100, 14:05, 14:13, 15:07:01-15:07:02, 15:45, 15:55, 15:68, 15:126, 15:207, 18:14, 27:07, 27:11, 27:24, 27:32-27:34, 27:43, 35:05:01-35:05:02, 35:22, 35:31, 35:51, 35:58, 35:72, 35:89, 35:97, 35:114, 37:09, 38:19, 39:03, 39:14, 39:24, 39:29, 39:37, 40:02:01-40:03, 40:05, 40:08-40:09, 40:13, 40:18-40:19, 40:24, 40:27, 40:29, 40:35, 40:37, 40:39-40:40, 40:50, 40:56-40:58, 40:71, 40:78, 40:82, 40:85, 40:89-40:91, 40:94, 40:97, 40:104-40:105, 40:107, 40:111, 40:115, 40:119, 40:122, 40:133Q, 40:142N-40:145, 40:157, 42:08, 42:14, 44:54, 44:106, 46:12, 48:04, 48:13, 48:24, 51:64, 51:81, 53:14, 55:04, 58:18, 58:27, C*03:05, C*03:25, C*03:27, C*08:08 |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-B*41 SSP typings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective lengths of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may  be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most wells, or a band of 800 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-B*41 subtyping.

In addition, wells number 4, 7, 10, 11 and 15 contain the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3The B*41:07 and 41:12 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 4.

4Due to the sharing of sequence motifs between HLA-B alleles some non-HLA-B*41 alleles will be amplified by primer mixes 1, 2, 4, 6 to 13, 15 and 16. In addition, the C*07:04:03 allele will be amplified by primer mix 12 and the C*03:05, C*03:25, C*03:27 and C*08:08 alleles will be amplified by primer mix 16.

The 41:08 and 40:136 alleles give rise to identical amplification patterns with the HLA-B*41 high resolution kit. These two alleles can be distinguished by the HLA-B low resolution and/or HLA-B*40 kits.

The B*41:09 and 45:08 alleles give rise to identical amplification patterns with the HLA-B*41 high resolution kit. These two alleles can be distinguished by the HLA-B low resolution and/or HLA-B*45 kits.

The 41:15 and 37:09 alleles give rise to identical amplification patterns with the HLA-B*41 high resolution kit. These two alleles can be distinguished by the HLA-B low resolution and/or HLA-B*37 kits.

5Short specific PCR fragments are less intense and not as sharp as longer specific bands.

6Primer mixes 3, 8, 9 and 14 may give rise to nonspecific amplifications.

7Primer mix 7 has tendencies of giving rise to primer dimer artefacts.

8Primer mix 4:   Specific PCR fragment of 90 bp in the B*41:12 and the B*15:46, 15:106, 35:47 and 40:10:01-40:10:02 alleles. Specific PCR fragment of 265 bp in the B*41:04 and 41:07 and in the B* 53:15 and 55:20 alleles.

9Primer mix 14: Specific PCR fragment of 115 bp in the B*41:13 allele. Specific PCR fragment of 215 bp in the B*41:16 allele.

10Primer mix 15: Specific PCR fragment of 155 bp in the B*41:11 and the B*13:30 alleles. Specific PCR fragment of 375 bp in the B*41:14 and the B*40:79 alleles.

1The 41:08 and 40:136 alleles give rise to identical amplification patterns with the HLA-B*41 high resolution kit. These two alleles can be distinguished by the HLA-B low resolution and/or HLA-B*40 kits.

The B*41:09 and 45:08 alleles give rise to identical amplification patterns with the HLA-B*41 high resolution kit. These two alleles can be distinguished by the HLA-B low resolution and/or HLA-B*45 kits.

The 41:15 and 37:09 alleles give rise to identical amplification patterns with the HLA-B*41 high resolution kit. These two alleles can be distinguished by the HLA-B low resolution and/or HLA-B*37 kits.

2The B*41:07 and 41:12 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 4.

3Primer mix 4:   Specific PCR fragment of 90 bp in the B*41:12 and the B*15:46, 15:106, 35:47 and 40:10:01-40:10:02 alleles. Specific PCR fragment of 265 bp in the B*41:04 and 41:07 and in the B* 53:15 and 55:20 alleles.

Primer mix 14: Specific PCR fragment of 115 bp in the B*41:13 allele. Specific PCR fragment of 215 bp in the B*41:16 allele.

Primer mix 15: Specific PCR fragment of 155 bp in the B*41:11 and the B*13:30 alleles. Specific PCR fragment of 375 bp in the B*41:14 and the B*40:79 alleles.