Gel Documentation Form and Worksheet

HLA-B\*41 (101.542-06/06u) Lot No: 5G8 Expiry Date: 2021-01-01

Name:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

Sample ID:\_\_\_\_\_\_\_\_\_\_\_\_\_\_

DNA Conc.(ng/ul):\_\_\_\_\_\_\_\_\_

Test Date:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

Tested By:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

Review Date:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

Reviewed By:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

*Interpretation:\_\_\_\_\_\_\_\_\_\_\_ Failed lanes: \_\_\_\_\_\_\_\_\_\_\_ Comments:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_*

**Gel Picture**

|  |
| --- |
| PHOTO DOCUMENT |



‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

**Notes:**

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 9 may have tendencies of unspecific amplifications.

Primer mix 24 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 430 base pairs.









**1**HLA-B\*41 alleles in bold lettering are listed as confirmed alleles on the on the IMGT/HLA web page [www.ebi.ac.uk/imgt/hla](http://www.ebi.ac.uk/imgt/hla), release 3.27.0, January 2017

**2**Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA

database release can be retrieved from web page <http://hla.alleles.org/alleles/deleted.html>.

**3**The following HLA-B\*41 primer mixes have two or more product sizes:

|  |  |  |  |
| --- | --- | --- | --- |
| Primer Mix | Size of spec. PCR product | Amplified HLA-B\*41 alleles | Other amplified HLA Class I alleles |
| **3** | 75 bp225 bp | \*41:03:01, 41:10, 41:22\*41:45N | \*07:143, 08:62, 15:210, 15:426, 18:01:01:01-18:01:12, 18:01:14-18:03, 18:05:01:01-18:15, 18:17N-18:24, 18:26-18:28, 18:30-18:42, 18:44:01-18:79, 18:81-18:97, 18:99-18:101, 18:103-18:106, 18:108-18:111, 18:113-18:119, 18:121-18:124, 18:126-18:143, 27:75, 35:162, 37:01:01:01-37:21, 37:23-37:40, 37:42N-37:54, 37:56-37:61, 37:63-37:70, 39:31:01:01-39:31:01:02, 39:101, 39:128, 40:149, 42:02:01:01-42:02:01:02, 42:09, 42:17-42:18, 48:22, 48:30, 51:116, **C\*06:77** |
| **4** | 90 bp 265 bp | \*41:12, 41:36\*41:04, 41:07 | \*15:46, 15:106, 35:47, 35:154, 40:10:01:01-40:10:02, 40:248, 40:282, 40:325, 40:351, 40:364, 44:34:02, 44:245, 44:269, 49:48, 53:30, 57:45 \*08:96, 08:107, 53:15, 55:20 |
| **5** | 230 bp255 bp | \*41:23, 41:27\*41:05 | \*07:143, 08:62, 18:03, 27:04:01-27:04:06, 27:06, 27:08, 27:11-27:12:01:02, 27:15, 27:20-27:21, 27:24-27:26, 27:31, 27:33, 27:36, 27:40, 27:61, 27:63, 27:66N, 27:68-27:69, 27:79, 27:86, 27:89w, 27:91, 27:100, 27:103, 27:105-27:109, 27:112-27:115, 27:120, 27:125, 27:138, 27:147, 27:149, 27:153-27:154, 27:165, 37:37, 39:31:01:01-39:31:01:02, 39:101, 39:128, 40:43, 40:50, 40:166, 42:02:01:01-42:02:01:02, 42:09, 42:17-42:18, 45:06, 46:14w, 50:52, 54:02, 55:16, 56:35, **C\*03:278w, C\*16:86w** |
| **14** | 115 bp180 bp 215 bp | \*41:13\*41:49\*41:16 |  |
| **17** | 85 bp245 bp | \*41:19\*41:49 | \*08:40, 08:165, 08:171, 27:83 |
| **19** | 130 bp155 bp | \*41:28\*41:11 | \*44:80\*13:30, 40:293, 44:219 |

**4**The B\*41:08 and the B\*40:136, 40:231, 40:279 alleles give rise to identical amplification patterns with the HLA-B\*41 high resolution kit. These alleles can be distinguished by the HLA-B low resolution and/or HLA-B\*40 kits.

The B\*41:46-41:47 and the B\*44:166 alleles give rise to identical amplification patterns with the HLA-B\*41 high resolution kit. These alleles can be distinguished by the HLA-B low resolution and/or HLA-B\*44 kits.

**5**The following HLA-B\*41 alleles can be distinguished by the different sizes of the HLA-specific PCR product:

|  |  |
| --- | --- |
| **Alleles** | **Primer mix** |
| B\*41:04, 41:36 | 4 |
| B\*41:10, 41:45N | 3 |

‘w’, might be weakly amplified.

‘?’, nucleotide sequence information not available for the primer matching sequence.

Changes in revision R01 compared to R00:

1. The size of the positive control band of primer mix 23 has been corrected to 800 bp in the Product Insert and the Worksheet.