HLA-B\*44 (101.563-24/03, -24u/03u) Lot No: 3R2 Expiry Date: 2026-08-01

Name:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

Sample ID:\_\_\_\_\_\_\_\_\_\_\_\_\_\_

DNA Conc.(ng/ul):\_\_\_\_\_\_\_\_\_

Test Date:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

Tested By:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

Review Date:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

Reviewed By:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

***Interpretation:\_\_\_\_\_\_\_\_\_\_\_ Failed lanes: \_\_\_\_\_\_\_\_\_\_\_ Comments:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_***

**Gel Picture**

|  |
| --- |
| PHOTO DOCUMENT |







Abbreviations

ICB: Internal Control Band

AmpS: Amplicon Size

**Notes:**

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mix 36 has a tendency giving rise to primer oligomer formation.

Primer mixes 25, 28, 29, 37, 39, 42, 45, 46, 51, 52, 55 and 58 may have tendencies of unspecific amplifications.

Primer mixes 31, 37 and 39 may give rise to a lower yield of HLA-specific PCR product than the other B\*44 primer mixes.

We assume that the fourth exon nucleotide sequences are conserved within allelic groups. Primer mix 25 amplifies the B\*44:27 and B\*44:31 alleles. The purpose of primer mixes 23 and 25 is to distinguish e.g. the B\*44:02,44:02 and B\*44:02,44:27 genotypes.

Primer mix 64 contains a negative control, which will amplify the majority of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 200 base pairs.































**1**HLA-B\*44 alleles in bold lettering are listed as confirmed alleles on the on the IMGT/HLA web page [www.ebi.ac.uk/imgt/hla](http://www.ebi.ac.uk/imgt/hla), release 3.29.0, August 2017.

**2**Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page <http://hla.alleles.org/alleles/deleted.html>.

**3**The following HLA-B\*44 primer mixes have two or more product sizes:

|  |  |  |  |
| --- | --- | --- | --- |
| Primer Mix | Size of spec. PCR product | Amplified HLA-B\*44 alleles | Other amplified HLA Class I alleles |
|  **5** | 75 bp 255 bp | \*44:04, 44:28:01-44:28:02, 44:56N, 44:76\*44:70  | \*07:20, 07:140, 07:153, 35:45, **C\*04:345, C\*05:102, C\*05:262, C\*14:61** |
|  **6** | 170 bp195 bp | \*44:39\*44:05:01:01-44:05:05, 44:10, 44:15:01:01-44:15:01:02, 44:17-44:18:01:02, 44:25, 44:31, 44:43:01-44:43:02, 44:58N, 44:62, 44:70, 44:75, 44:77-44:78, 44:82, 44:107, 44:117, 44:123, 44:134, 44:136, 44:139-44:140, 44:144, 44:158, 44:184, 44:213, 44:226, 44:230, 44:236, 44:250-44:251, 44:256, 44:263-44:264, 44:288, 44:308, 44:330, 44:339-44:340, 44:355, 44:465, 44:504, 44:506 | \*15:561\*07:20, 07:140, 07:153, 15:14, 15:91, 15:131, 15:161, 15:518, 15:551, 15:592, 15:634, 18:56, 18:169, 35:45, 35:71, 35:169, 35:309, 35:452, 35:545, 39:173, 40:132, 40:178, 45:01:01:01-45:02, 45:05-45:07, 45:09, 45:11-45:23, 45:25, 45:27-45:30, 46:17, 49:20, 50:02:01:01-50:02:01:02, 50:21, 50:56, 51:23, 51:108, 52:15, 53:22, 55:122, 56:74, 57:07, 57:93, 57:119, 58:07, 82:01:01:01-82:03, **C\*02:82, C\*04:345,** **C\*05:102, C\*05:184, C\*05:262, C\*06:165, C\*07:979, C\*08:93, C\*14:61, C\*15:25** |
|  **9** | 120 bp 195 bp | \*44:23:01:01N-44:23:01:02N\*44:10, 44:15:01:01-44:15:01:02, 44:18:01:01-44:18:01:02, 44:140, 44:339-44:340, 44:355 | \*35:545, 45:01:01:01-45:01:12, 45:05-45:07, 45:11-45:23, 45:25, 45:27-45:28N, 45:30, 49:20, 50:02:01:01-50:02:01:02, 50:21, 50:56, 55:122, 56:74, 82:01:01:01-82:03 |
| **10** | 125 bp 150 bp175 bp | \*44:24, 44:181 \*44:92\*44:237N | \*13:132, 41:55, **C\*02:167, C\*03:504, C\*04:459, C\*05:205, C\*06:284, C\*07:828** |
| **18** | 115 bp 155 bp 260 bp | \*44:31, 44:169, 44:182\*44:74\*44:19N, 44:82, 44:96, 44:153  | \*40:132, 40:178\*57:07, 57:93, 57:119 |
| **19** | 150 bp 190 bp | \*44:11\*44:30 |  |
| **21** | 100 bp 180 bp250 bp | \*44:32, 44:326 \*44:16, 44:21, 44:64:01-44:64:02, 44:124, 44:150, 44:347\*44:86  | \*27:148\*07:28, 27:227, 40:414, 40:498, **C\*02:97**\*51:246, 52:56, 53:27 |
| **22** | 205 bp290 bp | \*44:177, 44:228\*44:26  | \*13:81, 35:183, 35:502, 51:209\*35:138, 47:07 |
| **26** | 100 bp 145 bp 225 bp | \*44:69:01-44:69:02\*44:33, 44:81\*44:85:01-44:85:02 | **\***38:51, **A\*24:370N**\*51:226, 58:107, **A\*24:137, A\*32:96** |
| **27** | 160 bp 255 bp | \*44:02:10, 44:34:01\*44:15:01:01-44:15:01:02, 44:18:01:01-44:18:01:02, 44:20, 44:47, 44:100, 44:197, 44:339, 44:355 | \*15:01:76, 15:32:02, 37:22:02, 40:01:11, 40:58, 49:03, 51:310, 52:01:02:01-52:01:02:09, 52:01:04, 52:01:09, 52:01:40, 52:01:43, 52:01:47, 52:01:49, 52:02:01, 52:03, 52:06:01-52:06:02, 52:09:01, 52:10:03, 52:11:02, 52:21:01, 52:31:02, 52:85, 52:97, 52:102, 52:104\*45:01:01:01-45:03, 45:05-45:07, 45:09, 45:11-45:23, 45:25, 45:27-45:30, 49:20, 50:02:01:01-50:02:01:02, 50:21, 50:56, 51:23, 51:108, 52:15, 55:122, 56:74 |
| **28** | 130 bp 165 bp | \*44:68\*44:35 |  |
| **29**  | 90 bp 255 bp | \*44:117, 44:251\*44:15:01:01-44:15:01:02, 44:55, 44:103, 44:188, 44:212 | **\***14:10, 15:460, 35:59:01, 35:242, 37:19:01, 38:20:01, 39:42, 39:106, **C\*03:26:02, C\*15:140, C\*15:152:03, C\*17:01:10****\***13:23, 13:55, 18:09, 27:01, 27:142, 37:10, 40:47, 40:96, 40:157, 40:201, 49:02, 49:70 |
| **31** | 110 bp 195 bp 230 bp 260 bp | \*44:94\*44:17, 44:43:01-44:43:02, 44:144, 44:230, 44:256, 44:308\*44:98\*44:99  | \*15:14, 15:91, 15:131, 15:161, 15:518, 15:551, 15:592, 15:634, 18:56, 18:169, 35:45, 35:71, 45:09, 46:17, 53:22, 57:93, 57:119, 58:07, **C\*02:82, C\*06:165, C\*07:979, C\*14:61****A\*02:432, A\*29:151, A\*30:29, A\*30:122, A\*31:61** |
| **33** | 165 bp195 bp | \*44:80-44:81\*44:16, 44:37:01-44:37:03, 44:64:01-44:64:02, 44:91, 44:132, 44:150, 44:190, 44:268, 44:275, 44:290, 44:347 | \*57:126\*07:28, 08:87, 15:77, 15:233, 18:166, 27:227, 40:414, 40:498, 49:61, 51:261, 51:274, 52:26, 57:04:01-57:04:03, 58:114, **A\*23:31, A\*24:106, C\*01:115, C\*02:97, C\*03:433, C\*03:514, C\*04:439, C\*06:306, C\*07:231, C\*14:79, C\*16:10** |
| **35** | 85 bp 185 bp | \*44:63\*44:04, 44:56N, 44:132, 44:137, 44:163:01-44:163:02, 44:190, 44:268, 44:446 | \*08:87, 18:166, **C\*01:115, C\*04:439, C\*06:306, C\*07:231, C\*14:79, C\*16:10** |
| **36** | 145 bp 235 bp | \*44:40, 44:44, 44:130, 44:156, 44:210:01-44:210:02\*44:142  | \*07:149, 07:453, 08:02, 08:117, 08:255, 08:271, 13:97, 15:115, 27:01, 27:142, 38:02:01:01-38:02:06, 38:02:08-38:04, 38:08, 38:15, 38:18, 38:23, 38:29, 38:35, 38:43-38:50, 38:62, 38:64, 38:72, 38:74-38:76, 38:79, 38:82, 38:84-38:85, 38:89, 38:92, 38:97-38:98, 38:160, 38:162-38:165N, 38:167, 38:172-38:173Q, 56:07, 59:08 |
| **40** | 235 bp 260 bp | \*44:10, 44:15:01:01-44:15:01:02, 44:18:01:01-44:18:01:02, 44:48, 44:256, 44:339-44:340, 44:355, 44:481\*44:31, 44:41:01-44:41:02, 44:54, 44:65, 44:106, 44:135, 44:158, 44:184, 44:213, 44:230, 44:236, 44:440 | \*35:45, 35:71, 35:169, 35:309, 35:452, 35:545, 45:01:01:01-45:03, 45:05-45:07, 45:09, 45:11-45:23, 45:25, 45:27-45:30, 49:20, 50:02:01:01-50:02:01:02, 50:21, 50:56, 53:22, 55:122, 56:74, 58:07, 82:01:01:01-82:03, **C\*02:82, C\*04:345, C\*05:102, C\*05:184, C\*05:262, C\*06:165, C\*07:979, C\*08:93, C\*14:61, C\*15:25****\***07:20, 07:140, 07:153, 15:14, 15:91, 15:131, 15:161, 15:518, 15:551, 15:578, 15:634, 18:56, 18:169, 27:47, 27:50:01-27:50:02, 39:173, 40:132, 40:178, 46:17, 58:07, 82:01:01:01-82:03, **C\*02:82, C\*04:345, C\*05:102, C\*05:184, C\*05:262, C\*06:165, C\*07:979, C\*08:93, C\*14:61** |
| **43** | 110 bp 165 bp | \*44:52N, 44:115\*44:29, 44:89, 44:285 |  |
| **44** | 85 bp 160 bp | \* 44:53:01-44:53:02, 44:182, 44:192:01-44:192:04, 44:343\*44:73, 44:76, 44:79:01-44:79:02, 44:146, 44:150, 44:275, 44:299 | **\***07:28, 08:87, 40:414, 40:498, **C\*01:115, C\*02:97, C\*03:433, C\*04:439, C\*06:306, C\*07:231, C\*14:79, C\*16:10**\*07:28, 15:77, 15:233, 49:61, 51:261, 52:26, **C\*01:115, C\*02:97, C\*03:433, C\*03:514, C\*04:439, C\*06:306, C\*14:79** |
| **45** | 220 bp 250 bp285 bp | \*44:36\*44:54, 44:106, 44:135, 44:158, 44:184, 44:213 \*44:173 | \*07:20, 07:140, 07:153\*13:169 |
| **46** | 175 bp450 bp | \*44:66, 44:267N\*44:56N  | \*27:208, 51:11N\*40:337N |
| **47** | 65 bp 115 bp140 bp195 bp | \*44:93, 44:139\*44:72, 44:204:01-44:204:02\*44:191\*44:58N  | \*40:392 |
| **48** | 105 bp 140 bp 175 bp260 bp595 bp | \*44:61N\*44:59:01-44:59:02, 44:136, 44:191 \*44:237N\*44:78, 44:120, 44:195N\*44:59:01-44:59:02, 44:136 | \*35:223, 40:392, 53:60\*15:454N, 35:204, **C\*03:147** |
| **50** | 95 bp160 bp195 bp 285 bp | \*44:111, 44:200\*44:171N\*44:137, 44:227\*44:148  | \*15:627, 58:76, **C\*03:251, C\*03:314, C\*05:10, C\*05:148, C\*08:44, C\*08:61, C\*08:82, C\*08:126, C\*08:209, C\*15:130, C\*15:229** |
| **51** | 120 bp200 bp240 bp | \*44:115-44:116\*44:217N\*44:259, 44:543 |  |
| **52** | 130 bp250 bp285 bp405 bp | \*44:143, 44:261\*44:86\*44:173\*44:118, 44:546  | \*51:246, 52:56, 53:27\*13:169\*39:46 |
| **53** | 210 bp240 bp | \*44:122, 44:451\*44:259, 44:543 |  |
| **54** | 175 bp270 bp | \*44:253, 44:489\*44:126:01-44:126:02, 44:159, 44:195N  | \*35:377, 35:469, **C\*03:337****\***15:394, 15:454N, 35:247:01-35:247:02, 35:510, 58:45:01-58:45:02, **C\*03:96,** **C\*15:52:01-15:52:02** |
| **55** | 135 bp255 bp | \*44:116\*44:110, 44:138Q, 44:149N, 44:211 |  |
| **56** | 110 bp250 bp280 bp | \*44:198N\*44:160Q \*44:104, 44:148  | **C\*03:265N**\*35:333Q, 37:16Q, **C\*15:96Q**\*15:627, 58:76, **C\*03:251, C\*03:314, C\*05:10, C\*05:148, C\*08:44, C\*08:61, C\*08:82, C\*08:126, C\*08:209, C\*15:130, C\*15:229** |
| **58** | 100 bp210 bp245 bp | \*44:187, 44:487\*44:87-44:88, 44:223\*44:271 |  |
| **61** | 210 bp255 bp | \*44:130, 44:156, 44:221, 44:224, 44:258\*44:15:01:01-44:15:01:02, 44:55, 44:103, 44:188, 44:212 | \*07:149, 07:453, 08:02, 08:117, 08:255, 08:271, 15:115, 18:09, 37:10, 38:02:01:01-38:02:06, 38:02:08-38:04, 38:08, 38:15, 38:18, 38:23, 38:29, 38:35, 38:43-38:45, 38:47-38:50, 38:62, 38:64, 38:72, 38:74-38:76, 38:79, 38:82, 38:84-38:85, 38:89, 38:92, 38:97-38:98, 38:160, 38:162-38:165N, 38:167, 38:172-38:173Q\*13:23, 13:55, 18:09, 27:01, 27:142, 37:10, 40:47, 40:96, 40:157, 40:201, 49:02, 49:70 |
| **62** | 115 bp175 bp300 bp | \*44:38\*44:253, 44:489\*44:71  | **\***13:80, 35:326, **C\*15:147**\*35:377, 35:469, **C\*03:337****C\*15:141** |

**4**The following HLA-B\*44 alleles can be distinguished by the different sizes of the HLA-specific PCR product:

|  |  |  |  |
| --- | --- | --- | --- |
| **Alleles** | **Primer mix** | **Alleles** | **Primer mix** |
| B\*44:21, 44:326 | 21 | B\*44:78, 44:136 | 48 |
| B\*44:58N, 44:139 | 47 | B\*44:79:01-44:79:02, 44:192:01 | 44 |
| B\*44:69:01, 44:85:01 | 26 | B\*44:94, 44:98 | 31 |
| B\*44:72, 44:93 | 47 | B\*44:118, 44:261 | 52 |
| B\*44:74, 44:169 | 18 | B\*44:160Q, 44:198N | 56 |

Abbreviations

w: might be weakly amplified.

?: nucleotide sequence information not available for the primer matching sequence.