PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® HLA-B*50

Prod. No:	101.548-06/06u

Lot No:    	89R

Expiry Date:    	2015-July-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 16 primer mixes used for HLA-B*50 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified HLA-B*50 alleles | Other amplified HLA Class I alleles3 |
| --- | --- | --- | --- | --- |
| 1 | 255 bp | 800 bp | *50:01:01-50:01:03, 50:04-50:11, 50:13-50:16, 50:18-50:19 | *07:78, 07:84, 08:09, 08:84, 13:02:01-13:04, 13:08Q-13:09, 13:14-13:16, 13:18-13:19, 13:27, 13:30-13:35, 13:37-13:38, 13:40-13:42, 13:44-13:49N, 13:53-13:56N, 13:58-13:59, 13:62, 15:04, 15:16:01-15:16:03, 15:42, 15:67, 15:83, 15:95, 15:137, 15:155, 15:222, 15:254, 27:14, 27:81, 35:37, 35:60, 39:06:01-39:06:02, 39:33-39:34, 39:50, 39:57, 39:62, 39:64, 40:06:01:01-40:06:06, 40:44, 40:53, 40:70, 40:75, 40:83, 40:86, 40:93, 40:95-40:96, 40:103, 40:109-40:110, 40:127, 40:131, 40:148, 40:159, 40:161-40:162, 40:165, 40:167, 40:177, 40:190, 41:01, 41:05-41:07, 41:09, 41:12, 41:14, 41:16-41:17, 41:20-41:22, 41:25, 42:04, 45:04, 45:08, 45:10, 46:11, 46:18, 46:32, 49:01:01-49:19N, 49:21-49:25, 51:01:01-51:03, 51:05, 51:07:01-51:22, 51:24:01-51:24:04, 51:26-51:41N, 51:43-51:44N, 51:48-51:55, 51:57-51:58, 51:60-51:61:02, 51:63, 51:65-51:69, 51:71-51:80, 51:82-51:107, 51:109-51:134, 51:136-51:138, 51:140-51:147, 52:01:01:01-52:14, 52:16-52:31, 54:01:01-54:05N, 54:07-54:08N, 54:10-54:28, 55:01:01-55:01:06, 55:01:08-55:03, 55:05, 55:07, 55:09-55:13, 55:15-55:48, 55:50, 55:52-55:57, 56:01:01-56:01:04, 56:05:01-56:08, 56:13-56:17, 56:19N-56:30, 56:33-56:38N, 58:08:01-58:08:02, 59:01:01:01-59:05, 73:01-73:02, 78:01:01-78:07 |
| 2 | 255 bp | 1070 bp | *50:02 | *44:15, 44:18, 44:20, 44:47, 44:100, 45:01-45:03, 45:05-45:07, 45:09, 45:11-45:14, 49:20, 51:23, 51:108, 52:15 |
| 34,5 | 80 bp | 800 bp | *50:04 | *15:18:03, 15:73, 15:224, 39:17, 40:48, 40:71, 44:140, 51:62, 56:04, A*02:339, C*03:12, C*03:19, C*03:102, C*15:26 |
| 44 | 105 bp | 1070 bp | *50:01:01-50:02, 50:04-50:08, 50:10-50:11, 50:13, 50:15, 50:18-50:19 | *07:78, 13:02:01-13:02:14, 13:03, 13:08Q-13:09, 13:14-13:16, 13:18-13:19, 13:27, 13:30-13:34, 13:37-13:38, 13:40-13:42, 13:44-13:45, 13:48-13:49N, 13:53-13:56N, 13:58, 15:42, 35:60, 44:15, 44:18, 45:01, 45:03-45:08, 45:10-45:14, 46:11, 46:18, 49:01:01-49:03, 49:06-49:17, 49:19N-49:25, 51:15, 52:25, 54:01:01-54:03, 54:05N, 54:07-54:08N, 54:10, 54:12-54:13, 54:16-54:28, 55:01:01-55:01:06, 55:01:08-55:03, 55:05, 55:07, 55:09-55:12, 55:15-55:16, 55:18-55:19, 55:21-55:22, 55:24-55:26, 55:29-55:31, 55:33-55:38, 55:40-55:41, 55:43, 55:45-55:48, 55:50, 55:52, 55:54-55:57, 56:01:01-56:01:04, 56:07-56:08, 56:13-56:14, 56:16-56:17, 56:19N-56:20:02, 56:23-56:30, 56:33-56:38N, 59:01:01:01-59:01:01:02, 59:04-59:05 |
| 5 | 145 bp | 1070 bp | *50:01:01-50:02, 50:04-50:05, 50:07-50:19 | *07:133, 15:46, 15:53, 15:106, 15:143, 15:212, 18:48, 35:19, 35:47, 35:63, 35:154, 35:217, 39:69, 40:01:01-40:11:02, 40:14:01-40:16, 40:18, 40:20, 40:22N-40:40, 40:42-40:45, 40:48-40:75, 40:77-40:92, 40:94-40:95, 40:97-40:108, 40:111-40:116, 40:118N-40:136, 40:138-40:156, 40:158-40:184, 40:186-40:187, 40:189-40:200, 40:202-40:207, 40:210-40:226, 41:01-41:04, 41:06-41:25, 44:09, 44:46, 44:75, 44:90, 44:129, 44:131, 45:01-45:14, 47:02, 47:03w |
| 66 | 160 bp | 1070 bp | *50:01:01-50:02, 50:04-50:08, 50:10-50:13, 50:15-50:19 | *13:01:01-13:03, 13:06-13:09, 13:11-13:12, 13:14-13:17, 13:19-13:20, 13:22:01-13:23, 13:25, 13:27-13:30, 13:32-13:34, 13:36-13:40, 13:42-13:45, 13:47-13:50, 13:52-13:58, 13:60-13:61, 13:63N, 15:42, 15:86, 15:224, 35:60, 39:17, 40:48, 40:71, 44:10, 46:11, 46:18, 49:01:01-49:03, 49:06-49:17, 49:19N-49:25, 51:15, 51:62, 51:106, 52:25, 54:01:01-54:03, 54:05N, 54:07-54:08N, 54:12-54:13, 54:17-54:19, 54:21-54:28, 55:01:01-55:03, 55:05, 55:07, 55:10-55:12, 55:15-55:16, 55:18-55:19, 55:22, 55:24-55:26, 55:29-55:31, 55:33-55:36, 55:38-55:45, 55:47-55:48, 55:50, 55:53-55:55N, 55:57, 56:01:01-56:02, 56:04, 56:07-56:08, 56:10, 56:14, 56:16-56:17, 56:19N-56:20:02, 56:23-56:30, 56:33-56:38N, 59:01:01:01-59:01:01:02, 59:04-59:05 |
| 7 | 150 bp | 1070 bp | *50:01:01-50:01:02, 50:01:03w, 50:02, 50:04-50:08, 50:10-50:13, 50:15-50:19 | *07:78w, 13:16w, 13:20w, 13:48w, 15:42w, 49:01:01, 49:01:03-49:03, 49:06-49:10, 49:12-49:17, 49:19N-49:25, 51:62w, 51:106w, 52:25w, 54:20w, 55:01:01w-55:01:10w, 55:03w, 55:05w, 55:09w, 55:11w, 55:15w, 55:24w-55:25w, 55:29w, 55:31w, 55:33w, 55:36w, 55:38w, 55:40w, 55:44w-55:45w, 55:52w-55:55Nw, 56:25w, 56:37w |
| 87 | 135 bp, 365 bp, 540 bp | 1070 bp | *50:05, 50:17 | *13:01:01-13:01:06, 13:06, 13:12-13:13, 13:17, 13:20, 13:22:01-13:23, 13:25-13:26, 13:28-13:29, 13:39, 13:43, 13:50-13:52, 13:57, 13:60-13:61, 13:63N, 15:20, 15:25:01, 15:25:03, 15:36, 15:62, 15:85, 15:106, 15:204, 15:240, 15:250, 35:28, 35:50, 35:63, 35:84, 35:162, 35:217, 40:04, 40:28, 40:59, 40:64, 40:68, 40:120, 40:129, 40:137, 40:160:02, 44:02:07, 44:03:02, 44:03:13, 44:03:19, 44:07, 44:28:01, 44:37:02, 44:43:02, 44:48, 44:64:01, 44:124, 44:128, 44:143, 48:02:01-48:02:03, 48:17, 48:25, 53:17:02, 53:28 |
| 94,8 | 105 bp, 150 bp | 800 bp | *50:06, 50:12 | *40:208 |
| 104 | 90 bp | 800 bp | *50:07 | *40:83, 40:170, 44:07 |
| 114,9 | 105 bp, 280 bp | 1070 bp | *50:08, 50:16 | *54:28 |
| 1210 | 145 bp, 180 bp | 1070 bp | *50:09, 50:18-50:19 | *14:19, 14:27, 15:137, 45:02, 51:107, 53:07, C*03:110, C*04:08, C*04:34, C*05:27, C*05:39, C*08:41, C*17:07 |
| 13 | 270 bp | 800 bp | *50:10 |  |
| 14 | 300 bp | 1070 bp | *50:11 |  |
| 15 | 135 bp | 800 bp | *50:13, 50:19 | *15:46, 15:53, 15:106, 35:47, 35:154, 40:10:01-40:10:02, 40:40, 40:49, 40:58, 40:110, 40:129, 40:164, 40:210, 40:219, 44:02:01:01-44:03:15, 44:03:17-44:05:04, 44:09-44:14, 44:16-44:17, 44:19N-44:39, 44:41:01-44:43:02, 44:45-44:48, 44:50-44:54, 44:56N, 44:58N-44:59:02, 44:61N-44:80, 44:82, 44:84-44:102, 44:104-44:110, 44:112-44:129, 44:132-44:133, 44:135-44:144, 44:146-44:155, 44:157-44:165, 47:01:01:01-47:08, 49:10, 49:24, 57:45 |
| 164,11 | 95 bp, 160 bp, 580 bp | 1070 bp | *50:14-50:15 | *08:102, 13:36, 13:62, 18:80, 27:04:03, 27:23, 27:92, 35:01:10, 35:04:02, 35:84, 35:90, 35:150, 35:184, 35:201, 35:217, 38:16, 40:01:06, 40:08, 40:25-40:26, 40:28, 40:68, 40:95, 40:106, 40:166w, 44:12, 44:62, 49:18, 51:56:01-51:56:02, 51:103, 51:119, 53:28, 55:01:04, 56:36w, 73:01w-73:02w, C*02:02:23 |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-B*50 SSP typings.

When the primers in a primer mix can give rise to HLA-specific PCR products of more than one length this is indicated if the size difference is more than 20 base pairs. Size differences of 20 base pairs or less are not given. For high resolution SSP kits the respective lengths of the HLA-specific PCR product(s) are given for the alleles amplified by these primer mixes.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most wells, or a band of 800 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-B*50 SSP subtyping. In addition, wells number 3, 9, 10, 13 and 15 contain the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3Due to the sharing of sequence motifs between HLA-B alleles non-HLA-B*50 alleles will be amplified by primer mixes 1 to 12, 15 and 16. In addition, the A*02:339, C*03:12, C*03:19, C*03:102 and C*15:26 alleles will be amplified by primer mix 3, the C*02:02:23 will be amplified primer mix 16, and the C*03:110, C*04:08, C*04:34, C*05:27, C*05:39, C*08:41 and C*17:07 alleles will be amplified by primer mix 12.

4HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

5Primer mix 3 has a tendency of giving rise to primer oligomer formation.

6Primer mix 6 may give rise to a lower yield of HLA-specific PCR product than the other B*50 primer mixes.

7Primer mix 8: Specific PCR fragment of 135 bp in the B *50:05 allele. Specific PCR fragment of 365 bp in the B*13:01:01-13:01:06, 13:06, 13:12-13:13, 13:17, 13:20, 13:22:01-13:22:02, 13:25-13:26, 13:28-13:29, 13:39, 13:43, 13:50-13:52, 13:57, 13:60-13:61, 13:63N, 15:20, 15:25:01, 15:25:03, 15:36, 15:62, 15:85, 15:106, 15:240, 15:250, 35:28, 40:129, 40:137, 40:160:02, 44:02:07, 44:03:02, 44:03:13, 44:03:19, 44:07, 44:28:01, 44:37:02, 44:43:02, 44:48, 44:64:01, 44:124, 44:128, 44:143, 48:02:01-48:02:03, 48:17, 48:25, 53:17:02 and 53:28 alleles. Specific PCR fragment of 540 bp in the B*35:50, 35:84, 35:162, 35:217, 40:68 and 40:160:02 alleles. Specific PCR fragments of 365 and 540 bp in the B*50:17 and the B*13:23, 15:204, 35:63, 40:04, 40:28, 40:59, 40:64 and 40:120 alleles.

8Primer mix 9: Specific PCR fragment of 105 bp in the B*50:12 allele. Specific PCR fragment of 150 bp in the B*50:06 and the B*40:208 alleles.

9Primer mix 11: Specific PCR fragment of 105 bp in the B*50:16 allele. Specific PCR fragment of 280 bp in the B*50:08 and the B*54:28 alleles.

10Primer mix 12: Specific PCR fragment of 145 bp in the B*50:18 and 50:19 and in the B*51:107and the C*03:110 alleles. Specific PCR fragment of 180 bp in the B*50:09 and in the B*14:19, 14:27, 15:137, 45:02 and 53:07 and the C*04:08, C*04:34, C*05:27, C*05:39, C*08:41 and C*17:07 alleles.

11Primer mix 16: Specific PCR fragment of 95 bp in the B*35:01:10, 35:04:02, 51:56:01-51:56:02, 53:28 and 55:01:04 and the C*02:02:23 alleles. Specific PCR fragment of 160 bp in the B*50:15 and the B*08:102, 13:36, 18:80, 27:23, 27:92, 35:84, 35:90, 35:150, 35:184, 35:201, 35:217, 38:16, 40:08, 40:25, 40:68, 40:106, 40:166w, 44:12, 56:36w and 73:01w-73:02w alleles. Specific PCR fragment of 580 bp in the B*13:62, 27:04:03, 40:01:06, 40:26, 40:95 and 44:62, alleles. Specific PCR fragment of 95 and 580 bp in the B*50:14 and the B*40:28, 49:18, 51:103 and 51:119 alleles.

‘w’, may be weakly amplified.

1The B*5003 allele has been shown to be identical to B*50:02.

2HLA-B*50 alleles in bold lettering are listed as confirmed alleles on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.10.0, October 2012.

3The B*50:14 and 40:95 alleles will give rise to identical amplification patterns with the HLA-B*50 subtyping kit. These alleles can be distinguished by the HLA-B low resolution kit and/or the HLA-B*40 subtyping kit.

4Primer mix 8: Specific PCR fragment of 135 bp in the B *50:05 allele. Specific PCR fragment of 365 bp in the B*13:01:01-13:01:06, 13:06, 13:12-13:13, 13:17, 13:20, 13:22:01-13:22:02, 13:25-13:26, 13:28-13:29, 13:39, 13:43, 13:50-13:52, 13:57, 13:60-13:61, 13:63N, 15:20, 15:25:01, 15:25:03, 15:36, 15:62, 15:85, 15:106, 15:240, 15:250, 35:28, 40:129, 40:137, 40:160:02, 44:02:07, 44:03:02, 44:03:13, 44:03:19, 44:07, 44:28:01, 44:37:02, 44:43:02, 44:48, 44:64:01, 44:124, 44:128, 44:143, 48:02:01-48:02:03, 48:17, 48:25, 53:17:02 and 53:28 alleles. Specific PCR fragment of 540 bp in the B*35:50, 35:84, 35:162, 35:217, 40:68 and 40:160:02 alleles. Specific PCR fragments of 365 and 540 bp in the B*50:17 and the B*13:23, 15:204, 35:63, 40:04, 40:28, 40:59, 40:64 and 40:120 alleles.

150 bp in the B*50:06 and the B*40:208 alleles.

Primer mix 11: Specific PCR fragment of 105 bp in the B*50:16 allele. Specific PCR fragment of 280 bp in the B*50:08 and the B*54:28 alleles.

Primer mix 12: Specific PCR fragment of 145 bp in the B*50:18 and 50:19 and in the B*51:107and the C*03:110 alleles. Specific PCR fragment of 180 bp in the B*50:09 and in the B*14:19, 14:27, 15:137, 45:02 and 53:07 and the C*04:08, C*04:34, C*05:27, C*05:39, C*08:41 and C*17:07 alleles.

Primer mix 16: Specific PCR fragment of 95 bp in the B*35:01:10, 35:04:02, 51:56:01-51:56:02, 53:28 and 55:01:04 and the C*02:02:23 alleles. Specific PCR fragment of 160 bp in the B*50:15 and the B*08:102, 13:36, 18:80, 27:23, 27:92, 35:84, 35:90, 35:150, 35:184, 35:201, 35:217, 38:16, 40:08, 40:25, 40:68, 40:106, 40:166w, 44:12, 56:36w and 73:01w-73:02w alleles. Specific PCR fragment of 580 bp in the B*13:62, 27:04:03, 40:01:06, 40:26, 40:95 and 44:62, alleles. Specific PCR fragment of 95 and 580 bp in the B*50:14 and the B*40:28, 49:18, 51:103 and 51:119 alleles.

‘w’, may be weakly amplified.