Gel Documentation Form and Worksheet

HLA-B*50 (101.548-06/06u) 	Lot No: 70V	Expiry Date: 2016-October-01

Name:___________________

Sample ID:_______________

DNA Conc.(ng/ul):_________

Test  Date:____________

Tested By:____________

Review Date:_______________

Reviewed By:_______________

Gel Picture

PHOTO DOCUMENT

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon size

Interpretation:_______________

Failed lanes:  _______________

Comments:    ___________________

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mix 3 has a tendency of giving rise to primer oligomer formation.

Primer mixes 6 and 8 may give rise to a lower yield of HLA-specific PCR product than the other B*50 primer mixes.

Primer mix 21 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by control primer pairs. PCR product sizes range from 75 to 200 base pairs. The PCR product generated by the control primer pair is 430 base pairs.

1HLA-B*50 alleles in bold lettering are listed as confirmed alleles on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.15.0, January 2014.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3Primer mix 8: Specific PCR fragment of 135 bp in the B*50:05 allele. Specific PCR fragment of 365 bp in the B*13:01:01-13:01:08, 13:06, 13:12:01-13:13, 13:17, 13:20, 13:22:01-13:22:02, 13:25-13:26:02, 13:28-13:29, 13:39, 13:43, 13:50-13:52, 13:57, 13:60-13:61, 13:63N, 13:73, 15:20, 15:25:01, 15:25:03, 15:36, 15:62, 15:85, 15:106, 15:240, 15:250, 15:271, 15:289, 35:28, 40:129, 40:137, 44:02:07, 44:03:02, 44:03:13, 44:03:19, 44:07, 44:28:01, 44:37:02, 44:43:02, 44:48, 44:64:01, 44:124, 44:128, 44:143, 44:181, 48:02:01-48:02:03, 48:17, 48:25, 53:17:02 and 53:28 alleles. Specific PCR fragment of 540 bp in the B*35:84, 35:162, 35:217, 35:231, 40:68 and 58:41 alleles. Specific PCR fragments of 365 and 540 bp in the B*50:17 and the B*13:23, 15:204, 35:63, 40:04, 40:28, 40:59, 40:64, 40:120 and 40:160:02 alleles.

150 bp in the B*50:06 and the B*40:208 alleles.

Primer mix 9: Specific PCR fragment of 105 bp in the B*50:12 allele. Specific PCR fragment of 150 bp in the B*50:06 and the B*40:208 alleles.

Primer mix 11: Specific PCR fragment of 105 bp in the B*50:16 allele. Specific PCR fragment of 280 bp in the B*50:08 and the B*54:28 alleles.

Primer mix 12: Specific PCR fragment of 140 bp in the B*50:18 and 50:19 and in the B*44:186, 51:107and the C*03:110 alleles. Specific PCR fragment of 180 bp in the B*50:09 and in the B*14:19, 14:27, 15:137, 37:39, 45:02 and 53:07 and in the C*04:08, C*04:34, C*04:147, C*05:27, C*05:39, C*08:41 and C*17:07 alleles.

Primer mix 16: Specific PCR fragment of 95 bp in the B*35:01:10, 35:04:02, 40:28, 51:56:01-51:56:02, 53:28, 55:01:04 and in the C*02:02:23 alleles. Specific PCR fragment of 160 bp in the B*50:15 and the B*07:177w, 08:102, 13:36, 18:80, 27:23, 27:92, 35:84, 35:90, 35:150, 35:184, 35:201, 35:217, 38:16, 38:46, 40:08, 40:25, 40:68, 40:106, 40:166w, 40:232, 44:12, 49:26, 51:103, 51:119, 56:36w and 73:01w-73:02w alleles. Specific PCR fragment of 580 bp in the B*13:62, 27:04:03, 40:01:06, 40:26, 40:28, 40:95 and 44:62 alleles. Specific PCR fragment of 95 and 580 bp in the B*50:14 and the B*40:28 and 49:18 alleles. Specific PCR fragment of 160 and 580 bp in the B*51:103 and 51:119 alleles

‘w’, may be weakly amplified.

4The B*50:33 and the B*41:21alleles will give rise to identical amplification patterns with the HLA-B*50 subtyping kit. These alleles can be distinguished by the HLA-B low resolution kit and/or the HLA-B*41 subtyping kit.

The HLA-B*50 subtyping kit cannot distinguish the silent mutations in the B*50:01:01-50:01:04 alleles.

Changes in revision R01 compared to R00:

- Primer mix 8 does not amplify the B*35:50 allele. This primer mix may also give rise to a lower yield of HLA-specific PCR product than the other HLA-B*50 high resolution primer mixes. This has been corrected in the Specificity and Interpretation Tables.