PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® HLA-B*53

Prod. No:	101.549-06/06u

Lot No:    	08F

Expiry Date:    	2010-August-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 16 primer mixes used for HLA-B*53 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified HLA-B*53 alleles | Other amplified HLA Class I alleles3 |
| --- | --- | --- | --- | --- |
| 16 | 145 bp | 800 bp | 530101-5302, 5304-5308, 5310, 5314-5316 | 3806, 3807, 4406, 4903, 510101-512402, 5126-5146, 5148, 5149, 5151-5153, 5155-5157, 520101-520104, 5203-5214, 580101-5802, 5804-5816, 5818, 5819 |
| 2 | 270 bp | 1070 bp | 5302, 5306 | 1822, 3521, 3524, 3581, 3704, 4028, 5104, 5146, 5156, 5714, 5809 |
| 36, 7 | 135, 295 bp | 800 bp | 5303, 5315 | 0712, 0714, 071801-071802, 370101-3704, 3707-3709, 3712, 3713, 4030, 4034, 4059, 4817 |
| 44 | 180 bp | 800 bp | 5304, 5307 | 1303, 1558, 1573, 9533, 9537, 3503, 3506, 3513, 3536, 3538, 3555, 3556, 3559, 3570, 3574, 3575, 3584, 3585, 3932, 3933, 4071, 4410, 4415, 4418, 4501, 4502, 4504-4507, 4611, 4901-4903, 5001, 5002, 5004, 511301-511302, 5115, 5214, 5403, 5601, 5602, 5604, 5607, 5608, 5611, 5613, 5614, 5616, 5617, 5620, 5622, 8201, 8202, Cw*0408, Cw*0434, Cw*1515 |
| 56 | 140 bp | 800 bp | 5305, 5309, 5311-5313, 5316 | 1809, 3710, 5154, 5209 |
| 68 | 205, 235 bp | 800 bp | 5302, 5305, 5306, 5312 | 1538, 1587, 9515, 180101-1809, 1812-1815, 1817N-1819, 1820 weakly,1822-1825, 1827-1829, 3521, 3524, 3943, 4415, 5137, 5145, 5209, 5714, 5809 |
| 75,6 | 130 bp | 800 bp | 5307, 5310 | 1532, 1818, 370101-370104, 3703N-3706, 3708, 3710, 3711, 3713, 3714, 3909 |
| 8 | 165 bp | 1070 bp | 5306, 5308 | 070201-0724, 0726-0747, 0749N, 0750, 0752-0766, 0806, 0820, 1316, 1320, 1401-1409, 15010101-150104, 150106-1504, 1506-1519, 1521, 1523-1530, 1532-1540,1542-1547, 1549, 1550, 1553, 1554, 1556-1558, 1560-1574, 1576-1582, 1585, 1587, 1589, 1590, 1592-1599, 9501-9504, 9506, 9508-9510, 9512, 9513, 9515-9522, 9525-9529, 9531-9535, 9537-9544, 9546-9549N, 1815, 1819, 1821, 270401-270402, 2706, 2710, 2715, 2718, 2720, 2721, 2724, 2725, 2740, 3511, 351401-351402, 3521, 3543, 3544, 3558, 3567, 3579, 3586, 3707, 3713, 3810, 3918, 3935, 3936, 4005, 4015, 4016, 4023, 4026, 4028, 4032, 4051, 460101-4605, 4607N, 4608, 4610, 4612, 4614-4616, 4805, 4808, 4815, 4901-4905, 5001, 5002, 5004, 510101-5104, 5106, 5107, 5111N-5114, 5116-5118, 5121-512402, 5126-5130, 5132-5139, 5141N, 5143, 5145, 5146, 5148-5152, 5155-5157, 520101-5202, 5204-5209, 5211-5214, 5406, 550101-550104, 5503, 5505, 5509, 5511, 5514, 5515, 5517, 5521, 5524, 5525, 5528, 5529, 5531, 5603, 560501-5606, 5621, 5709, 5713, 5806, 5808, 5819, 7801-7803, 7805, 7806 |
| 96 | 75 bp | 1070 bp | 530101-5305, 5307, 5309-5316 | 0821, 0825, 1303, 1304, 1505, 1520, 1531, 1548, 1551, 1555, 1584, 1591, 9507, 9523, 9536, 2738, 350101-3510, 3512, 3513, 3516-352002, 3522-3530, 3532, 3534, 3536-3542, 3545-3557, 3559, 3561-3564, 356801-3572, 3574-3578, 3580-3585, 3590-3592, 3594, 3932, 3933, 4063, 4071, 4092, 44020101-440303, 4405-4415, 4417-4420, 4422-4430, 4432-4455, 4457-4463, 4501-4507, 4509, 4609, 4611, 4613, 4802, 5105, 5108-510902, 5115, 5119, 5120, 5140, 5142, 5144N, 5154, 5203, 5210, 5403, 5508, 5601, 5602, 5604, 5607-5609, 5611, 5613-5617, 5620, 5622, 570101-570102, 5702-5708, 5710-5712, 5714-5717, 580101-5802, 5804, 5807, 5809-5818, 7804, 8201, 8202, 8301, Cw*0336, Cw*0434 |
| 10 | 150 bp | 1070 bp | 530101-5303, 5305, 5309, 5310, 5311weakly, 5312-5316 | 1304, 1310 weakly, 1505, 1520, 1531 weakly, 1552, 1555 weakly,1584weakly, 1591, 9507, 9514, 9523, 9524, 9536, 180101-1812, 1813weakly, 1814, 1817N, 1818, 1820 weakly, 1822-1825, 1827-1829, 2741, 350101-350107, 3505, 3507, 350801-350803weakly, 3510, 3515-3517, 3519-352002, 3523-3530, 3532, 3535, 3537, 3540N-3542, 3545weakly, 3546-3554, 3557, 3561 weakly, 3563, 3564, 3566weakly, 356801-3569, 3571, 3572, 3576-3578, 3580, 3582, 3590-3594, 3907, 3943, 4003, 4020, 4038, 4052, 4059, 4060 weakly , 4443, 4606, 4613, 4802, 4814, 5409, 5414, 5523, 5532, 5609, 5618, 570101-570103, 5706, 5708, 5710, 5711, 5714-5716, 580101-5802, 5804, 5807, 5809-5818 |
| 116 | 85 bp | 800 bp | 530101-5303, 5305, 5306, 5308-5316 | 1304, 1310 weakly, 15010101-150104, 150106-150108, 1502weakly, 1503-1505, 1506 weakly, 1507, 1508, 151101-1512, 1513weakly, 1514-1516, 1518-1520, 1521 weakly, 1523, 1524, 1525 weakly, 1526N-1529, 1531 weakly, 1532, 1534, 1535, 1536weakly, 1538-1540, 1543, 1544 weakly, 1546, 1547, 1549-1554, 1555 weakly, 1556, 1557, 1560-1562, 1564-1571, 1574-1576, 1579N-1582, 1584, 1585, 1587, 1588 weakly, 1589 weakly, 1591-1598, 9501-9505, 9506 weakly, 9508-9511N, 9512 weakly, 9513-9515, 9517-9520, 9521weakly, 9522-9527, 9528 weakly, 9529weakly, 9531, 9532, 9534-9536, 9538-9540, 9542-9549N, 180101-1815, 1817N-1819, 1820 weakly, 1822-1825, 1827-1829, 2741, 350101-350107, 3505, 3507-350803, 3510, 3511, 351401-3517, 3519-3521, 3523-3530, 3532, 3535, 3537, 3540N-3543, 3545-3552, 3554, 3557, 3558, 3561-3563, 3566-3569, 3571, 3572, 3576-3580, 3582, 3586, 3590-3594, 3907, 3943, 4417, 4443, 4509, 460101-4605, 4607N-4610, 4612-4616, 4802, 4814, 4904, 4905 weakly, 5137, 5145, 5406, 5409, 5414, 5514, 5523, 5532, 5603, 5609, 5618, 570101-570103, 5706, 5708, 5710, 5711, 5714-5716, 580101-5802, 5804-5807, 5809-5819 |
| 126 | 105 bp | 1070 bp | - | 130201-1303, 1308Q, 1309, 1314-1316, 1318, 1319, 1542, 3560, 4415, 4418, 4501, 4503-4508, 4611, 4901-4903, 5001, 5002, 5004, 5115, 5401-5403, 5405N, 5407, 5408N, 5410, 5412, 5413, 550101-5503, 5505, 5507, 5509-5512, 5515, 5516, 5518, 5519, 5521, 5522, 5524-5526, 5529-5531, 5601, 5607, 5608, 5613, 5614, 5616, 5617, 5619N, 5620, 5623, 5901 |
| 136 | 90bp | 1070 bp | - | 580101-5802, 5804-5819 |
| 146, 9 | 95, 130 bp | 800 bp | 5310, 5314 | 0709, 0711, 0717, 0828, 0835, 1405, 1507, 1555, 1568, 9526, 1814, 3505, 3551, 3558, 3566, 3572, 3903, 3924, 3937, 4003, 4206, 4209, 4612, 4807, 4814, 5818, Cw*0808 |
| 156 | 105 bp | 1070 bp | 5309 | 180101-1806, 1808-1815, 1817N-1828, 352001-352002, 3528, 3572, 370101-3705, 3707-3714, 4903, 5107, 520101-520104, 5203-5205, 5207-5214, 7805, 7806 |
| 166 | 115 bp | 1070 bp | 530101-5310, 5312-5315 | 0726, 0734, 0743, 0748, 0813, 0825, 1301-1304, 1306-1317, 1319-1322, 1401-140202, 1404-1409, 1502, 1503, 1505, 1506, 1509, 1510, 1513, 1516-1518, 1520, 1521, 1523, 1525, 1529, 1531, 1536, 1537, 1539, 1540, 1542, 1544, 1548, 1552, 1555, 1561, 1562, 1564, 1567, 1569, 1572, 1574, 1580, 1586, 1588-1591, 1593, 1595, 1598, 9503, 9506-9508, 9512, 9514, 9515, 9519, 9521, 9523, 9524, 9527, 9531-9534, 9536, 9538, 9539, 180101-1812, 1814, 1815, 1817N, 1818, 1820, 1822-1829, 2701-270402, 270502-2721, 2723, 2724, 2726-2743, 350101-3507, 350901-3513, 3515-3517, 3519-3537, 3539-3542, 3546-3560, 3563-3565Q, 356801-3572, 3574-3578, 3581-3585, 3588, 3590-3594, 3702, 380101-3817, 39010101-39010102L, 390103-3907, 3909, 3910, 3912-3917, 391901-3920, 3922-3932, 3934, 3935, 3937-3945, 400101-401403, 4018-4022N, 4024-4031, 4033-4040, 4042-4092, 440301-440303, 4407, 4410, 4413, 4426, 4429-4432, 4436-4440, 4443, 4445-4447, 4450, 4454, 4457, 4460, 4461N, 446401-446402, 4606, 4608, 4611, 4613, 47010101-4705, 480101-4804, 4806, 4807, 4809-4818, 4901-4905, 5001, 5002, 5004, 510101-5104, 5106, 5107, 510901-5119, 5121-512402, 5126-5128, 5130-5135, 5137-5143, 5145, 5146, 5148-5153, 5156, 5157, 520101-5214, 5401-5405N, 5407-5409, 5411-5415, 550101-5505, 5507, 5508, 5510-5519, 5522-5532, 5601, 5602, 5604-5612, 5614-5623, 570101-570103, 570301-570302, 5706-5708, 5710, 5711, 5714-5717, 580101-5802, 5804-5819, 5901- 5903, 670101-6702, 7301, 7801-7806, 8101-8104N |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-B*53 SSP typings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective lengths of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may  be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most wells, or a band of 800 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA- B*53 subtyping.

In addition, wells number 3 to 7, 11 and 14 contain the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3Due to the sharing of sequence motifs between HLA-B alleles non-HLA-B*53 alleles will be amplified by all 16 primer mixes. In addition, the Cw*0336 and Cw*0434 allele will be amplified by primer mix 9, the Cw*0408, the Cw*0434 and Cw*1515 allelles will be amplified by primer mix 4, and the Cw*0808 allele will be amplified by primer mix 14.

4Primer mix 3 may give rise to a specific PCR fragment longer than 500 bp, that should be disregarded in the interpretation of HLA-B*53 subtypings.

5Primer mix 7 has a tendency of giving rise to nonspecific amplifications and also to primer oligomer formation.

6Short specific PCR fragments are less intense and not as sharp as longer specific bands.

7Primer mix 3: Specific PCR fragment of 135 bp in the B*5303 and 370101-3704, 3707-3709, 3712 and 3713 alleles. Specific PCR fragment of 295 bp in the B*5315 and the 0712, 0714, 071801-071802, 4030, 4034, 4059 and 4817 alleles.

8Primer mix 6: Specific PCR fragment of 205 bp in the B*5302 and 5306 and the 1538, 180101-1809, 1812-1815, 1817N-1819, 1820weakly, 1822-1825, 1827-1829, 3521, 3524, 3943, 5137, 5145, 5714 and 5809 alleles. Specific PCR fragment of 235 bp in the the B*5305, 5312 and the 1587, 9515, 5209, 4415 and 5209 alleles.

9Primer mix 14: Specific PCR fragment of 95 bp in the B*5314, 0709, 0711, 0717, 0828, 0835, 1405, 1507, 1555, 1568, 9526, 1814, 3505, 3551, 3558, 3566, 3572, 3903, 3924, 3937, 4003, 4206, 4209, 4612, 4807, 4814, 5818 and the Cw*0808 alleles. Specific PCR fragment of 130 bp in the B*5310 allele.

1Primer mix 3: Specific PCR fragment of 135 bp in the B*5303 and 370101-3704, 3707-3709, 3712 and 3713 alleles. Specific PCR fragment of 295 bp in the B*5315 and the 0712, 0714, 071801-071802, 4030, 4034, 4059 and 4817 alleles.

8Primer mix 6: Specific PCR fragment of 205 bp in the B*5302 and 5306 and the 1538, 180101-1809, 1812-1815, 1817N-1819, 1820weakly, 1822-1825, 1827-1829, 3521, 3524, 3943, 5137, 5145, 5714 and 5809 alleles. Specific PCR fragment of 235 bp in the the B*5305, 5312 and the 1587, 9515, 5209, 4415 and 5209 alleles.

Primer mix 14: Specific PCR fragment of 95 bp in the B*5314, the 0709, 0711, 0717, 0828, 0835, 1405, 1507, 1555, 1568, 9526, 1814, 3505, 3551, 3558, 3566, 3572, 3903, 3924, 3937, 4003, 4206, 4209, 4612, 4807, 4814, 5818 and the Cw*0808 alleles. Specific PCR fragment of 130 bp in the B*5310 allele.

‘w’, might be weakly amplified.