PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® HLA-B*56

Prod. No:	101.571-06/06u

Lot No:    	85R

Expiry Date:    	2015-July-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 24 primer mixes used for HLA-B*56 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified HLA-B*56 alleles3 | Other amplified HLA Class I alleles4 |
| --- | --- | --- | --- | --- |
| 15 | 125 bp | 800 bp | *56:01:01-56:02, 56:07-56:11, 56:13-56:17, 56:19N, 56:22-56:26, 56:28N-56:30, 56:33-56:38N | *13:01:01-13:04, 13:06-13:09, 13:11-13:20, 13:22:01-13:23, 13:25-13:34, 13:36-13:61, 13:63N, 14:22, 15:16:01-15:16:03, 15:20, 15:34, 15:42, 15:62, 15:67, 15:85-15:86, 15:95, 15:137, 15:194, 15:222, 15:254, 18:22, 18:52, 27:14, 35:01:01:01-35:05:02, 35:06-35:08:06, 35:10-35:11:03, 35:13-35:15, 35:19-35:21, 35:23-35:30, 35:33-35:38, 35:40N-35:42:02, 35:45-35:50, 35:52, 35:54-35:58, 35:60, 35:62-35:63, 35:65Q-35:66, 35:68:01-35:72, 35:74-35:78, 35:80-35:85, 35:89-35:98, 35:100-35:101:02, 35:103-35:112, 35:115-35:116, 35:119-35:124, 35:126-35:134N, 35:136-35:148, 35:150-35:163, 35:166-35:173N, 35:175-35:178, 35:180-35:184, 35:186-35:195, 35:197-35:198, 35:200, 35:202-35:204, 35:206-35:207, 35:209-35:214, 35:216N-35:220, 40:28, 44:10, 44:15, 44:18, 44:48, 45:01-45:14, 46:11, 46:18, 46:33, 48:02:01-48:02:03, 48:25, 49:01:01-49:25, 50:01:01, 50:01:03-50:02, 50:05-50:19, 51:15, 51:56:01-51:56:02, 53:01:01-53:06, 53:08:01-53:29, 54:01:01-54:05N, 54:07-54:08N, 54:10, 54:12-54:28, 55:01:01-55:01:05, 55:01:07-55:02:03, 55:02:05-55:03, 55:05, 55:07, 55:09-55:13, 55:15-55:16, 55:18-55:19, 55:21-55:37, 55:39-55:48, 55:50, 55:52-55:57, 57:05, 58:01:01-58:01:02, 58:01:04-58:02, 58:04-58:07, 58:09, 58:11-58:19, 58:21-58:26, 58:28-58:29, 58:31N-58:39N, 59:01:01:01-59:01:01:02, 59:03-59:05, 82:01-82:03, C*02:56, C*06:20, C*07:81, C*07:168 |
| 2 | 400 bp | 1070 bp | *56:01:01-56:13, 56:15-56:32, 56:34-56:38N | *07:65, 13:54, 27:77, 35:76, 35:78, 40:166?, 44:90, 54:01:01-54:06, 54:08N-54:23, 54:25-54:28, 55:01:01-55:05, 55:07-55:17, 55:19-55:33, 55:35-55:46, 55:48-55:57, 81:01, 81:03?-81:04N?, 81:06?, 82:01-82:03, 83:01 |
| 320 | 180 bp | 1070 bp | *56:01:01-56:02, 56:04, 56:07-56:08, 56:13-56:14, 56:16-56:17, 56:19N-56:20:02, 56:24-56:30, 56:33-56:38N | *13:03, 13:48, 15:73, 40:71, 44:10, 44:15, 44:18, 44:140, 45:01, 45:04-45:07, 45:11-45:14, 46:11, 49:01:01-49:03, 49:06, 49:08-49:17, 49:19N-49:25, 50:01:01-50:02, 50:04-50:08, 50:10-50:13, 50:15-50:19, 51:15, 51:62, 51:106, 52:25, 54:03, 59:04, 82:01-82:03 |
| 45 | 105 bp | 800 bp | *56:02, 56:04, 56:10 | *13:11, 15:18:03, 15:73, 15:86, 15:224, 39:17, 40:48, 40:71, 51:62, 82:01-82:03, C*02:06, C*02:47, C*07:289, C*12:15, C*15:37, C*15:58 |
| 5 | 285 bp | 1070 bp | *56:03 | *40:10:02, 54:01:02, 55:35 |
| 65,9 | 105 bp, 150 bp | 800 bp | *56:05:01, 56:32 | *07:09w, 07:11w, 07:162w, 15:03:01-15:03:03, 15:16:01-15:16:03, 15:18:01-15:18:05, 15:23, 15:29, 15:39:01-15:40, 15:47:01w-15:47:02w, 15:49w, 15:61-15:62, 15:64:01-15:64:02, 15:65w, 15:67, 15:69, 15:72, 15:74, 15:80, 15:93, 15:95, 15:98, 15:103, 15:108, 15:115, 15:119, 15:127, 15:131-15:132, 15:134, 15:138, 15:153, 15:158, 15:161, 15:173, 15:176, 15:194-15:195, 15:197-15:198, 15:210, 15:220, 15:221w, 15:223, 15:226N, 15:238, 15:242, 15:254-15:255, 15:263, 18:15, 18:21w, 18:30w, 18:57, 27:05:18, 35:11:01-35:11:02, 35:11:03w, 35:21, 35:58, 35:99w, 35:188, 39:63, 46:03w, 46:08, 48:25, 49:04, 51:37, 51:45, 53:06, 53:08:01-53:08:02, 53:28, 55:14, 57:13w, 57:22, 58:06, 58:19, C*03:175, C*15:09, C*15:19, C*15:27, C*15:61 |
| 75,6,10 | 120 bp, 170 bp | 800 bp | *56:05:01-56:06, 56:21, 56:33 | *13:62, 15:21, 15:44, 15:66, 15:189, 35:26, 35:82, 35:85, 35:135, 38:28, 40:95, 49:18, 50:14, 51:01:01-51:01:03, 51:01:05-51:01:12, 51:01:14-51:02:03, 51:02:05-51:03, 51:05, 51:07:01-51:09:02, 51:11N-51:12, 51:13:02-51:14, 51:16-51:23, 51:24:03-51:24:04, 51:26-51:37, 51:39-51:41N, 51:43-51:44N, 51:48-51:55, 51:57-51:58, 51:60-51:61:02, 51:63, 51:65-51:80, 51:82-51:91, 51:93-51:134, 51:136-51:138, 51:140-51:147, 52:01:01:01-52:01:04, 52:01:06-52:01:14, 52:01:16-52:13, 52:15-52:24, 52:26-52:31, 55:01:04, 58:08:01-58:08:02, 78:01:01-78:07 |
| 8 | 395 bp | 1070 bp | *56:08, 56:14 | *35:186 |
| 911 | 145 bp, 205 bp | 800 bp | *56:07, 56:31 | *07:149, 08:02, 15:38:01-15:38:02, 15:115, 15:185, 18:01:01:01-18:01:13, 18:01:15-18:09, 18:12-18:15, 18:17N-18:19, 18:22-18:25, 18:27-18:34, 18:36-18:40, 18:42-18:55, 18:57, 18:59-18:60, 18:62-18:83, 27:01, 35:21, 35:24:01-35:24:02, 35:188, 35:190, 38:02:01-38:04, 38:08, 38:15, 38:18, 38:23, 38:29, 38:35, 39:43, 44:40, 44:44, 44:130, 44:156, 51:37, 51:45, 53:02, 53:06, 53:28, 57:14, 58:09 |
| 107 | 150 bp | 800 bp | *56:09, 56:11-56:12 | *08:49, 13:01:01-13:01:06, 13:06-13:07N, 13:12-13:13, 13:17, 13:20, 13:22:01-13:23, 13:25-13:26, 13:28-13:29, 13:36, 13:39, 13:43, 13:50-13:52, 13:57, 13:60-13:61, 13:63N, 14:10, 15:02:01-15:02:06, 15:13:01-15:13:02, 15:20-15:21, 15:25:01-15:25:03, 15:36, 15:44, 15:62, 15:77, 15:80, 15:85, 15:88-15:89, 15:106, 15:112, 15:121, 15:139, 15:144, 15:154, 15:165, 15:170, 15:194, 15:204, 15:213-15:214, 15:223, 15:240, 15:250, 18:22, 18:69, 27:19, 27:30, 35:01:01:01-35:04:03, 35:06-35:08:06, 35:10-35:17, 35:19-35:21, 35:23-35:30, 35:33-35:36, 35:38-35:42:02, 35:45-35:50, 35:52, 35:54-35:57, 35:59, 35:61:01-35:63, 35:65Q, 35:69-35:71, 35:74, 35:76-35:78, 35:80-35:85, 35:90-35:96, 35:98, 35:100-35:101:02, 35:103-35:113, 35:115-35:116, 35:120-35:126, 35:128-35:134N, 35:136-35:150, 35:152-35:173N, 35:175-35:184, 35:186-35:198, 35:200-35:204, 35:206-35:220, 37:01:01-37:01:07, 37:01:09, 37:03N-37:06, 37:08, 37:10-37:11, 37:13-37:37, 38:20, 39:42, 40:28, 44:02:01:01-44:14, 44:16-44:17, 44:19N, 44:21-44:30, 44:32-44:40, 44:42-44:46, 44:48-44:52N, 44:55-44:64:02, 44:66-44:98, 44:101-44:105, 44:107-44:134, 44:136-44:137, 44:139-44:157, 44:159-44:165, 46:33, 48:02:01-48:02:03, 48:25, 50:17, 51:04, 51:42, 51:46, 51:56:01-51:56:02, 51:139, 53:01:01-53:13, 53:15-53:29, 55:14, 57:01:01-57:01:04, 57:01:06-57:10, 57:12, 57:14-57:20, 57:22-57:30, 57:32-57:46, 57:48-57:60, 58:01:01-58:01:02, 58:01:04-58:01:12, 58:04-58:05, 58:09-58:15, 58:17N, 58:19, 58:21-58:24, 58:28-58:29, 58:31N-58:37, 58:39N, 83:01, C*03:102 |
| 115 | 105 bp | 800 bp | *56:01:01-56:01:04, 56:07-56:08, 56:13-56:14, 56:16-56:17, 56:19N-56:20:02, 56:23-56:30, 56:33-56:38N | *07:78, 13:02:01-13:02:14, 13:03, 13:08Q-13:09, 13:14-13:16, 13:18-13:19, 13:27, 13:30-13:34, 13:37-13:38, 13:40-13:42, 13:44-13:45, 13:48-13:49N, 13:53-13:56N, 13:58, 15:42, 35:60, 44:15, 44:18, 45:01, 45:03-45:08, 45:10-45:14, 46:11, 46:18, 49:01:01-49:03, 49:06-49:17, 49:19N-49:25, 50:01:01-50:02, 50:04-50:08, 50:10-50:11, 50:13, 50:15, 50:18-50:19, 51:15, 52:25, 54:01:01-54:03, 54:05N, 54:07-54:08N, 54:10, 54:12-54:13, 54:16-54:28, 55:01:01-55:01:03, 55:01:05-55:01:06, 55:01:08-55:03, 55:05, 55:07, 55:09-55:12, 55:15-55:16, 55:18-55:19, 55:21-55:22, 55:24-55:26, 55:29-55:31, 55:33-55:38, 55:40-55:41, 55:43, 55:45-55:48, 55:50, 55:52, 55:54-55:57, 59:01:01:01-59:01:01:02, 59:04-59:05 |
| 127,12 | 150 bp, 180 bp | 800 bp | *56:10, 56:23, 56:25, 56:37 | *07:78, 13:06, 13:16, 13:20, 13:31w, 13:48, 13:53, 15:42, 15:73w, 15:86, 15:224, 35:60, 46:18, 51:62, 51:106, 52:25, 54:01:01-54:02, 54:05N, 54:07-54:08N, 54:10, 54:12-54:13, 54:17-54:25, 54:27-54:28, 55:01:01-55:03, 55:05, 55:07, 55:09-55:12, 55:15-55:16, 55:18-55:19, 55:21, 55:24-55:26, 55:29-55:31, 55:33-55:40, 55:42-55:45, 55:47-55:48, 55:50, 55:52-55:57, 59:01:01:01-59:01:01:02, 59:05 |
| 135,7,13 | 70 bp, 180 bp | 1070 bp | *56:11, 56:22, 56:30 | *15:58, 15:133, 15:229, 15:253, 35:03:01-35:03:05, 35:03:07-35:03:12, 35:06, 35:13, 35:36, 35:38, 35:55-35:56, 35:59, 35:70, 35:74-35:75, 35:84-35:85, 35:98, 35:106, 35:109, 35:127-35:128, 35:136, 35:150-35:153, 35:155-35:156, 35:160, 35:163, 35:167, 35:169, 35:179, 35:181, 35:193, 35:195, 35:198, 35:204-35:205, 38:30, 39:32-39:33, 39:74, 51:13:01-51:13:02, 51:92, 52:14, 53:04, 58:34 |
| 145,8 | 80 bp | 800 bp | *56:04, 56:27 | *15:18:03, 15:73, 15:224, 39:17, 40:48, 40:71, 44:140, 50:04, 51:62, 55:38, A*02:339w, C*03:12w, C*03:19w, C*03:102w, C*15:26w |
| 1514 | 160 bp, 380 bp | 1070 bp | *56:13, 56:28N | *44:15, 44:18, 44:140, 45:01, 45:04-45:08, 45:10-45:14, 55:56, 82:01-82:03 |
| 16 | 160 bp | 800 bp | *56:36 | *08:87, 40:166, 44:02:01:01-44:02:26, 44:06, 44:08-44:09, 44:11-44:12, 44:16, 44:19N-44:24, 44:27:01-44:27:02, 44:33-44:34:02, 44:41:01-44:41:02, 44:44, 44:48-44:49, 44:52N-44:53, 44:55, 44:59:01-44:59:02, 44:63, 44:66-44:68, 44:71-44:74, 44:80, 44:83-44:84, 44:86-44:87, 44:89-44:91, 44:93, 44:95, 44:97, 44:99-44:102, 44:104, 44:106, 44:112-44:113, 44:116, 44:118-44:119, 44:121, 44:126:01-44:127, 44:131-44:132, 44:137-44:138Q, 44:142Q, 44:145, 44:148-44:149N, 44:151-44:152, 44:162, 45:06, 83:01 |
| 175,7,15 | 80 bp, 180 bp | 1070 bp | *56:05:01-56:06, 56:15, 56:21, 56:38N | *07:20, 07:24, 07:60, 07:100, 07:131, 08:21, 08:25, 13:62, 15:09-15:10:02, 15:30, 15:37, 15:45, 15:48, 15:63, 15:90, 15:99, 15:150, 15:243, 15:248, 27:76, 35:02:01-35:02:05, 35:04:01-35:04:03, 35:09:01-35:09:03, 35:12:01-35:12:03, 35:18, 35:22, 35:34, 35:39, 35:44, 35:81, 35:83, 35:95, 35:129N, 35:149, 35:154, 35:162, 35:172, 35:182-35:184, 35:199, 35:201, 35:211, 35:220, 40:05, 40:26, 40:28, 40:63, 40:92, 40:174, 41:17, 44:62, 44:77, 44:82, 44:107, 44:123, 49:18, 50:14, 51:01:01-51:01:14, 51:01:16-51:01:31, 51:01:33-51:09:02, 51:11N-51:12, 51:18-51:20, 51:22-51:24:04, 51:26-51:30, 51:32-51:33, 51:35, 51:38-51:41N, 51:43-51:44N, 51:46, 51:48-51:61:02, 51:64-51:80, 51:83-51:84, 51:86-51:91, 51:94-51:96, 51:98N-51:100, 51:102-51:105, 51:107-51:142, 51:145-51:148, 52:01:01:01-52:07, 52:09-52:13, 52:15-52:24, 52:27-52:31, 53:19, 55:08, 55:51, 57:02:01-57:03:02, 57:05, 57:07, 57:09, 57:12, 57:17, 57:28N, 57:39, 57:42, 57:46, 57:57, 58:08:01-58:08:02, 58:28, 78:01:01-78:06, C*03:81, C*03:152 |
| 186 | 295 bp | 800 bp | *56:16 | *40:149, 41:22, 49:25, 54:21, 55:10, 57:58, 82:01-82:03 |
| 195,16 | 75 bp, 185 bp | 800 bp | *56:19N, 56:29 | *07:176, 39:40:01N-39:40:02N |
| 20 | 215 bp | 1070 bp |  | *07:02:19, 15:01:04, 37:01:04, 40:94, 51:01:18, 54:01:01, 54:01:03-54:28, 58:01:08 |
| 215,17 | 110 bp, 170 bp | 1070 bp | *56:17, 56:23, 56:26 | *07:37, 27:77 |
| 225,18 | 105 bp, 175 bp | 800 bp | *56:06, 56:18, 56:31-56:32, 56:36 | *07:65, 07:134, 08:37, 08:89, 15:50, 15:69, 15:93, 15:186, 15:199, 18:01:01:01-18:01:13, 18:01:15-18:01:17, 18:02-18:09, 18:11-18:15, 18:17N-18:19, 18:22-18:25, 18:27-18:34, 18:37-18:40, 18:42-18:49, 18:51-18:55, 18:59-18:60, 18:62-18:83, 35:35, 35:76, 35:93, 35:188, 35:190, 37:35, 39:07, 39:43, 40:166, 42:09, 44:90, 45:06, 46:03, 46:21:01-46:21:02, 46:29, 54:09, 54:14, 55:14, 55:23, 57:26, 58:05, 58:35, C*06:89, C*07:123, C*07:173, C*15:09, C*15:19, C*15:27, C*15:61, C*16:20 |
| 235,19 | 85 bp, 350 bp | 1070 bp | *56:20:01-56:20:02, 56:24 | *07:78, 40:48 |
| 247 | 380 bp | 1070 bp | *56:21 | *08:03, 08:52, 08:78, 40:13, 40:19, 40:109, 40:117, 48:18, 49:18, 51:01:01-51:01:07, 51:01:09-51:01:14, 51:01:16-51:12, 51:16-51:24:04, 51:26-51:36, 51:38-51:41N, 51:43-51:44N, 51:46, 51:48-51:49, 51:51-51:53, 51:55-51:61:02, 51:64-51:77, 51:79-51:84, 51:86-51:89, 51:91, 51:93-51:96, 51:98N-51:105, 51:107-51:134, 51:136-51:148, 52:01:01:01-52:11, 52:13, 52:15-52:19, 52:21-52:24, 52:27-52:31, 53:19, 57:02:01-57:03:02, 57:05, 57:07, 57:09, 57:17, 57:39, 57:42, 57:46, 57:57, 58:08:01-58:08:02, 58:27-58:28, 59:02, A*24:146 |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA B*56 high resolution SSP typings.

When the primers in a primer mix can give rise to HLA-specific PCR products of more than one length this is indicated if the size difference is more than 20 base pairs. Size differences of 20 base pairs or less are not given. For high resolution SSP kits the respective lengths of the HLA-specific PCR product(s) are given for the alleles amplified by these primer mixes.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may  be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1 070 base pairs, for most wells, or a band of 800 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-B*56 subtyping.

In addition, wells number 4, 6, 7, 9 to 12, 14, 16, 18, 19 and 22 contain the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3For several HLA Class I alleles 1st and/or 4th exon(s) and beyond, as well as intron nucleotide sequences, are not available. In these instances it is not known whether some of the primers of the SSP sets are completely matched with the target sequences or not. We assume that unknown sequences in these regions are conserved within allelic groups.

4Due to the sharing of sequence motifs between HLA-B alleles non-HLA-B*56 alleles will be amplified by all primer mixes. In addition, the A*02:339, C*03:12, C*03:19 and C*15:26 alleles will be weakly amplified by primer mix 14, the A*24:146 allele will be amplified by primer mix 24, the C*02:06, C*02:47, C*07:289, C*12:15, C*15:37 and C*15:58 alleles will be amplified by primer mix 4, the C*02:56, C*06:20, C*07:81 and C*07:168 alleles will be amplified by primer mix 1, the C*03:81 and C*03:152 alleles will be amplified by primer mix 17, the C*03:102 allele by primer mix 10 and weakly by primer mix 14, the C*03:175 allele will be amplified by primer mix 6, the C*06:89, C*07:123, C*07:173 and C*16:20 alleles will be amplified by primer mix 22 and the C*15:09, C*15:19, C*15:27 and C*15:61 alleles will be amplified by primer mixes 6 and 22.

5HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

6Primer mix 7 and 18 may give rise to a lower yield of HLA-specific PCR product than the other B*56 primer mixes.

7Primer mixes 10, 12, 13, 17 and 24 may have tendencies of unspecific amplifications.

8Primer mixes 14 has a tendency to giving rise to primer oligomer formation.

9Primer mix 6: Specific PCR fragment of 105 bp in the B*56:05:01 and the B*27:05:18 and 39:63 alleles. Specific PCR fragment of 150 bp in the B*56:32 and in the B*07:09w, 07:11w, 07:162w, 15:03:01-15:03:03, 15:16:01-15:16:03, 15:18:01-15:18:05, 15:23, 15:29, 15:39:01-15:40, 15:47:01w-15:47:02w, 15:49w, 15:61-15:62, 15:64:01-15:64:02, 15:65w, 15:67, 15:69, 15:72, 15:74, 15:80, 15:93, 15:95, 15:98, 15:103, 15:108, 15:115, 15:119, 15:127, 15:131-15:132, 15:134, 15:138, 15:153, 15:158, 15:161, 15:173, 15:176, 15:194-15:195, 15:197-15:198, 15:210, 15:220, 15:221w, 15:223, 15:226N, 15:238, 15:242, 15:254-15:255, 15:263, 18:15, 18:21w, 18:30w, 18:57, 35:11:01-35:11:02, 35:11:03w, 35:21, 35:58, 35:99w, 35:188, 46:03w, 46:08, 48:25, 49:04, 51:37, 51:45, 53:06, 53:08:01-53:08:02, 53:28, 55:14, 57:13w, 57:22, 58:06 and 58:19 and the C*03:175, C*15:09, 15:19, 15:27, 15:61alleles.

10Primer mix 7: Specific PCR fragment of 120 bp in the B*56:05:01-56:06 and 56:21 and the B*13:62, 40:95, 49:18, 50:14, 51:01:01-51:01:03, 51:01:05-51:01:12, 51:01:14-51:02:03, 51:02:05-51:03, 51:05, 51:07:01-51:09:02, 51:11N-51:12, 51:13:02-51:14, 51:16-51:23, 51:24:03-51:24:04, 51:26-51:37, 51:39-51:41N, 51:43-51:44N, 51:48-51:55, 51:57-51:58, 51:60-51:61:02, 51:63, 51:65-51:80, 51:82-51:91, 51:93-51:134, 51:136-51:138, 51:140-51:147, 52:01:01:01-52:01:04, 52:01:06-52:01:14, 52:01:16-52:13, 52:15-52:24, 52:26-52:31, 55:01:04, 58:08:01-58:08:02 and 78:01:01-78:07 alleles. Specific PCR fragment of 170 bp in the B*56:33 and the B*15:21, 15:44, 15:66, 15:189, 35:26, 35:82, 35:85, 35:135 and 38:28 alleles.

11Primer mix 9: Specific PCR fragment of 145 bp in the B*56:07 and the B*07:149, 08:02, 15:115, 27:01, 38:02:01-38:04, 38:08, 38:15, 38:18, 38:23, 38:29, 38:35, 44:40, 44:44, 44:130 and 44:156 alleles. Specific PCR fragment of 205 bp in the B*56:31 and the B*15:38:01-15:38:02, 15:185, 18:01:01:01-18:01:13, 18:01:15-18:09, 18:12-18:15, 18:17N-18:19, 18:22-18:25, 18:27-18:34, 18:36-18:40, 18:42-18:55, 18:57, 18:59-18:60, 18:62-18:83, 35:21, 35:24:01-35:24:02, 35:188, 35:190, 39:43, 51:37, 51:45, 53:02, 53:06, 53:28, 57:14 and 58:09 alleles.

12Primer mix 12: Specific PCR fragment of 150 bp in the B*56:25 and 56:37 and the B*07:78, 13:16, 13:20, 13:31w, 13:48, 15:73w, 51:62, 51:106, 52:25, 55:09 and 55:24 alleles. Specific PCR fragment of 180 bp in the B*56:10 and 56:23 and the B*13:06, 13:53, 15:86, 15:224, 35:60, 46:18, 54:01:01-54:02, 54:05N, 54:07-54:08N, 54:10, 54:12-54:13, 54:17-54:19, 54:21-54:25, 54:27-54:28, 55:02:01-55:02:06, 55:07, 55:10, 55:12, 55:16, 55:18-55:19, 55:26, 55:30, 55:34-55:35, 55:37, 55:39, 55:42-55:43, 55:47-55:48, 55:50, 55:56-55:57, 59:01:01:01-59:01:01:02 and 59:05 alleles.Specific PCR fragment of 150 and 180 bp in the B*15:42, 54:20, 55:01:01-55:01:10, 55:03, 55:05, 55:11, 55:15, 55:21, 55:25, 55:29, 55:31, 55:33, 55:36, 55:38, 55:40, 55:44-55:45 and 55:52-55:55N alleles.

13Primer mix 13: Specific PCR fragment of 70 bp in the B*56:30 allele. Specific PCR fragment of 180 bp in the B*56:11 and 56:22 and the B*15:58, 15:133, 15:229, 15:253, 35:03:01-35:03:05, 35:03:07-35:03:12, 35:06, 35:13, 35:36, 35:38, 35:55-35:56, 35:59, 35:70, 35:74-35:75, 35:84-35:85, 35:98, 35:106, 35:109, 35:127-35:128, 35:136, 35:150-35:153, 35:155-35:156, 35:160, 35:163, 35:167, 35:169, 35:179, 35:181, 35:193, 35:195, 35:198, 35:204-35:205, 38:30, 39:32-39:33, 39:74, 51:13:01-51:13:02, 51:92, 52:14, 53:04 and 58:34 alleles.

14Primer mix 15: Specific PCR fragment of 160 bp in the B*56:13 and the B*44:15, 44:18, 44:140, 45:01, 45:04-45:08, 45:10-45:14, 55:56 and 82:01-82:03 alleles. Specific PCR fragment of 380 bp in the B*56:28N allele.

15Primer mix 17: Specific PCR fragment of 80 bp in the B*56:38N allele. Specific PCR fragment of 180 bp in the B*56:05:01-56:06, 56:15 and 56:21 and in the B*07:20, 07:24, 07:60, 07:100, 07:131, 08:21, 08:25, 13:62, 15:09-15:10:02, 15:30, 15:37, 15:45, 15:48, 15:63, 15:90, 15:99, 15:150, 15:243, 15:248, 27:76, 35:02:01-35:02:05, 35:04:01-35:04:03, 35:09:01-35:09:03, 35:12:01-35:12:03, 35:18, 35:22, 35:34, 35:39, 35:44, 35:81, 35:83, 35:95, 35:129N, 35:149, 35:154, 35:162, 35:172, 35:182-35:184, 35:199, 35:201, 35:211, 35:220, 40:05, 40:26, 40:28, 40:63, 40:92, 40:174, 41:17, 44:62, 44:77, 44:82, 44:107, 44:123, 49:18, 50:14, 51:01:01-51:01:14, 51:01:16-51:01:31, 51:01:33-51:09:02, 51:11N-51:12, 51:18-51:20, 51:22-51:24:04, 51:26-51:30, 51:32-51:33, 51:35, 51:38-51:41N, 51:43-51:44N, 51:46, 51:48-51:61:02, 51:64-51:80, 51:83-51:84, 51:86-51:91, 51:94-51:96, 51:98N-51:100, 51:102-51:105, 51:107-51:142, 51:145-51:148, 52:01:01:01-52:07, 52:09-52:13, 52:15-52:24, 52:27-52:31, 53:19, 55:08, 55:51, 57:02:01-57:03:02, 57:05, 57:07, 57:09, 57:12, 57:17, 57:28N, 57:39, 57:42, 57:46, 57:57, 58:08:01-58:08:02, 58:28 and 78:01:01-78:06 and the C*03:81 and C*03:152 alleles.

16Primer mix 19: Specific PCR fragment of 75 bp in the B*56:29 and the B*07:176 alleles. Specific PCR fragment of 185 bp in the B*56:19N and the B*39:40:01N-39:40:02N alleles.

17Primer mix 21: Specific PCR fragment of 110 bp in the B*56:17 and 56:23 and in the B*07:37 and 27:77 alleles. Specific PCR fragment of 170 bp in the B*56:26 allele.

18Primer mix 22: Specific PCR fragment of 105 bp in the B*56:06 and the B*07:65, 07:134 and 35:76 alleles. Specific PCR fragment of 175 bp in the B*56:36 and the B*40:166, 44:90 and 45:06 alleles.

19Primer mix 23: Specific PCR fragment of 85 bp in the B56:20:01-56:20:02 and in the B*07:78 and 40:48 alleles. Specific PCR fragment of 350 bp in the B*56:24 allele.

20Primer mix 3 may amplify the C*03:19 and C*03:102 alleles.

’w’, might be weakly amplified.

1HLA-B*56 alleles in bold lettering are listed as confirmed alleles on the on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.10.0, October 2012.

2The HLA-B*56 primer set cannot separate the B*56:09 and B*35:78 alleles. These alleles can be distinguished by the HLA-B low resolution and/or B*35 kits.

3The HLA-B*56 primer set cannot separate the B*56:18 and B*07:65 alleles. These alleles can be distinguished by the HLA-B low resolution and/or B*07 kits.

4The B*56:13 and 56:28N alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 15.

5The B*56:17 and 56:26 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 21.

6The B*56:19N and 56:29 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 19.

7Primer mix 6: Specific PCR fragment of 105 bp in the B*56:05:01 and the B*27:05:18 and 39:63 alleles. Specific PCR fragment of 150 bp in the B*56:32 and in the B*07:09w, 07:11w, 07:162w, 15:03:01-15:03:03, 15:16:01-15:16:03, 15:18:01-15:18:05, 15:23, 15:29, 15:39:01-15:40, 15:47:01w-15:47:02w, 15:49w, 15:61-15:62, 15:64:01-15:64:02, 15:65w, 15:67, 15:69, 15:72, 15:74, 15:80, 15:93, 15:95, 15:98, 15:103, 15:108, 15:115, 15:119, 15:127, 15:131-15:132, 15:134, 15:138, 15:153, 15:158, 15:161, 15:173, 15:176, 15:194-15:195, 15:197-15:198, 15:210, 15:220, 15:221w, 15:223, 15:226N, 15:238, 15:242, 15:254-15:255, 15:263, 18:15, 18:21w, 18:30w, 18:57, 35:11:01-35:11:02, 35:11:03w, 35:21, 35:58, 35:99w, 35:188, 46:03w, 46:08, 48:25, 49:04, 51:37, 51:45, 53:06, 53:08:01-53:08:02, 53:28, 55:14, 57:13w, 57:22, 58:06 and 58:19 and the C*03:175, C*15:09, 15:19, 15:27, 15:61alleles.

Primer mix 7: Specific PCR fragment of 120 bp in the B*56:05:01-56:06 and 56:21 and the B*13:62, 40:95, 49:18, 50:14, 51:01:01-51:01:03, 51:01:05-51:01:12, 51:01:14-51:02:03, 51:02:05-51:03, 51:05, 51:07:01-51:09:02, 51:11N-51:12, 51:13:02-51:14, 51:16-51:23, 51:24:03-51:24:04, 51:26-51:37, 51:39-51:41N, 51:43-51:44N, 51:48-51:55, 51:57-51:58, 51:60-51:61:02, 51:63, 51:65-51:80, 51:82-51:91, 51:93-51:134, 51:136-51:138, 51:140-51:147, 52:01:01:01-52:01:04, 52:01:06-52:01:14, 52:01:16-52:13, 52:15-52:24, 52:26-52:31, 55:01:04, 58:08:01-58:08:02 and 78:01:01-78:07 alleles. Specific PCR fragment of 170 bp in the B*56:33 and the B*15:21, 15:44, 15:66, 15:189, 35:26, 35:82, 35:85, 35:135 and 38:28 alleles.

Primer mix 9: Specific PCR fragment of 145 bp in the B*56:07 and the B*07:149, 08:02, 15:115, 27:01, 38:02:01-38:04, 38:08, 38:15, 38:18, 38:23, 38:29, 38:35, 44:40, 44:44, 44:130 and 44:156 alleles. Specific PCR fragment of 205 bp in the B*56:31 and the B*15:38:01-15:38:02, 15:185, 18:01:01:01-18:01:13, 18:01:15-18:09, 18:12-18:15, 18:17N-18:19, 18:22-18:25, 18:27-18:34, 18:36-18:40, 18:42-18:55, 18:57, 18:59-18:60, 18:62-18:83, 35:21, 35:24:01-35:24:02, 35:188, 35:190, 39:43, 51:37, 51:45, 53:02, 53:06, 53:28, 57:14 and 58:09 alleles.

Primer mix 12: Specific PCR fragment of 150 bp in the B*56:25 and 56:37 and the B*07:78, 13:16, 13:20, 13:31w, 13:48, 15:73w, 51:62, 51:106, 52:25, 55:09 and 55:24 alleles. Specific PCR fragment of 180 bp in the B*56:10 and 56:23 and the B*13:06, 13:53, 15:86, 15:224, 35:60, 46:18, 54:01:01-54:02, 54:05N, 54:07-54:08N, 54:10, 54:12-54:13, 54:17-54:19, 54:21-54:25, 54:27-54:28, 55:02:01-55:02:06, 55:07, 55:10, 55:12, 55:16, 55:18-55:19, 55:26, 55:30, 55:34-55:35, 55:37, 55:39, 55:42-55:43, 55:47-55:48, 55:50, 55:56-55:57, 59:01:01:01-59:01:01:02 and 59:05 alleles.Specific PCR fragment of 150 and 180 bp in the B*15:42, 54:20, 55:01:01-55:01:10, 55:03, 55:05, 55:11, 55:15, 55:21, 55:25, 55:29, 55:31, 55:33, 55:36, 55:38, 55:40, 55:44-55:45 and 55:52-55:55N alleles.

Primer mix 13: Specific PCR fragment of 70 bp in the B*56:30 allele. Specific PCR fragment of 180 bp in the B*56:11 and 56:22 and the B*15:58, 15:133, 15:229, 15:253, 35:03:01-35:03:05, 35:03:07-35:03:12, 35:06, 35:13, 35:36, 35:38, 35:55-35:56, 35:59, 35:70, 35:74-35:75, 35:84-35:85, 35:98, 35:106, 35:109, 35:127-35:128, 35:136, 35:150-35:153, 35:155-35:156, 35:160, 35:163, 35:167, 35:169, 35:179, 35:181, 35:193, 35:195, 35:198, 35:204-35:205, 38:30, 39:32-39:33, 39:74, 51:13:01-51:13:02, 51:92, 52:14, 53:04 and 58:34 alleles.

Primer mix 15: Specific PCR fragment of 160 bp in the B*56:13 and the B*44:15, 44:18, 44:140, 45:01, 45:04-45:08, 45:10-45:14, 55:56 and 82:01-82:03 alleles. Specific PCR fragment of 380 bp in the B*56:28N allele.

Primer mix 17: Specific PCR fragment of 80 bp in the B*56:38N allele. Specific PCR fragment of 180 bp in the B*56:05:01-56:06, 56:15 and 56:21 and in the B*07:20, 07:24, 07:60, 07:100, 07:131, 08:21, 08:25, 13:62, 15:09-15:10:02, 15:30, 15:37, 15:45, 15:48, 15:63, 15:90, 15:99, 15:150, 15:243, 15:248, 27:76, 35:02:01-35:02:05, 35:04:01-35:04:03, 35:09:01-35:09:03, 35:12:01-35:12:03, 35:18, 35:22, 35:34, 35:39, 35:44, 35:81, 35:83, 35:95, 35:129N, 35:149, 35:154, 35:162, 35:172, 35:182-35:184, 35:199, 35:201, 35:211, 35:220, 40:05, 40:26, 40:28, 40:63, 40:92, 40:174, 41:17, 44:62, 44:77, 44:82, 44:107, 44:123, 49:18, 50:14, 51:01:01-51:01:14, 51:01:16-51:01:31, 51:01:33-51:09:02, 51:11N-51:12, 51:18-51:20, 51:22-51:24:04, 51:26-51:30, 51:32-51:33, 51:35, 51:38-51:41N, 51:43-51:44N, 51:46, 51:48-51:61:02, 51:64-51:80, 51:83-51:84, 51:86-51:91, 51:94-51:96, 51:98N-51:100, 51:102-51:105, 51:107-51:142, 51:145-51:148, 52:01:01:01-52:07, 52:09-52:13, 52:15-52:24, 52:27-52:31, 53:19, 55:08, 55:51, 57:02:01-57:03:02, 57:05, 57:07, 57:09, 57:12, 57:17, 57:28N, 57:39, 57:42, 57:46, 57:57, 58:08:01-58:08:02, 58:28 and 78:01:01-78:06 and the C*03:81 and C*03:152 alleles.

Primer mix 19: Specific PCR fragment of 75 bp in the B*56:29 and the B*07:176 alleles. Specific PCR fragment of 185 bp in the B*56:19N and the B*39:40:01N-39:40:02N alleles.

Primer mix 21: Specific PCR fragment of 110 bp in the B*56:17 and 56:23 and in the B*07:37 and 27:77 alleles. Specific PCR fragment of 170 bp in the B*56:26 allele.

Primer mix 22: Specific PCR fragment of 105 bp in the B*56:06 and the B*07:65, 07:134 and 35:76 alleles. Specific PCR fragment of 175 bp in the B*56:36 and the B*40:166, 44:90 and 45:06 alleles.

Primer mix 23: Specific PCR fragment of 85 bp in the B56:20:01-56:20:02 and in the B*07:78 and 40:48 alleles. Specific PCR fragment of 350 bp in the B*56:24 allele.

‘w’, may be weakly amplified.

Change in revision R01 compared to R00:

- Primer mix 3 may amplify the C*03:19 and C*03:102 alleles. A foot note has been added in the Specificity Table.