HLA-B\*57 (101.567-12/12u) Lot No: 6N7 Expiry Date: 2025-12-01

Name:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

Sample ID:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

DNA Conc.(ng/ul):\_\_\_\_\_\_\_\_\_

Test Date: \_\_\_\_\_\_\_\_\_\_\_

Tested By: \_\_\_\_\_\_\_\_\_\_\_\_

Review Date:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

Reviewed By:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

***Interpretation:\_\_\_\_\_\_\_\_\_\_\_ Failed lanes: \_\_\_\_\_\_\_\_\_\_ Comments:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_***

**Gel Picture**

|  |
| --- |
| PHOTO DOCUMENT |





Abbreviations

ICB: Internal Control Band

AmpS: Amplicon size

**Notes:**

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Interpretation Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mix 9 may have tendencies of unspecific amplifications.

Primer mix 15 has a tendency to giving rise to primer oligomer formation.

Primer mix 23 may give rise to a lower yield of HLA-specific PCR product than the other HLA-B\*57 resolution primer mixes.

Primer mix 32 contains a negative control, which will amplify the majority of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 200 base pairs.









**1**HLA-B\*57 alleles in bold lettering are listed as confirmed alleles on the on the IMGT/HLA web page [www.ebi.ac.uk/imgt/hla](http://www.ebi.ac.uk/imgt/hla), release 3.25.0, July 2016.

**2**Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page <http://hla.alleles.org/alleles/deleted.html>.

**3**The following HLA-B\*57 primer mixes have two or more product sizes:

|  |  |  |  |
| --- | --- | --- | --- |
| Primer Mix | Size of spec. PCR product | Amplified HLA-B\*57 alleles | Other amplified HLA Class I alleles |
| **6** | 95 bp 180 bp205 bp | \*57:04:01-57:04:03, 57:32\*57:06\*57:30, 57:129 | \*44:153 |
| **7** | 105 bp170 bp295 bp | \*57:15\*57:37\*57:55, 57:94 | \*14:104, 44:517 |
| **8** | 90 bp 200 bp 250 bp | \*57:16\*57:49\*57:07, 57:93, 57:119  | \*44:153 |
| **11** | 215 bp250 bp | \*57:09, 57:24, 57:119\*57:14:01-57:14:02, 57:43 | \*08:181 |
| **14** | 135 bp195 bp | \*57:17 \*57:10  | **\***44:189, **C\*07:239****\***07:219 |
| **15** | 110 bp145 bp | \*57:29, 57:33\*57:11   | \*14:61, 15:214, 18:81, 35:250, 37:93, 39:145, 39:158, 40:427, 44:311, 44:450, 51:165, 55:108\*14:01:01:01-14:02:08, 14:02:10-14:04, 14:07N, 14:09, 14:11-14:12, 14:14-14:36, 14:38-14:52, 14:54, 14:57-14:70Q, 14:72N-14:95, 14:97-14:101N, 14:103-14:104, 15:510, 18:44:01-18:44:02, 35:475, 39:79, 44:440, 46:84, 50:51, 54:38, 58:02:01:01-58:02:02, 58:06-58:07, 58:25, 58:38, 58:43, 58:60, 58:104, 58:132, **C\*01:32:01-01:32:02, C\*06:20, C\*14:82, C\*16:98, C\*16:102, C\*16:110** |
| **16** | 85 bp 160 bp | \*57:12\*57:18  | \*08:119, 14:24, 35:226, 39:92, 58:64 |
| **17** | 140 bp 210 bp | \*57:13, 57:31:01-57:31:02\*57:13, 57:22, 57:27, 57:57, 57:63  | \*07:227, 08:181, 40:30, 40:34, **C\*06:72****\***07:227, 55:14, **C\*06:72w** |
| **19** | 90 bp240 bp | \*57:04:01-57:04:03, 57:41 \*57:13, 57:25, 57:43  | \*44:153, **C\*06:72****\***07:227, 40:30, 40:34 |
| **20** | 90 bp 240 bp | \*57:20\*57:26  | \*08:181, 55:14, **C\*06:72** |
| **21** | 120 bp150 bp | \*57:33, 57:98Q\*57:21  | \*44:450\*35:127, 50:54, 55:100 |
| **22** | 95 bp140 bp | \*57:34\*57:40, 57:98Q | \*58:54\*14:20, 15:461, 44:317 |
| **23** | 75 bp 100 bp465 bp | \*57:35\*57:36\*57:05, 57:67:01  | \*13:01:18, 13:02:06, 13:26:01, 13:72, 13:127, 15:16:03, 15:20, 15:34:02, 15:85, 15:194, 15:393, 15:407, 15:587, 15:599, 35:46, 35:207, 46:01:07, 46:33, 58:36, **C\*01:73** |
| **24** | 230 bp270 bp | \*57:77\*57:40  | \*44:36, **A\*03:174, A\*11:65, A\*24:247, C\*01:170, C\*04:83****\***08:262, 15:461, 40:284, 44:317, **C\*07:511, C\*15:121, C\*16:162** |

**4**The following HLA-B\*57 alleles can be distinguished by the different sizes of the HLA-specific PCR product:

|  |  |
| --- | --- |
| **Alleles** | **Primer mix** |
| B\*57:15, 57:37, 57:55 | 7 |
| B\*57:25, 57:41 | 19 |

Abbreviations

w: might be weakly amplified.