PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® HLA-B*57:01

Prod. No:	101.572-12/12u

Lot No:    	02S

Expiry Date:    	2015-August-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 15+1 primer mixes used for HLA-B*57:01 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified HLA-B*57:01 alleles | Other amplified HLA-B alleles3,4 |
| --- | --- | --- | --- | --- |
| 15 | 90 bp | 800 bp | *57:01:01-57:01:14 | *57:02:01-57:15, 57:17-57:19, 57:21-57:35, 57:37-57:44, 57:46-57:50, 57:52-57:60, 58:36 |
| 2 | 220 bp | 800 bp | *57:01:01-57:01:14 | *57:03:01-57:03:02, 57:06-57:08, 57:10, 57:14-57:18, 57:20-57:23, 57:25-57:27, 57:29, 57:31-57:41, 57:43-57:51, 57:53-57:60, 35:208, 40:30, 40:34, 44:153, 55:14, 58:14 |
| 35,7 | 95 bp, 170 bp, 215 bp | 800 bp |  | *57:04, 57:06, 57:18, 57:27, 57:32, 44:153 |
| 45 | 100 bp | 1070 bp |  | *57:15, 57:20, 57:29, 07:120, 15:214, 18:81, 40:150 |
| 55,8 | 90 bp, 165 bp, 245 bp | 800 bp |  | *57:07, 57:16, 57:23, 57:26, 57:34, 44:153, 55:14, C*06:72 |
| 65,9 | 90 bp, 210 bp | 1070 bp |  | *57:02:01-57:03:02, 57:07-57:09, 57:12, 57:17, 57:39, 57:42, 57:46, 57:57, 40:30, 40:34 |
| 75,6,10 | 100 bp, 140 bp, 175 bp, 215 bp, 240 bp | 1070 bp |  | *57:09, 57:13-57:14, 57:24-57:25, 57:31, 57:50, 57:53, 35:208, 40:30, 40:34, 44:153, 55:14, 58:14, C*06:72 |
| 8 | 185 bp | 1070 bp |  | *57:10, 57:44 |
| 95,11 | 110 bp, 150 bp | 800 bp |  | *57:21, 57:33, 57:40, 14:20, 35:127 |
| 105,6,12 | 90 bp, 170 bp, 205 bp, 240 bp | 1070 bp |  | *57:04, 57:13, 57:22, 57:37, 57:41, 57:43, 57:57, 44:153, 55:14, C*06:72 |
| 115 | 100 bp | 1070 bp | *57:01:01-57:01:14 | *57:06, 57:08, 57:10, 57:13-57:16, 57:18-57:27, 57:29-57:31, 57:33-57:38, 57:40w, 57:41, 57:43-57:45, 57:47-57:52, 57:54-57:56, 57:58-57:60, 35:208, 55:14, 58:14, C*06:72 |
| 125,6,13 | 75 bp, 100 bp, 140 bp | 1070 bp |  | *57:08, 57:35-57:36, 57:38, 57:60, 15:87 |
| 135,14 | 105 bp, 155 bp, 200 bp | 1070 bp |  | *57:45, 57:49, 57:51, 57:58, 58:12, A*02:42, A*02:285, A*02:310 |
| 145 | 125 bp | 1070 bp |  | *57:47, 57:54, 57:59, 15:33, 15:116, 15:248, 40:63, 40:92, 49:22, 51:126, C*03:87, C*03:129, C*05:27, C*05:39, C*17:07 |
| 1515 | 145 bp, 185 bp, 295 bp | 1070 bp |  | *57:48, 57:55-57:56, A*33:12 |
| 1616 | - | - |  | Negative control |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-B*57:01 SSP typings.

When the primers in a primer mix can give rise to HLA-specific PCR products of more than one length this is indicated if the size difference is more than 20 base pairs. Size differences of 20 base pairs or less are not given. For high resolution SSP kits the respective lengths of the HLA-specific PCR product(s) are given for the alleles amplified by these primer mixes.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most wells, or a band of 800 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-B*57:01 subtyping.

In addition, wells number 2, 3, 5 and 9 contain the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3Due to the sharing of sequence motifs between HLA-B alleles some non-HLA-B*57 alleles will be amplified by primer mixes 1 to 7 and 9 to 14. In addition, primer mix 13 will amplify the A*02:42, A*02:285 and A*02:310 alleles, primer mix 15 will amplify the A*33:12 allele, primer mix 14 will amplify the C*03:87, C*03:129, C*05:27, C*05:39 and C*17:07 alleles, and primer mixes 5, 7, 10 and 11 will amplify the C*06:72 allele.

4For several HLA Class I alleles 1st and/or 4th exon(s) and beyond, as well as intron nucleotide sequences, are not available. In these instances it is not known whether some of the primers of the SSP sets are completely matched with the target sequences or not. We assume that unknown sequences in these regions are conserved within allelic groups.

5HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

6Primer mixes 7, 10 and 12 may have tendencies of unspecific amplifications.

7Primer mix 3: Specific PCR fragment of 95 bp in the B*57:04 and 57:32 and the B*44.153 alleles. Specific PCR fragment of 170 bp in the B*57:06 and B*57:18 alleles. Specific PCR fragment of 215 bp in the B*57:27 allele.

8Primer mix 5: Specific PCR fragment of 90 bp in the B*57:16 and 57:34 alleles. Specific PCR fragment of 165 bp in the B*57:23 allele. Specific PCR fragment of 245 bp in the B*57:07 and 57:26 and in the B*44:153, 55:14 and the C*06:72 alleles.

9Primer mix 6: Specific PCR fragment of 90 bp in the B*57:02:01-57:03:02, 57:07, 57:09, 57:12, 57:17, 57:39, 57:42, 57:46 and 57:57 and the B*40:30 and 40:34 alleles. Specific PCR fragment of 210 bp in the B*57:08 allele.

10Primer mix 7: Specific PCR fragment of 100 bp in the B *57:53 allele. Specific PCR fragment of 140 bp in the B*57:31 and in the C*06:72 alleles. Specific PCR fragment of 175 bp in the B*57:14 and 57:50 and the B*35:208, 44:153, 55:14 and 58:14 alleles. Specific PCR fragment of 215 bp in the B*57:09 and 57:24 alleles. Specific PCR fragment of 240 bp in the B*57:25 allele. Specific PCR fragment of 140 and 240bp in the B*57:13 and in the B*40:30 and 40:34 alleles.

11Primer mix 9: Specific PCR fragment of 110 bp in the B*57:33 allele. Specific PCR fragment of 150 bp in the B*57:21 and 57:40 and the B*14:20 and 35:127 alleles.

12Primer mix 10: Specific PCR fragment of 90 bp in the B *57:04, 57:41 and the B*44:153 alleles. Specific PCR fragment of 170 bp in the B*57:37 allele. Specific PCR fragment of 205 bp in the B*57:13, 57:22, 57:57 and the B*55:14 alleles. Specific PCR fragment of 240 bp in the B*57:43 allele. Specific PCR fragment of 90 and 205 bp in the C*06:72 allele.

13Primer mix 12: Specific PCR fragment of 75 bp in the B*57:35 allele. Specific PCR fragment of 100 bp in the B *57:36 allele. Specific PCR fragment of 140 bp in the B*57:08, 57:38 and 57:60 and the B*15:87 alleles.

14Primer mix 13: Specific PCR fragment of 105 bp in the B*57:45 and 57:51 and in the A*02:285 alleles. Specific PCR fragment of 155 bp in the B*57:58 and in the B*58:12 and the A*02:42 and A*02:310 alleles. Specific PCR fragment of 200 bp in the B*57:49 allele.

15Primer mix 15: Specific PCR fragment of 145 bp in the B *57:56 allele. Specific PCR fragment of 185 bp in the B*57:48 and in the A*33:12 alleles. Specific PCR fragment of 295 bp in the B*57:55 allele.

16Well 16 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by control primer pairs. PCR product sizes range from 75 to 200 base pairs. The PCR product generated by the control primer pair is 430 base pairs.

1Primer mix 3: Specific PCR fragment of 95 bp in the B*57:04 and 57:32 and the B*44.153 alleles. Specific PCR fragment of 170 bp in the B*57:06 and B*57:18 alleles. Specific PCR fragment of 215 bp in the B*57:27 allele.

Primer mix 5: Specific PCR fragment of 90 bp in the B*57:16 and 57:34 alleles. Specific PCR fragment of 165 bp in the B*57:23 allele. Specific PCR fragment of 245 bp in the B*57:07 and 57:26 and in the B*44:153, 55:14 and the C*06:72 alleles.

Primer mix 6: Specific PCR fragment of 90 bp in the B*57:02:01-57:03:02, 57:07, 57:09, 57:12, 57:17, 57:39, 57:42, 57:46 and 57:57 and the B*40:30 and 40:34 alleles. Specific PCR fragment of 210 bp in the B*57:08 allele.

Primer mix 7: Specific PCR fragment of 100 bp in the B *57:53 allele. Specific PCR fragment of 140 bp in the B*57:31 and in the C*06:72 alleles. Specific PCR fragment of 175 bp in the B*57:14 and 57:50 and the B*35:208, 44:153, 55:14 and 58:14 alleles. Specific PCR fragment of 215 bp in the B*57:09 and 57:24 alleles. Specific PCR fragment of 240 bp in the B*57:25 allele. Specific PCR fragment of 140 and 240bp in the B*57:13 and in the B*40:30 and 40:34 alleles.

Primer mix 9: Specific PCR fragment of 110 bp in the B*57:33 allele. Specific PCR fragment of 150 bp in the B*57:21 and 57:40 and the B*14:20 and 35:127 alleles.

Primer mix 10: Specific PCR fragment of 90 bp in the B *57:04, 57:41 and the B*44:153 alleles. Specific PCR fragment of 170 bp in the B*57:37 allele. Specific PCR fragment of 205 bp in the B*57:13, 57:22, 57:57 and the B*55:14 alleles. Specific PCR fragment of 240 bp in the B*57:43 allele. Specific PCR fragment of 90 and 205 bp in the C*06:72 allele.

Primer mix 12: Specific PCR fragment of 75 bp in the B*57:35 allele. Specific PCR fragment of 100 bp in the B *57:36 allele. Specific PCR fragment of 140 bp in the B*57:08, 57:38 and 57:60 and the B*15:87 alleles.

Primer mix 13: Specific PCR fragment of 105 bp in the B*57:45 and 57:51 and in the A*02:285 alleles. Specific PCR fragment of 155 bp in the B*57:58 and in the B*58:12 and the A*02:42 and A*02:310 alleles. Specific PCR fragment of 200 bp in the B*57:49 allele.

Primer mix 15: Specific PCR fragment of 145 bp in the B *57:56 allele. Specific PCR fragment of 185 bp in the B*57:48 and in the A*33:12 alleles. Specific PCR fragment of 295 bp in the B*57:55 allele.

2Primer mix 16 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by control primer pairs. PCR product sizes range from 75 to 200 base pairs. The PCR product generated by the control primer pair is 430 base pairs.

Change in revision R01 compared to R00:

- Primer mix 11 amplifies weakly the B*57:40 allele. This has been corrected in the Specificity and Interpretation Tables.