Gel Documentation Form and Worksheet

HLA-B*57:01 (101.572-12/12u) 	Lot No: 10V	Expiry Date: 2016-June-01

Name:________________

Sample ID:_______________

DNA Conc.(ng/ul):________

Test  Date: ________      _

Tested By: ____      _____

Review Date:_______________

Reviewed By:______________ _

Gel Picture

PHOTO DOCUMENT

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon size

Interpretation:_______________

Failed lanes:    _________

Comments:    ___________________

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 7 and 10 may have tendencies of unspecific amplifications.

Well 16 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by control primer pairs. PCR product sizes range from 75 to 200 base pairs. The PCR product generated by the control primer pair is 430 base pairs.

1Primer mix 3: Specific PCR fragment of 95 bp in the B*57:04 and 57:32 and the B*44.153 alleles. Specific PCR fragment of 170 bp in the B*57:06 and B*57:18 alleles. Specific PCR fragment of 215 bp in the B*57:27 allele.

Primer mix 5: Specific PCR fragment of 90 bp in the B*57:16 and 57:34 alleles. Specific PCR fragment of 165 bp in the B*57:23 allele. Specific PCR fragment of 200 bp in the B*57:46 and 57:64 alleles. Specific PCR fragment of 245 bp in the B*57:07 and 57:26 and in the B*44:153, 55:14 and the C*06:72 alleles.

Primer mix 6: Specific PCR fragment of 90 bp in the B*57:02:01-57:03:02, 57:07, 57:09, 57:12, 57:17, 57:39, 57:42, 57:46 57:57 and 57:63 and the B*40:30 and 40:34 alleles. Specific PCR fragment of 205 bp in the B*57:08 and 57:65 alleles.

Primer mix 7: Specific PCR fragment of 100 bp in the B *57:53 allele. Specific PCR fragment of 140 bp in the B*57:31 and and in the C*06:72 alleles. Specific PCR fragment of 175 bp in the B*57:14 and 57:50 and the B*35:208, 44:153, 55:14 and 58:14 alleles. Specific PCR fragment of 215 bp in the B*57:09 and 57:24 alleles. Specific PCR fragment of 240 bp in the B*57:25 allele. Specific PCR fragment of 140 and 240bp in the B*57:13 and in the B*40:30 and 40:34 alleles.

Primer mix 9: Specific PCR fragment of 110 bp in the B*57:33 allele. Specific PCR fragment of 150 bp in the B*57:21 and 57:40 and the B*14:20 and 35:127 alleles.

Primer mix 10: Specific PCR fragment of 90 bp in the B *57:04, 57:41 and the B*44:153 alleles. Specific PCR fragment of 170 bp in the B*57:37 allele. Specific PCR fragment of 205 bp in the B*57:13, 57:22, 57:57 and 57:63 and the B*55:14 alleles. Specific PCR fragment of 240 bp in the B*57:43 allele. Specific PCR fragment of 90 and 205 bp in the C*06:72 allele.

Primer mix 12: Specific PCR fragment of 75 bp in the B*57:35 allele. Specific PCR fragment of 100 bp in the B *57:36 allele. Specific PCR fragment of 140 bp in the B*57:08, 57:38 and 57:60 and the B*15:87 alleles.

Primer mix 13: Specific PCR fragment of 105 bp in the B*57:45 and 57:51 and in the A*02:285 alleles. Specific PCR fragment of 155 bp in the B*57:58 and in the B*58:12 and the A*02:42 and A*02:310 alleles. Specific PCR fragment of 200 bp in the B*57:49 allele.

Primer mix 15: Specific PCR fragment of 145 bp in the B *57:56 allele. Specific PCR fragment of 185 bp in the B*57:48 and in the A*33:12 alleles. Specific PCR fragment of 295 bp in the B*57:55 allele.

2Primer mix 16 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by control primer pairs. PCR product sizes range from 75 to 200 base pairs. The PCR product generated by the control primer pair is 430 base pairs.