HLA-B\*57:01 (101.572-12/12u) Lot No: 0S7 Expiry Date: 2027-05-01

Name:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_ Test Date: \_\_\_\_\_\_\_\_\_\_\_\_ Review Date:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

Sample ID:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_ Tested By: \_\_\_\_\_\_\_\_\_\_\_\_\_ Reviewed By:\_\_\_\_\_\_\_\_\_\_\_\_\_\_

DNA Conc.(ng/ul):\_\_\_\_\_\_\_\_\_

***Interpretation:\_\_\_\_\_\_\_\_\_\_\_ Failed lanes: \_\_\_\_\_\_\_\_\_\_\_ Comments:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_***

**Gel Picture**

|  |
| --- |
| PHOTO DOCUMENT |



Abbreviations

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon size

**Notes:**

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mix 2 may give rise to a lower yield of HLA-specific PCR product than the other B\*57:01 primer mixes.

Primer mix 18 has tendency of unspecific amplifications.

Primer mix 24 contains a negative control, which will amplify a majority of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 200 base pairs.







**1**HLA-B alleles listed on the IMGT/HLA web page 2023-01-12, release 3.51.0, [www.ebi.ac.uk/imgt/hla](http://www.ebi.ac.uk/imgt/hla).

**2**Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page <http://hla.alleles.org/alleles/deleted.html>.

**3**The following HLA-B\*57 primer mixes have two or more product sizes. Note that no HLA-B\*57:01 alleles are amplified in any of these primer mixes.

|  |  |  |
| --- | --- | --- |
| Primer Mix | Size of spec. PCR product | Amplified Class I alleles |
| **3** | 95 bp  170 bp  215 bp | \*57:04:01-57:04:03, 57:32, **44:153**  \*57:06, 57:18  \*57:27 |
| **4** | 90 bp  105 bp  170 bp | \*57:20, 57:62  \*57:15, 57:29, **08:301, 14:61, 15:214, 18:81, 35:250, 37:93, 39:145, 39:158, 40:427, 44:311, 51:165, 55:108**  \*57:91, **15:626, 53:59** |
| **5** | 90 bp  165 bp  200 bp  245 bp | \*57:16, 57:34, **58:54**  \*57:23  \*57:46, 57:64  \*57:07, 57:26,57:93, 57:119, **08:181, 44:153, 55:14, C\*06:72** |
| **6** | 90 bp  205 bp | \*57:02:01:01-57:03:05, 57:07, 57:09, 57:12, 57:17, 57:39, 57:42, 57:46, 57:57, 57:63, 57:66, 57:70, 57:80, 57:84, 57:94, 57:96, 57:101, 57:124, 57:137, 57:155, **08:181, 40:30, 40:34**  \*57:08, 57:65, **15:403** |
| **7** | 100 bp  140 bp  185 bp  215 bp  250 bp | \*57:53  \*57:13, 57:31:01-57:31:02, **07:227, 08:181, 40:30, 40:34, C\*06:72**  \*57:50  \*57:09, 57:24, 57:119, **08:181**  \*57:13, 57:14:01-57:14:02, 57:25, **07:227, 40:30, 40:34** |
| **8** | 160 bp  210 bp | \*57:87  \*57:10, 57:39, 57:51, 57:73, **07:219, 15:446, 58:41** |
| **9** | 110 bp  150 bp | \*57:33, **44:450**  \*57:21, 57:40, 57:74, **14:20, 15:461, 27:188, 35:127, 44:317, 50:54, 51:186, 53:41, 55:100, 56:80** |
| **10** | 205 bp  240 bp | \*57:13, 57:22, 57:57, 57:63, 57:161w, **07:227, 55:14, C\*06:72w**  **\***57:43 |
| **12** | 75 bp  140 bp | \*57:35, **15:625**  \*57:08, 57:60, **15:87, 15:403** |
| **13** | 105 bp  135 bp  200 bp | \*57:45, 57:51, 57:69, 57:71, 57:88, **58:111, A\*02:285, A\*02:799**  \*57:58, 57:80, 57:90, **07:219, 58:12, A\*02:42:01-02:42:02, A\*02:310, A\*02:597**  \*57:49 |
| **14** | 115 bp  135 bp  205 bp | \*57:54, 57:96, **15:33, 15:248, 49:22, 51:126, C\*03:87:01-03:87:02, C\*03:414, C\*05:27, C\*05:39, C\*05:184, C\*08:115, C\*08:182**  \*57:47  \*57:77, **15:617, 44:36, C\*03:232** |
| **15** | 145 bp  185 bp  295 bp | \*57:56, 57:78, **15:340, 35:300, 40:218, 44:307, 45:21, 46:28, 51:58, 52:82, 58:83**  \*57:48, **A\*23:72, A\*33:12**  \*57:55, 57:94, **14:104, 44:517** |
| **17** | 90 bp  180 bp | \*57:04:01-57:04:03, 57:41, **44:153, C\*06:72**  \*57:68 |
| **18** | 170 bp  275 bp  300 bp | \*57:37, 57:79N  \*57:81, **44:13, 44:67, 51:170**  \*57:86, **51:356, 53:25** |
| **19** | 100 bp  165 bp  215 bp | \*57:76  \*57:92  \*57:72 |
| **20** | 240 bp  340 bp | \*57:44, **18:167,** **44:249**  \*57:75, **35:374, 40:327?, 52:81** |
| **21** | 100 bp  135 bp  485 bp | \*57:36  \*57:38  \*57:05 |

Abbreviations

‘w’, might be weakly amplified.

‘?’, nucleotide sequence of the primer matching region not known.