PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® HLA-B*58

Prod. No:	101.568-06/06u

Lot No:    	50N

Expiry Date:    	2014-October-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 24 primer mixes used for HLA-B*58 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified HLA-B*58 alleles3 | Other amplified HLA Class I alleles4 |
| --- | --- | --- | --- | --- |
| 15 | 90 bp | 800 bp | *58:01:01-58:02, 58:05-58:35 |  |
| 2 | 150 bp | 800 bp | *58:01:01-58:01:09, 58:04-58:05, 58:09-58:15, 58:17N, 58:19, 58:21-58:24, 58:28-58:36 | *08:49, 13:01:01-13:01:06, 13:06-13:07N, 13:12-13:13, 13:17, 13:20, 13:22:01-13:23, 13:25-13:26, 13:28-13:29, 13:36, 13:39, 13:43, 13:50-13:52, 14:10, 15:02:01-15:02:05, 15:13:01-15:13:02, 15:20-15:21, 15:25:01-15:25:03, 15:36, 15:44, 15:62, 15:77, 15:80, 15:85, 15:88-15:89, 15:106, 15:112, 15:121, 15:139, 15:144, 15:154, 15:165, 15:170, 15:194, 15:204, 15:213-15:214, 15:223, 18:22, 27:19, 27:30, 35:01:01:01-35:04:03, 35:06-35:08:05, 35:10-35:17, 35:19-35:21, 35:23-35:30, 35:33-35:36, 35:38-35:42:02, 35:45-35:50, 35:52, 35:54-35:57, 35:59, 35:61-35:63, 35:65Q, 35:69-35:71, 35:74, 35:76-35:78, 35:80-35:85, 35:90-35:96, 35:98, 35:100-35:101:02, 35:103-35:113, 35:115-35:116, 35:120-35:126, 35:128-35:134N, 35:136-35:150, 35:152-35:173N, 35:175-35:184, 35:186, 37:01:01-37:01:07, 37:03N-37:06, 37:08, 37:10-37:11, 37:13-37:31, 38:20, 39:42, 40:28, 44:02:01:01-44:14, 44:16-44:17, 44:19N, 44:21-44:30, 44:32-44:40, 44:42-44:46, 44:48-44:52N, 44:55-44:64:02, 44:66-44:98, 44:101-44:105, 44:107-44:134, 44:136-44:137, 44:139, 48:02:01-48:02:02, 48:25, 51:04, 51:42, 51:46, 51:56:01-51:56:02, 53:01:01-53:13, 53:15-53:27, 55:14, 56:09, 56:11-56:12, 57:01:01-57:01:04, 57:01:06-57:10, 57:12, 57:14-57:20, 57:22-57:30, 57:32-57:46, 57:48-57:52, 83:01, C*03:102 |
| 3 | 240 bp | 800 bp | *58:02, 58:06-58:07, 58:25 | *14:19, 14:27, 57:11, C*01:35, C*03:17, C*03:71, C*14:20 |
| 410 | 190 bp, 230 bp | 1070 bp | *58:04, 58:15 |  |
| 55 | 90 bp | 1070 bp | *58:05 | *15:124, 40:77, 40:87:01-40:87:02, 48:14, 48:26, C*07:81, C*07:168 |
| 6 | 170 bp | 1070 bp | *58:06, 58:08:01-58:08:02, 58:19 | *07:02:01-07:24, 07:26-07:47, 07:49N-07:50, 07:52-07:144, 08:20, 08:53, 08:79, 13:16, 13:20, 13:31, 13:48, 14:01:01-14:28, 15:01:01:01-15:01:04, 15:01:06-15:04, 15:06-15:19, 15:21, 15:23-15:30, 15:32-15:40, 15:42-15:47, 15:49-15:50, 15:53-15:54, 15:56-15:58, 15:60-15:74, 15:76-15:82, 15:85, 15:87, 15:89-15:90, 15:92-15:99, 15:101-15:104, 15:106, 15:108-15:110, 15:112-15:113, 15:115-15:122, 15:125-15:129, 15:131-15:135, 15:137-15:144, 15:146-15:150, 15:152-15:154, 15:156-15:161, 15:163-15:175, 15:177-15:178, 15:180-15:184, 15:187, 15:189-15:199, 15:201-15:217, 15:219-15:221, 15:223, 15:225-15:234, 15:236, 15:238, 18:15, 18:19, 18:21, 18:30, 18:57, 27:04:01-27:04:03, 27:06, 27:10, 27:15, 27:18, 27:20-27:21, 27:24-27:25, 27:40, 27:54, 27:63, 27:66N, 27:68-27:69, 27:79, 27:86, 35:11:01-35:11:02, 35:14:01-35:14:02, 35:21, 35:43:01-35:44, 35:58, 35:67, 35:79, 35:86, 35:96, 35:99, 35:102, 35:117-35:118, 35:135, 35:167, 35:43:02, 37:07, 37:13, 38:10, 38:32, 39:18, 39:35-39:36, 40:05, 40:15-40:16, 40:23, 40:26, 40:28, 40:32, 40:51, 40:95, 40:98, 40:148, 40:158, 40:161, 40:174, 44:76, 44:79, 46:01:01-46:05, 46:07N-46:08, 46:10, 46:12, 46:14-46:17, 46:20, 46:22-46:24, 46:27-46:30, 48:05, 48:08, 48:15, 48:25, 49:01:01-49:10, 49:12-49:20, 50:01:01-50:02, 50:04-50:15, 51:01:01-51:04, 51:06:01-51:07:02, 51:11N-51:14, 51:16-51:18, 51:21-51:24:04, 51:26-51:30, 51:32-51:39, 51:41N, 51:43, 51:45-51:46, 51:48-51:52, 51:55-51:72, 51:74-51:80, 51:82-51:92, 51:94-51:96, 51:98N-51:107, 51:109-51:114, 51:116-51:119, 51:121, 51:123-51:126, 52:01:01:01-52:02, 52:04-52:09, 52:11-52:18, 52:20-52:27, 53:06, 53:08:01-53:08:02, 54:06, 54:20, 55:01:01-55:01:08, 55:03, 55:05, 55:09, 55:11, 55:14-55:15, 55:17, 55:21, 55:24-55:25, 55:28-55:29, 55:31, 55:33, 55:36, 55:38, 55:40, 55:44-55:45, 55:52-55:53, 56:03, 56:05:01-56:06, 56:21, 56:25, 56:32, 57:09, 57:13, 57:22, 78:01:01-78:03, 78:05-78:07 |
| 75,7 | 80 bp | 1070 bp | *58:06, 58:08:01-58:08:02, 58:19 | *07:02:13w, 07:68:02w, 13:16, 13:20, 13:48, 14:01:01-14:02:06, 14:03w, 14:04-14:11, 14:12w, 14:13-14:28, 15:03:02, 15:13:02, 15:39:02, 15:42, 18:15, 35:11:01, 35:21, 35:58, 35:96, 35:167, 38:10, 38:32, 39:35, 40:28, 40:95, 44:76w, 44:79, 48:25, 49:01:01, 49:01:03-49:10, 49:12-49:20, 50:01:01-50:01:02, 50:02, 50:04-50:15, 51:01:01-51:01:24, 51:01:26-51:02:04, 51:03-51:04, 51:06:01-51:07:02, 51:11N-51:14, 51:16-51:18, 51:21-51:24:01, 51:24:04, 51:26-51:28, 51:29w, 51:30, 51:32-51:35, 51:37-51:39, 51:41N, 51:43, 51:45-51:46, 51:48-51:53, 51:55w, 51:56:01-51:60, 51:62-51:70, 51:74-51:80, 51:82w, 51:83-51:86, 51:88-51:90, 51:92, 51:94-51:96, 51:98N-51:107, 51:109-51:114, 51:116-51:121, 51:123-51:125, 52:01:01:01-52:01:06, 52:01:08-52:02, 52:04-52:09, 52:11-52:18, 52:20, 52:22-52:23, 52:24w, 52:25-52:27, 53:06, 53:08:01, 54:20w, 55:01:01-55:01:08, 55:03, 55:05, 55:09w, 55:11, 55:14-55:15, 55:17, 55:24-55:25, 55:28-55:29, 55:31, 55:33, 55:36, 55:38, 55:40, 55:44-55:45, 55:52w, 55:53-55:54, 56:05:01-56:06, 56:21, 56:25, 56:32, 78:01:01-78:03, 78:05-78:07, C*03:119, C*07:165 |
| 8 | 250 bp | 1070 bp | *58:07 |  |
| 98 | 160 bp | 800 bp | *58:08:01, 58:23 | A*02:01:33 |
| 105,11 | 115 bp, 270 bp | 800 bp | *58:09, 58:24 | *14:10, 18:22, 35:21, 35:24:01-35:24:02, 35:81, 35:96, 35:109, 35:157, 37:04:01-37:04:02, 40:28, 51:04, 51:46, 51:56:01-51:56:02, 53:02, 53:06, 57:14, C*15:18, C*15:39 |
| 119 | 205 bp | 1070 bp | *58:02, 58:07, 58:25 | *18:44, 57:11, A*24:64w, C*01:06w, C*16:43w |
| 12 | 235 bp | 1070 bp | *58:10N |  |
| 13 | 165 bp | 1070 bp | *58:01:01-58:02, 58:04-58:10N, 58:12-58:19, 58:21-58:36 | *15:20, 15:228, 35:01:01:01-35:01:27, 35:01:29-35:30, 35:32:01-35:42:02, 35:44-35:56, 35:58-35:59, 35:61-35:72, 35:74-35:78, 35:80-35:86, 35:88-35:184, 35:186, 48:02:01-48:02:02, 51:01:01-51:24:04, 51:26-51:41N, 51:43-51:46, 51:48-51:126, 52:01:01:01-52:06:02, 52:08-52:27, 53:01:01-53:27, 56:05:01-56:05:02, 56:21, 78:01:01-78:07, 81:03, 83:01? |
| 147,8 | 165 bp | 1070 bp | *58:11, 58:20 | *07:02:01-07:57, 07:59-07:144, 08:01:01-08:05, 08:07-08:80, 13:01:01-13:04, 13:06-13:23, 13:25-13:52, 14:01:01-14:28, 15:01:01:01-15:01:04, 15:01:06-15:19, 15:21, 15:23-15:40, 15:42-15:58, 15:60-15:99, 15:101, 15:103-15:129, 15:131-15:227, 15:229-15:238, 18:01:01-18:15, 18:17N-18:68, 27:01-27:21, 27:23-27:86, 35:31, 35:35, 35:43:01, 35:46, 35:57, 35:60, 35:79, 35:87, 35:96, 35:98, 35:43:02, 37:01:01-37:31, 38:01:01-38:37, 39:01:01:01-39:01:01:02L, 39:01:03-39:20, 39:22-39:45, 39:47-39:68, 40:01:01-40:16, 40:18-40:40, 40:42-40:180, 41:01-41:19, 42:01:01-42:02, 42:04-42:16, 44:02:01:01-44:65, 44:67-44:117, 44:119-44:139, 45:01-45:13, 46:01:01-46:30, 47:01:01:01-47:08, 48:01:01-48:01:03, 48:02:02-48:27, 49:01:01-49:20, 50:01:01-50:02, 50:04-50:15, 51:42, 52:07, 54:01:01-54:24, 55:01:01-55:05, 55:07-55:54, 56:01:01-56:04, 56:06-56:20:02, 56:22-56:32, 57:01:01-57:01:10, 57:01:12-57:52, 59:01:01:01-59:05, 67:01:01-67:03, 81:01-81:02, 81:04N-81:05, 82:01-82:03, 83:01? |
| 15 | 155 bp | 1070 bp | *58:12 |  |
| 16 | 145 bp | 1070 bp | *58:13 | *49:06 |
| 175,7,8,12 | 55 bp, 225 bp, 375 bp | 1070 bp | *58:16, 58:18, 58:22 | *07:12, 07:14, 07:18:01-07:18:02, 07:137, 35:08:01-35:08:05, 35:18, 35:45, 35:61, 35:80, 35:100, 35:105, 35:142, 35:156, 35:158, 35:164, 35:172, 35:174, 35:176, 44:04, 44:28:01-44:28:02, 44:56N, 44:76, 53:11, 53:16, 53:24, 57:02:01-57:02:02, 57:04-57:05, 57:12-57:13, 57:19, 57:28N, 57:30, 57:42, C*01:49-01:50, C*03:28, C*03:49, C*03:117, C*04:80, C*04:100, C*05:10, C*15:08 |
| 185,13 | 70 bp, 140 bp | 1070 bp | *58:17N, 58:25 |  |
| 195 | 95 bp | 800 bp | *58:18 | *07:09, 07:11, 07:17, 08:28, 08:35, 08:37, 15:07:01-15:07:02, 15:55, 15:68, 15:126, 15:207, 18:14, 35:05:01-35:05:02, 35:51, 35:58, 35:66, 35:72, 35:89, 35:97, 35:114, 40:03, 40:105, 42:09, 46:12, 48:14, 53:14 |
| 205,14 | 85 bp, 170 bp | 1070 bp | *58:21, 58:29 | *18:51 |
| 21 | 440 bp | 1070 bp | *58:14 |  |
| 22 | 180 bp | 1070 bp | *58:30 | *35:70 |
| 2315 | 150 bp, 260 bp | 1070 bp | *58:26, 58:31N | *35:75, 44:138Q |
| 245,8 | 95 bp | 1070 bp | *58:27-58:28 | *07:05:01-07:06, 07:32, 07:34, 07:40, 07:53, 07:69, 07:80, 07:90, 07:97, 07:105, 07:112, 07:123, 07:137-07:138, 07:140, 08:01:01-08:05, 08:08N, 08:10-08:11, 08:13, 08:15, 08:17-08:27, 08:29, 08:31-08:34, 08:36, 08:38-08:48, 08:50-08:68, 08:70-08:80, 13:13, 13:21, 14:13, 27:07:01-27:07:02, 27:11, 27:24, 27:33, 27:43, 35:02:01-35:02:04, 35:04:01-35:04:03, 35:09:01-35:09:03, 35:12:01-35:12:03, 35:18, 35:22, 35:31, 35:81, 35:83, 35:87-35:88, 35:95, 35:129N, 35:149, 35:154, 35:157, 35:162, 35:172, 35:182-35:184, 37:12, 39:14, 40:02:01-40:02:12, 40:04-40:05, 40:08, 40:13, 40:15-40:16, 40:19, 40:28-40:30, 40:34-40:35, 40:37, 40:39-40:40, 40:45, 40:50, 40:56-40:58, 40:64, 40:68, 40:78, 40:80, 40:82, 40:85, 40:89-40:91, 40:94, 40:97, 40:99, 40:104, 40:107, 40:111, 40:115, 40:119, 40:122, 40:129, 40:133Q, 40:137, 40:142N-40:145, 40:157, 40:160, 40:164, 40:169, 40:173-40:174, 40:176, 41:02:01-41:02:04, 41:10-41:11, 41:13, 41:15, 41:19, 42:01:01-42:02, 42:05:01-42:05:02, 42:07-42:08, 42:10-42:15, 44:62, 44:77, 44:82, 48:01:01-48:01:03, 48:04-48:06, 48:09-48:13, 48:15-48:16, 48:18-48:20, 48:22, 48:24, 48:27, 51:04, 51:46, 51:56:01-51:56:02, 51:64, 51:81, 53:19, 55:04, 55:49, 55:51, 56:12, 57:02:01-57:03:02, 57:05, 57:07, 57:09, 57:12, 57:17, 57:39, 57:42, 57:46, 81:01-81:05, C*15:24 |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-B*58 SSP typings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective lengths of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most wells, or a band of 800 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA- B*58 subtyping.

In addition, wells number 2, 3, 9, 10 and 19 contain the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3For many HLA-B alleles fourth exon or intron nucleotide sequences are not available. We assume that unknown sequences in the fourth exon and the introns are conserved within allelic groups.

4Due to the sharing of sequence motifs between HLA-B alleles some non-HLA-B*58 alleles will be amplified by primer mixes 2, 3, 5 to 7, 10, 11, 13, 14, 16, 17, 19, 20 and 22 to 24. In addition, the A*02:01:33 allele will be amplified by primer mix 9, the A*24:64, C*01:06 and C*16:43 alleles will be weakly amplified by primer mix 11, the C*01:35, C*03:17, C*03:71 and C*14:20 alleles will be amplified by primer mix 3, the C*01:49-01:50, C*03:28, 03:49, 03:117, C*04:80, 04:100, C*05:10, C*15:08 alleles will be amplified by primer mix 17, the C*03:102 allele will be amplified by primer mix 2, the C*03:119 and C*07:165 alleles will be amplified by primer mix 7, the C*07:81 and C*07:168 alleles will be amplified by primer mix 5, the C*15:18 and C*15:39 alleles will be amplified by primer mix 10 and the C*15:24 allele will be amplified by primer mix 24.

5Short specific PCR fragments are less intense and not as sharp as longer specific bands.

7Primer mixes 7, 14 and 17 have a tendency of giving rise to nonspecific amplifications.

8Primer mixes 9, 14, 17 and 24 have a tendency of primer dimer formation.

9Primer mix 11 may give a lower yield of HLA-specific PCR product than the other HLA-B*58 primer mixes.

10Primer mix 4: Specific PCR fragment of 190 bp in the B*58:04 allele. Specific PCR fragment of 230 bp in the B*58:15 allele.

11Primer mix 10: Specific PCR fragment of 115 bp in the B*58:24 and the C*15:18 alleles. Specific PCR fragment of 270 bp in the B*58:09 and the B*14:10, 18:22, 35:21, 35:24:01-35:24:02, 35:81, 35:96, 35:109, 35:157, 37:04:01-37:04:02, 40:28, 51:04, 51:46, 51:56:01-51:56:02, 53:02, 53:06 and 57:14 and in the C*15:39 alleles.

12Primer mix 17: Specific PCR fragment of 225 bp in the B*58:22 and the B*07:12, 07:14, 07:18:01-07:18:02, 07:137, 35:08:01-35:08:02, 35:08:04-35:08:05, 35:18, 35:45, 35:61, 35:80, 35:100, 35:105, 35:142, 35:156, 35:158, 35:164, 35:172, 35:174, 35:176, 44:04, 44:28:01-44:28:02, 44:56N, 44:76, 53:11, 53:16, 53:24, 57:02:01-57:02:02, 57:04-57:05, 57:12-57:13, 57:19, 57:28N, 57:30 and 57:42 and in the C*01:50, C*03:28, 03:49, 03:117, C*04:80, 04:100, C*05:10 and C*15:08 alleles. Specific PCR fragment of 375 bp in the B*58:16 and 58:18 alleles. Specific PCR fragment of 55 and 225 bp in the B*35:08:03 and the C*01:49 alleles.

13Primer mix 18: Specific PCR fragment of 70 bp in the B*58:25 allele. Specific PCR fragment of 140 bp in the B*58:17N allele.

14Primer mix 20: Specific PCR fragment of 85 bp in the B*58:29 allele. Specific PCR fragment of 170 bp in the B*58:21 and the B*18:51 alleles.

15Primer mix 23: Specific PCR fragment of 150 bp in the B*58:26 and the B*35:75 and 44:138Q alleles. Specific PCR fragment of 260 bp in the B*58:31N allele.

‘?’, nucleotide sequence information not available for the primer matching sequence.

‘w’, might be weakly amplified.

1The B*5803 allele was never assigned.

2HLA-B*58 alleles in bold lettering are listed as confirmed alleles on the on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.7.0, January 2012.

3The B*58:09 and B*58:24 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 10.

4The B*58:21 and B*58:29 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 20.

5The B*58:36 and the B*15:20 as well as most B*35 and B*53 alleles give rise to identical amplification patterns with the HLA-B*58 subtyping kit. These alleles can be distinguished by e.g. the HLA-B low resolution kit and/or the HLA-B*15, HLA-B*35 and HLA-B*53 subtyping kits.

6Primer mix 4: Specific PCR fragment of 190 bp in the B*58:04 allele. Specific PCR fragment of 230 bp in the B*58:15 allele.

Primer mix 10: Specific PCR fragment of 115 bp in the B*58:24 and the C*15:18 alleles. Specific PCR fragment of 270 bp in the B*58:09 and the B*14:10, 18:22, 35:21, 35:24:01-35:24:02, 35:81, 35:96, 35:109, 35:157, 37:04:01-37:04:02, 40:28, 51:04, 51:46, 51:56:01-51:56:02, 53:02, 53:06 and 57:14 and in the C*15:39 alleles.

Primer mix 17: Specific PCR fragment of 225 bp in the B*58:22 and the B*07:12, 07:14, 07:18:01-07:18:02, 07:137, 35:08:01-35:08:02, 35:08:04-35:08:05, 35:18, 35:45, 35:61, 35:80, 35:100, 35:105, 35:142, 35:156, 35:158, 35:164, 35:172, 35:174, 35:176, 44:04, 44:28:01-44:28:02, 44:56N, 44:76, 53:11, 53:16, 53:24, 57:02:01-57:02:02, 57:04-57:05, 57:12-57:13, 57:19, 57:28N, 57:30 and 57:42 and in the C*01:50, C*03:28, 03:49, 03:117, C*04:80, 04:100, C*05:10 and C*15:08 alleles. Specific PCR fragment of 375 bp in the B*58:16 and 58:18 alleles. Specific PCR fragment of 55 and 225 bp in the B*35:08:03 and the C*01:49 alleles.

Primer mix 18: Specific PCR fragment of 70 bp in the B*58:25 allele. Specific PCR fragment of 140 bp in the B*58:17N allele.

Primer mix 20: Specific PCR fragment of 85 bp in the B*58:29 allele. Specific PCR fragment of 170 bp in the B*58:21 and the B*18:51 alleles.

Primer mix 23: Specific PCR fragment of 150 bp in the B*58:26 and the B*35:75 and 44:138Q alleles. Specific PCR fragment of 260 bp in the B*58:31N allele.

‘w’, might be weakly amplified.

‘?’, nucleotide sequence information not available for the primer matching sequence.