PHOTO DOCUMENT

Olerup SSP AB

Olerup SSPTM HLA-B*07

Prod. No:	101.512-24/03 – including Taq pol.

101.512-24u/03u– without Taq pol.

Lot No:    	Y40

Expiry Date:    	2009-September-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ug/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 32 primer mixes used for HLA-B*07 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified HLA-B*07 alleles3 | Other amplified HLA Class I alleles3,4 |
| --- | --- | --- | --- | --- |
| 15 | 400 bp | 800 bp | 070201-0707, 0709, 0710weakly, 0711, 0712, 0714-0726, 0728-0731, 0733-0736, 0739-0749N, 0751-0760 | 3910, 3916, 3917, 3920, 4201, 4202, 4204-4209, 4506, 670101-670102, 8102 |
| 2 | 285 bp | 1070 bp | 070201-0706, 0708-0711, 0713, 0715-0717, 0720-0732, 0734-0739, 0741-0752, 0754-0759 | 3566, 3707, 4015, 4016, 4031, 4032, 4045, 420501-420502, 480101-480102, 4805-4812, 4814-4816, 4818, 4819, 8101-8103 |
| 3 | 235 bp | 1070 bp | 070201-0704, 0707-0713, 0715, 0716, 071801-071802, 0720-0727, 0729-0731, 0735-0739, 0741-0752, 0754-0759 | 9538, 3566, 3707, 4031-4033, 4052, 4059, 4606, 4808 |
| 46,7 | 105 bp | 800 bp | 070201-0703, 070501-071802, 0720-0724, 0727-0732, 0735-0739, 0741, 0742, 0744-0747, 0749N-0752, 0754, 0756-0759 | 3566, 3707, 4015, 4016, 4023, 4032, 4805, 4808 |
| 57,8 | 220 bp | 1070 bp | 070201-070207, 0704-0707, 0709-0715, 0717-0726, 0728, 0730, 0731, 0733-0736, 0739-0749N, 0751-0760 | 2701-270402, 270502-2711, 2713-2715, 2717, 2719-2721, 2724-2728, 2730-2738, 4201, 4202, 4204-4206, 4208, 4209, 5402, 5510, 670101-6702, 7301, 8101-8103 |
| 67,8 | 215 bp | 1070 bp | 0703, 0708, 0716, 0727, 0732, 0738, 0750 | 1314, 1401, 1407N, 1568, 1571, 180101-1803, 1805-1815, 1817N-1824, 1826, 1827, 2712, 2716, 2718, 2723, 2729, 3550, 370101-3713, 380101-3804, 3806-3817, 39010101-39010102L, 390103-3903, 3905-390602, 3907?, 3908-3911, 391301-3920, 3922-3943, 400201-400602, 4008, 4009, 4011, 4013, 4018-4020, 4024, 4026-4029, 4035, 4037, 4039, 4040, 4044, 4050, 4056, 4057, 4064, 4068, 4070, 4071, 4074, 4075, 4078, 4207, 480101-4819 |
| 76,9 | 110, 215 bp | 1070 bp | 0704, 0719, 0725, 0729, 0752 | 080101-0805, 0807-0812, 0814-0819N, 0821-0824, 0826-0830N, 0832-0834, 4101-4105, 4107, 4108, 4201, 4202, 4204-4207, 4209, 5520 |
| 810 | 160, 235 bp | 1070 bp | 070501-0706, 0732, 0734, 0759 | 400101-400105, 4007, 4010, 4012, 401401-4016, 4021-4023, 4025, 4030, 4034, 4036, 4038, 4043, 4045-4049, 4054, 4055, 4061-4063, 4065-4067, 4069, 4072-4074, 4076, 4077, 4108, 420501-420502, 4431, 480101-480102, 480301-480302, 4805-4807, 4809-4812, 4814-4816, 4818, 4819, 8101-8103 |
| 97,8,11 | 240, 365, 390 bp | 1070 bp | 070501-070503, 0749N, 0758 |  |
| 1012 | 285, 415 bp | 800 bp | 0707, 0712, 0714, 071801-071802, 0747 | 9538, 400101-400105, 4007, 4010, 4012, 401401-401403, 4021-4023, 4025, 4033, 4036, 4038, 4042, 4043, 4046-4049, 4052, 4054, 4055, 4059-4063, 4065-4067, 4069, 4072-4074, 4076, 4077, 4108, 4431, 4606, 480301-480302, 4817, Cw*0738 |
| 11 | 210 bp | 1070 bp | 0708, 0710, 0732, 0738 | 1807, 2701-270402, 270502-270508, 270510-2711, 2713-2715, 2717, 2719-2721, 2723-2728, 2730-2738, 3550, 3806, 3807, 4008, 4013, 4068, 4806, 7301 |
| 12 | 230 bp | 1070 bp | 0709, 0711, 0717 | 0828, 9538, 3566, 4038, 4052, 4059, 4060, 4209, 4606, 4814, Cw*0738 |
| 136,7,13 | 85, 290 bp | 1070 bp | 0712, 0714, 071801-071802, 0731, 0734, 0743 | 4030, 4034, 4059, 4108, 420501-420502, 4817, Cw*0738 |
| 146,14 | 110, 395 bp | 1070 bp | 0713, 0726, 0734, 0743, 0748 | 0813, 0825, 9538, 2736, 400101-400105, 4007, 4010, 4012, 401401-401403, 4021, 4022N, 4025, 4030, 4031, 4033, 4034, 4036, 4038, 4042, 4043, 4045-4049, 4052, 4054, 4055, 4059-4063, 4065-4067, 4069, 4072-4074, 4076, 4077, 4431, 4606, 480101-480102, 480301-480302, 4806, 4807, 4809-4812, 4814-4818, 8101-8103, Cw*0738 |
| 1515 | 385, 425 bp | 1070 bp | 0711, 0713, 0715, 0754, 0757 | 0806, 0815, 0817, 3920, 3927, 4073, 4506, 6702, 7301 |
| 168 | 230 bp | 1070 bp | 0716, 0727, 0737, 0738, 0750 | 1314, 1568, 1571, 180101, 180103, 1802, 1805-1815, 1817N-1824, 1826, 1827, 2701, 2718, 2723, 2729, 3550, 370101-3713, 380101-380202, 3804, 3806-3812, 3815-3817, 3905, 3907?, 3908, 3911, 391301-391302, 3920, 3937, 400201-400602, 4008, 4009, 4011, 4013, 4018-4020, 4024, 4026-4029, 4035, 4037, 4039, 4040, 4044, 4056, 4057, 4064, 4068, 4070, 4071, 4074, 4075, 4078, 480101-4819, 5710 |
| 177,16 | 345, 410 bp | 1070 bp | 0710, 0711, 0727, 0746, 0757 | 0806, 0817, 3817, 3920, 7301 |
| 187,17 | 265, 385, 435 bp | 1070 bp | 0712, 0736, 0738, 0754 | 0803, 3566, 380101-380102, 3805-3807, 3809-3814, 3816, 4059, 4060, 4418, 4506, 4817, 4818, 4901, 4903-4905, Cw*0738 |
| 197,18 | 205, 250, 295 bp | 800 bp | 0720, 0735 | 1514, 1591, 9531, 3545, 3571, 4410, 4414, 4417, 4431, 4443, 4445, 4451, 4454, 5123, 5707, 5807, 8201, 8202 |
| 2019 | 160, 205, 275 bp | 1070 bp | 0721, 0751, 0759 |  |
| 2120 | 100, 165 bp | 1070 bp | 0722, 0752 |  |
| 2217 | 380 bp | 1070 bp | 0723 |  |
| 236 | 85 bp | 1070 bp | 0720, 0724 | 4063 |
| 246 | 125 bp | 1070 bp | 0725, 0739, 0748 | 2736, 3566, 400101-400105, 4007, 4010, 4012, 401401-401403, 4021, 4022N, 4025, 4030, 4031, 4033, 4034, 4036, 4038, 4042, 4043, 4045-4049, 4052, 4054, 4055, 4059-4063, 4065, 4066, 4069, 4072-4074, 4076, 4077, 4108, 420501-420502, 4431, 4606, 480101-480102, 480301-480302, 4806, 4807, 4809-4812, 4814, 4816-4819, 8101-8103 |
| 257,21 | 220, 425 bp | 800 bp | 0720, 0724, 0727, 0745, 0760 | 0821, 0825, 1507, 1545, 1555, 1568, 9526, 3505, 3522, 3551, 3558, 3572, 3817, 4005, 4071, 4454, 4612, Cw*0305, Cw*0325, Cw*0327 |
| 266, 22 | 95, 220, 290 bp | 800 bp | 0720, 0724, 0728, 0742, 0751, 0760 | 0814, 0821, 0825, 1507,1545, 1555, 1568, 9526, 2732, 3505, 3522, 3551, 3558, 3572, 4005, 4071, 4454, 4612, Cw*0305, Cw*0325, Cw*0327, Cw*0808 |
| 2723 | 165, 215 bp | 800 bp | 0729, 0741, 0756 | 1576, 9501, 3576, 5401-5406, 5408N-5413, 550101-5505, 5507-5509, 5511-5517, 5519-5528, 5601-5604, 560502, 5607, 5609-5613, 5615, 5618-5621 |
| 286,7,24 | 100, 170bp | 800 bp | 0730, 0756 | 1576, 9501, 3576, 5401-5406, 5408N-5413, 550101-5505, 5507-5509, 5511-5517, 5519-5528, 5601-5604, 560502, 5607, 5609-5613, 5615, 5618-5621 |
| 296,7,25 | 95, 305 bp | 800 bp | 0731, 0734, 0743, 0744, 0758 | 4108, 420501-420502, Cw*0738 |
| 3020 | 190 bp | 800 bp | 070501-0706, 0732, 0740, 0753 | 1301-130204, 1307N-1309, 1311, 1313-1321, 2707, 2711, 2720, 2724, 2727, 2733, 3531, 400101-400203, 4004, 40060101-4008, 4010-4016, 4019, 4021-4023, 4025, 4027, 4029, 4030, 4034-4038, 4040, 4043, 4045-4050, 4053-4058, 4061, 4062, 4064-4070, 4072-4078, 4431, 480101-480102, 480301-4807, 4809-4816, 4818, 4819, 5110, 5116, 5131, 5134, 5208, 5509, 5522, 5524, 8101-8103 |
| 3126 | 215, 280 bp | 800 bp | 0719, 0733, 0753, 0755, 0760 | 1405, 1507, 1545, 1555, 1568, 9526, 1814, 2707, 2711, 2724, 2732-2734, 3505, 3522, 3531, 3551, 3558, 3572, 3709, 3903, 3914, 3924, 3929, 3937, 400201-4003, 4005, 4008, 4009, 4013, 4018, 4019, 4024, 4027, 4029, 4035, 4037, 4039, 4040, 4050, 4056-4058, 4071, 4078, 4208, 4454, 4612, 4804, 4813, 4819, 5504, Cw*0305, Cw*0325, Cw*0327, Cw*0808 |
| 3227 | 190 bp | 1070 bp | 0734 | 080101-0806, 0808N-0810, 0812, 0813, 0815-0820, 0822-0824, 0826, 0827, 0829,-0834, 1306, 1542, 1583, 1586, 1826, 3560, 3712, 380101, 380201-3807, 3809-3817, 39010101-39010102L, 390103-390602, 3908-3914, 3916, 3918-3920, 3922-3924, 3926-3928, 3930, 3931, 3934-3937, 3939, 3941, 3942, 4039, 4051, 4101-410302, 4105-4108, 4201, 4202, 4204-4208, 5121, 5136, 5401, 5402, 5404, 5405N, 5407, 5408N, 5410-5413, 550101-5505, 5507, 5510-5513, 5515-5521, 5525, 5526, 5528, 5610, 5612, 5901, 5902, 670101-6702 |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-B*07 SSP subtypings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 bp or more. Size differences shorter than 20 bp are not given. For high resolution SSP kits the length of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers. PCR fragments longer than the control band may sometimes be observed. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most tubes, or a band of 800 base pairs, for some tubes.

Tube number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-B*07 subtyping.

In addition, tubes number 4, 10, 19, 25, 26, 27, 28, 29, 30 and 31 contain the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3For several HLA-B alleles 1st exon nucleotide sequences are not available. In these instances it is not known whether some of the primers of the SSP sets are completely matched with the target sequences or not. We assume that unknown sequences in the 1st exon are conserved within allelic groups.

4Due to the sharing of sequence motifs between HLA-B alleles non-HLA-B*07 alleles will be amplified by primer mixes 1 to 8, 10 to 19, 24, 26 and 29 to 32. In addition, the Cw*0305, Cw*0325 and Cw*0327 alleles will be amplified by primer mixes 35, 36 and 31, the Cw*0738 allele will be amplified by primer mixes 10, 12, 13, 14, 18 and 29 and the Cw*0808 allele will be amplified by primer mixes 26 and 31.

5Primer mix 1 may give rise to a weak false band of approx. 600 base pairs.

6Specific PCR fragments shorter than 125 base pairs are less intense and not as sharp as longer specific bands.

7Primer mixes 4, 5, 6, 9, 13, 17, 18, 19, 25, 28 and 29 may give rise to a primer oligomer artefact. Most pronounced for primer mixes 17 and 28.

8Primer mixes 5, 6, 9 and 16 may yield somewhat less specific PCR product than the other HLA-B*07 primer mixes.

9Primer mix 7: Specific PCR fragment of 110 bp in the B*0704, B*0719, B*0725, B*0753 and the B*080101-0805, B*0807-0812, B*0814-0819N, B*0821-0824, B*0826-0830N, B*0832-0834, B*4101-4105, B*4107, B*4108, B*4201, B*4202, B*4204 to B*4207, B*4209 and B*5520 alleles. Specific PCR fragment of 215 bp in the B*0729 allele.

10Primer mix 8: Specific PCR fragment of 160 bp in the B*0759 allele. Specific PCR fragment of 235 bp in the B*070501-0706, B*0732, B*0734 and the B*400101-400105, B*4007, B*4010, B*4012, B*401401-4016, B*4021-4023, B*4025, B*4030, B*4034, B*4036, B*4038, B*4043, B*4045-4049, B*4054, B*4055, B*4061-4063, B*4065-4067, B*4069, B*4072-4074, B*4076, B*4077, B*4108, B*420501-420502, B*4431, B*480101-480102, B*480301-480302, B*4805-4807, B*4809-4812, B*4814-4816, B*4818, B*4819 and B*8101-8103 alleles.

11Primer mix 9: Specific PCR fragment of 240 bp in B*0749 allele. Specific PCR fragment of 365 bp in the B*070501-070503 alleles. Specific PCR fragment of 390 bp in the B*0758 allele.

12Primer mix 10: Specific PCR fragment of 285 bp in the B*0707, B*0712, B*0714, B*071801-071802 and the B*400101-400105, B*4007, B*4010, B*4012, B*401401-401403, B*4021-4023, B*4025, B*4033, B*4036, B*4038, B*4042, B*4043, B*4046-4049, B*4052, B*4054, B*4055, B*4059-4063, B*4065-4067, B*4069, B*4072-4074, B*4076, B*4077, B*4108, B*4431, B*4606, B*480301-480302, B*4817 and Cw*0738 alleles. Specific PCR fragment of 415 bp in B*0747 allele.

13Primer mix 13: Specific PCR fragment of 85 bp in the B*0731, B*0734, B*0743 and the B*4108, B*420501-420502 and Cw*0738 alleles. Specific PCR fragment of 290 bp in the B*0712, B*0714, B*071801-071802 and the B*4030, B*4034, B*4059 and B*4817 alleles.

14Primer mix 14: Specific PCR fragment of 110 bp in the B*0726, B*0734, B*0743, B*0748 and the B*0813, B*0825, B*9538, B*2736, B*400101-400105, B*4007, B*4010, B*4012, B*401401-140103, B*4021, B*4022N, B*4025, B*4030, B*4031, B*4033, B*4034, B*4036, B*4038, B*4042, B*4043, B*4045-4049, B*4052, B*4054, B*4055, B*4059-4063, B*4065-4067, B*4069, B*4072-4074, B*4078, B*4077, B*4431, B*4606, B*480101-480102, B*480301-480302, B*4806, B*4807, B*4809-4812, B*4814-4818, B*8101-8103 and Cw*0738 alleles. Specific PCR fragment of 395 bp in the B*0713 allele.

15Primer mix 15: Specific PCR fragment of 385 bp in the B*0754 and the B*4506 alleles. Specific PCR fragment of 425 bp in the B*0711, B*0713, B*0715, B*0757 and the B*0806, B*0815, B*0817, B*3920, B*3927, B*4073, B*6702 and B*7301 alelles.

16Primer mix 17: Specific PCR fragment of 345 bp in the B*0746 alleles.Specific PCR fragment of 410 bp in the the B*0710, B*0711, B*0727, B*0757 and B*0806, B*0817, B*3817, B*3920 and B*7301 alleles.

17Primer mix 18: Specific PCR fragment of 265 bp in the B*0712 and the B*4059, B*4060 and Cw*0738 alleles. Specific PCR fragment of 385 bp in the B*0754 allele. Specific PCR fragment of 435 bp in the B*0736, B*0738 and the B*0803, B*3566, B*380101-380102 B*3805-3807, B*3809-3814, B*3816, B*4418, B*4506, B*4817, B*4818, B*4901 and B*4903-4905 alleles.

18Primer mix 19: Specific PCR fragment of 250 bp in the B*0720 and the B*1514, B*1591, B*9531, B*3545, B*3571, B*4410, B*4414, B*4417, B*4431, B*4443, B*4445, B*5123, B*5707, B*5807, B*8201 and B*8202. Specific PCR fragment of 250 bp in the B*4454 allele. Specific PCR fragment of 290 bp in the B*0735 allele.

19Primer mix 20: Specific PCR fragment of 160 bp in the B*0759 allele. Specific PCR fragment of 205 bp in the B*0751 allele. Specific PCR fragment of 275 bp in the B*0721 allele.

20Primer mix 21: Specific PCR fragment of 100 bp in the B*0752 allele. Specific PCR fragment of 165 bpö in the B*0722 allele.

21Primer mix 25: Specific PCR fragment of 220 bp in the B*0720, B*0724, B*0745, B*0760 and the  B*0821, B*0825, B*1507, B*1545, B*1555, B*1568, B*9526, B*3505, B*3522, B*3551, B*3558, B*3572, B*4005, B*4071, B*4454, B*4612, Cw*0305, Cw*0325 and Cw*0327 alleles. Specific PCR fragment of 425 bp in the B*0727 and B*3817 alleles.

22Primer mix 26: Specific PCR fragment of 95 bp in the B*0728 and B*0814 and B*2732 alleles. Specific PCR fragment of 220 bp in the B*0720, B*0724, B*0751, B*0760 and the  B*0821, B*0825, B*1507, B*1545, B*1555, B*1568, B*9526, B*3505, B*3522, B*3551, B*3558, B*3572, B*4005, B*4071, B*4454, B*4612, Cw*0305, Cw*0325, Cw*0327 and Cw*0808 alleles. Specific PCR fragment of 290 bp in the B*0742 allele.

23Primer mix 27: Specific PCR fragment of 165 bp in the B*0741, B*0756 and the B*1576, B*9501, B*3576, B*5401-5406, B*5408N-5413, B*550101-5505, B*5507-5509, B*5511-5517, B*5519-5528, B*5601-5604, B*560502, B*5607, B*5609-5613, B*5615, B*5618-5621 alleles. Specific PCR fragment of 215 bp in the B*0729 allele.

24Primer mix 28: Specific PCR fragment of 100 bp in the B*0730 allele. Specific PCR fragment of 170 bp in the B*0756 and the B*1576, B*9501, B*3576, B*5401-5406, B*5408N-5413, B*550101-5505, B*5507-5509, B*5511-5517, B*5519-5528, B*5601-5604, B*560502, B*5607, B*5609-5613, B*5615, B*5618-5621 alleles.

25Primer mix 29: Specific PCR fragment of 95 bp in the B*0731, B*0734, B*0743, B*0744 and the B*4108, B*420501-420502 and Cw*0738 alleles. Specific PCR fragment of 305 bp in the B*0758 allele.

26Primer mix 31: Specific PCR fragment of 215 bp in the B*0755 and the B*1507, B*1545, B*1568, B*9526, B*4612 and B*4819 alleles. Specific PCR fragment of 280 bp in the B*0719, B*0733, B*0753, B*0760 and the B*1405, B*1555, B*1814, B*2707, B*2711, B*2724, B*2732-2734, B*3505, B*3522, B*3531, B*3551, B*3558, B*3572, B*3709, B*3903, B*3914, B*3924, B*3929, B*3937, B*400201-4003, B*4005, B*4008, B*4009, B*4013, B*4018, B*4019, B*4024, B*4027, B*4029, B*4035, B*4037, B*4039, B*4040, B*4050, B*4056-4058, B*4071, B*4078, B*4208, B*4454, B*4804, B*4813, B*5504, Cw*0305, Cw*0325, Cw*0327 and Cw*0808 alleles.

27Primer mix 32 may give rise to a faint band in the B*1401 to 1407N, B*3701, B*3703-3712, B*3917, Cw*0103, Cw*04010101 to 040102, Cw*0403 to 0407, Cw*0409  to 0417, Cw*0501, C*0503, Cw*0505 to 0512, Cw*0609, Cw*080101 to 0805, Cw*0807, Cw*0808, Cw*0810, Cw*0812 and Cw*1801 to 1802.

‘?’, nucleotide sequence information not available for the primer matching sequence.

4The sequence of B*0701 has been shown to be in error.

5The B*0750, B*3707 and B*4808 alleles will give rise to identical amplification patterns. These alleles can e.g. be distinguished by the HLA-B low resolution primer set.

6Primer mix 7: Specific PCR fragment of 110 bp in the B*0704, B*0719, B*0725, B*0753 and the B*080101-0805, B*0807-0812, B*0814-0819N, B*0821-0824, B*0826-0830N, B*0832-0834, B*4101-4105, B*4107, B*4108, B*4201, B*4202, B*4204 to B*4207, B*4209 and B*5520 alleles. Specific PCR fragment of 215 bp in the B*0729 allele.

Primer mix 8: Specific PCR fragment of 160 bp in the B*0759 allele. Specific PCR fragment of 235 bp in the B*070501-0706, B*0732, B*0734 and the B*400101-400105, B*4007, B*4010, B*4012, B*401401-4016, B*4021-4023, B*4025, B*4030, B*4034, B*4036, B*4038, B*4043, B*4045-4049, B*4054, B*4055, B*4061-4063, B*4065-4067, B*4069, B*4072-4074, B*4076, B*4077, B*4108, B*420501-420502, B*4431, B*480101-480102, B*480301-480302, B*4805-4807, B*4809-4812, B*4814-4816, B*4818, B*4819 and B*8101-8103 alleles.

Primer mix 9: Specific PCR fragment of 240 bp in B*0749 allele. Specific PCR fragment of 365 bp in the B*070501-070503 alleles. Specific PCR fragment of 390 bp in the B*0758 allele.

Primer mix 10: Specific PCR fragment of 285 bp in the B*0707, B*0712, B*0714, B*071801-071802 and the B*400101-400105, B*4007, B*4010, B*4012, B*401401-401403, B*4021-4023, B*4025, B*4033, B*4036, B*4038, B*4042, B*4043, B*4046-4049, B*4052, B*4054, B*4055, B*4059-4063, B*4065-4067, B*4069, B*4072-4074, B*4076, B*4077, B*4108, B*4431, B*4606, B*480301-480302, B*4817 and Cw*0738 alleles. Specific PCR fragment of 415 bp in B*0747 allele.

Primer mix 13: Specific PCR fragment of 85 bp in the B*0731, B*0734, B*0743 and the B*4108, B*420501-420502 and Cw*0738 alleles. Specific PCR fragment of 290 bp in the B*0712, B*0714, B*071801-071802 and the B*4030, B*4034, B*4059 and B*4817 alleles.

Primer mix 14: Specific PCR fragment of 110 bp in the B*0726, B*0734, B*0743, B*0748 and the B*0813, B*0825, B*9538, B*2736, B*400101-400105, B*4007, B*4010, B*4012, B*401401-140103, B*4021, B*4022N, B*4025, B*4030, B*4031, B*4033, B*4034, B*4036, B*4038, B*4042, B*4043, B*4045-4049, B*4052, B*4054, B*4055, B*4059-4063, B*4065-4067, B*4069, B*4072-4074, B*4078, B*4077, B*4431, B*4606, B*480101-480102, B*480301-480302, B*4806, B*4807, B*4809-4812, B*4814-4818, B*8101-8103 and Cw*0738 alleles. Specific PCR fragment of 395 bp in the B*0713 allele.

Primer mix 15: Specific PCR fragment of 385 bp in the B*0754 and the B*4506 alleles. Specific PCR fragment of 425 bp in the B*0711, B*0713, B*0715, B*0757 and the B*0806, B*0815, B*0817, B*3920, B*3927, B*4073, B*6702 and B*7301 alelles.

Primer mix 17: Specific PCR fragment of 345 bp in the B*0746 alleles.Specific PCR fragment of 410 bp in the the B*0710, B*0711, B*0727, B*0757 and B*0806, B*0817, B*3817, B*3920 and B*7301 alleles.

Primer mix 18: Specific PCR fragment of 265 bp in the B*0712 and the B*4059, B*4060 and Cw*0738 alleles. Specific PCR fragment of 385 bp in the B*0754 allele. Specific PCR fragment of 435 bp in the B*0736, B*0738 and the B*0803, B*3566, B*380101-380102 B*3805-3807, B*3809-3814, B*3816, B*4418, B*4506, B*4817, B*4818, B*4901 and B*4903-4905 alleles.

Primer mix 19: Specific PCR fragment of 250 bp in the B*0720 and the B*1514, B*1591, B*9531, B*3545, B*3571, B*4410, B*4414, B*4417, B*4431, B*4443, B*4445, B*5123, B*5707, B*5807, B*8201 and B*8202. Specific PCR fragment of 250 bp in the B*4454 allele. Specific PCR fragment of 290 bp in the B*0735 allele.

Primer mix 20: Specific PCR fragment of 160 bp in the B*0759 allele. Specific PCR fragment of 205 bp in the B*0751 allele. Specific PCR fragment of 275 bp in the B*0721 allele.

Primer mix 21: Specific PCR fragment of 100 bp in the B*0752 allele. Specific PCR fragment of 165 bpö in the B*0722 allele.

Primer mix 25: Specific PCR fragment of 220 bp in the B*0720, B*0724, B*0745, B*0760 and the  B*0821, B*0825, B*1507, B*1545, B*1555, B*1568, B*9526, B*3505, B*3522, B*3551, B*3558, B*3572, B*4005, B*4071, B*4454, B*4612, Cw*0305, Cw*0325 and Cw*0327 alleles. Specific PCR fragment of 425 bp in the B*0727 and B*3817 alleles.

Primer mix 26: Specific PCR fragment of 95 bp in the B*0728 and B*0814 and B*2732 alleles. Specific PCR fragment of 220 bp in the B*0720, B*0724, B*0751, B*0760 and the  B*0821, B*0825, B*1507, B*1545, B*1555, B*1568, B*9526, B*3505, B*3522, B*3551, B*3558, B*3572, B*4005, B*4071, B*4454, B*4612, Cw*0305, Cw*0325, Cw*0327 and Cw*0808 alleles. Specific PCR fragment of 290 bp in the B*0742 allele.

Primer mix 27: Specific PCR fragment of 165 bp in the B*0741, B*0756 and the B*1576, B*9501, B*3576, B*5401-5406, B*5408N-5413, B*550101-5505, B*5507-5509, B*5511-5517, B*5519-5528, B*5601-5604, B*560502, B*5607, B*5609-5613, B*5615, B*5618-5621 alleles. Specific PCR fragment of 215 bp in the B*0729 allele.

Primer mix 28: Specific PCR fragment of 100 bp in the B*0730 allele. Specific PCR fragment of 170 bp in the B*0756 and the B*1576, B*9501, B*3576, B*5401-5406, B*5408N-5413, B*550101-5505, B*5507-5509, B*5511-5517, B*5519-5528, B*5601-5604, B*560502, B*5607, B*5609-5613, B*5615, B*5618-5621 alleles.

Primer mix 29: Specific PCR fragment of 95 bp in the B*0731, B*0734, B*0743, B*0744 and the B*4108, B*420501-420502 and Cw*0738 alleles. Specific PCR fragment of 305 bp in the B*0758 allele.

Primer mix 31: Specific PCR fragment of 215 bp in the B*0755 and the B*1507, B*1545, B*1568, B*9526, B*4612 and B*4819 alleles. Specific PCR fragment of 280 bp in the B*0719, B*0733, B*0753, B*0760 and the B*1405, B*1555, B*1814, B*2707, B*2711, B*2724, B*2732-2734, B*3505, B*3522, B*3531, B*3551, B*3558, B*3572, B*3709, B*3903, B*3914, B*3924, B*3929, B*3937, B*400201-4003, B*4005, B*4008, B*4009, B*4013, B*4018, B*4019, B*4024, B*4027, B*4029, B*4035, B*4037, B*4039, B*4040, B*4050, B*4056-4058, B*4071, B*4078, B*4208, B*4454, B*4804, B*4813, B*5504, Cw*0305, Cw*0325, Cw*0327 and Cw*0808 alleles.

‘w’,  might be weakly amplified.

‘?’, nucleotide sequence information not available for the primer matching sequence.