Gel Documentation Form and Worksheet

HLA-C high resolution for frequent alleles

(101.707-12/12u) 	Lot No: 9D5	Expiry Date: 2019-01-01

Name:__________________

Sample ID:______________

DNA Conc.(ng/ul):_________

Test  Date:_______________

Tested By:_______________

Review Date:_______________

Reviewed By:_______________	_

Interpretation:_______________	                    Failed lanes: ____________                           Comments:_____________

Gel Picture

PHOTO DOCUMENT

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 3, 5, 10, 18, 21, 40, 44, 45, 57, 61, 62 and 86 may have tendencies of unspecific amplifications, most pronounced in primer mix 10.

Primer mixes 6, 19, 20, 36, 78 and 80 have a tendency to giving rise to primer oligomer formation.

Primer mixes 1, 10, 12, 17, 27, 53, 64, 72, 77, 82, 86, 87 and 93 may give rise to a lower yield of HLA-specific PCR product than the other HLA-C high resolution primer mixes.

Primer mixes 10, 44, 92 and 93 may give rise to a long unspecific amplification product. This band should be disregarded when interpreting the typing results.

Primer mix 64 may give rise to unspecific amplification in C*12 alleles.

Primer mix 19 may give rise to a lower yield of HLA-specific PCR product than the other HLA-C high resolution primer mixes for the C*03 alleles.

Primer mix 96 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 430 base pairs.

1HLA-C alleles listed on the IMGT/HLA web page 2016-January-19, release 3.23.0, www.ebi.ac.uk/imgt/hla.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3Primer mix 2: Specific PCR fragment of 105 bp in the C*04:05 and 07:54 alleles. Specific PCR fragment of 215 bp in the C*01:05, 01:77 05:111, 07:37, 07:216, 07:307, 14:08 and 16:06 and in the B*07:77 and 07:193 alleles.

Primer mix 3: Specific PCR fragment of 65 bp in the C*08:99 allele. Specific PCR fragment of 100 bp in the C*01:27 allele.  Specific PCR fragment of 200 bp in the C*01:06 allele. Specific PCR fragment of 225 bp in the C*07:06 and 07:18-07:19 alleles. Specific PCR fragment of 290 bp in the C*01:11 allele. Specific PCR fragment of 365 bp in the C*15:13 and 15:103 alleles.

Primer mix 4: Specific PCR fragment of 100 bp in the C*02:34 and 16:09 alleles. Specific PCR fragment of 200 bp in the C*01:07:01-01:07:02, 06:10, 06:43:01 and 14:24:02 alleles.

Primer mix 5: Specific PCR fragment of 130 bp in the A*26:55, A*29:27 and B*18:90 alleles.

Specific PCR fragment of 155 bp in the C*02:02:28, 03:03:29, 03:04:19, 03:113, 04:01:43, 05:01:17, 05:09:02, 05:17, 07:03, 07:17:01-07:17:02, 07:23, 07:140-07:141:02, 07:304, 07:307, 08:15:02, 08:72:02 and 12:13 and in the B*39:09:03 and B*58:01:17 alleles. Specific PCR fragment of 215 bp in the C*01:08 and 12:19 alleles. Specific PCR fragment of 300 bp in the C*07:31:01-07:31:02 and 07:177 alleles. Specific PCR fragment of 130 bp and 275 bp in the C*06:27 and 16:47 alleles. Specific PCR fragment of 155 bp and 300 bp in the C*06:02:23, 12:03:10 and 16:02:11 alleles.

Primer mix 6: Specific PCR fragment of 105 bp in the C*03:24 and 15:90 alleles. Specific PCR fragment of 150 bp in the C*01:20 allele.

Specific PCR fragment of 190 bp in the C*01:60, 07:05, 07:177 and 07:364 and in the B*27:36, B*48:21 and B*48:26 alleles. Specific PCR fragment of 210 bp in the C*01:04, 01:09 and 06:23 and in the B*15:129, 15:342, B*27:25 and B*27:115 alleles. Specific PCR fragment of 265 bp in the C*01:15 and 03:33 alleles.

Primer mix 8: Specific PCR fragment of 100 bp in the C*02:03, 02:16:02, 02:18, 03:03:22, 03:04:34 and 06:07 and in the B*27:34, B*35:01:30w, B*40:02:07, 40:06:02, B*51:01:52w, B*57:01:16w and 57:03:02w alleles. Specific PCR fragment of 135 bp in the C*02:20 allele. Specific PCR fragment of 595 bp in the C*07:179 allele.

Primer mix 10: Specific PCR fragment of 105 bp in the C*07:46 and 07:356 and in the B*56:08 and 56:14 alleles. Specific PCR fragment of 130 bp in the C*01:115, 02:07, 02:97, 16:10 and 16:67 alleles. Specific PCR fragment of 270 bp in the C*06:161, 08:06 and 16:33 alleles.

Primer mix 11: Specific PCR fragment of 75 bp in the C*02:08, 02:42, 03:11:01-03:11:02, 03:18:02, 03:37:01-03:37:02, 03:64:01, 03:301, 06:03:01, 07:314:02 12:03:23, 15:10:02-15:10:03 and 15:113 and in the B*58:74 alleles. Specific PCR fragment of 160 bp in the C*02:104, 03:46 03:248, 04:16 and 15:100 alleles. Specific PCR fragment of 505 bp in the C*07:06 and 07:343 alleles.

Primer mix 12: Specific PCR fragment of 70 bp in the C*07:53 and 07:216 alleles. Specific PCR fragment of 280 bp in the C*02:02:01-02:02:03, 02:02:05-02:02:29, 02:03, 02:04w, 02:05:01-02:13, 02:14:01w, 02:15-02:25Q, 02:26:02-02:32, 02:34-02:40:02, 02:42-02:86, 02:88-02:106, 02:108-02:112, 03:02:01-03:02:09, 03:02:11-03:02:14, 03:04:01:01-03:04:25, 03:04:27-03:10, 03:14-03:17, 03:19, 03:23-03:29, 03:32-03:38:02, 03:40:01-03:42, 03:44-03:48, 03:51, 03:54, 03:57:01-03:57:02, 03:60, 03:63-03:64:01, 03:65, 03:70-03:74, 03:77-03:78, 03:80, 03:82, 03:84, 03:87:01-03:87:02, 03:89-03:95, 03:98, 03:100-03:101, 03:104-03:111, 03:113-03:115, 03:117-03:118, 03:121N, 03:123, 03:125, 03:128-03:131, 03:134-03:136, 03:138-03:140, 03:143, 03:145-03:149, 03:153-03:155, 03:157, 03:159, 03:162-03:164, 03:169Q, 03:172-03:174, 03:178-03:181, 03:183-03:184:02, 03:186:01-03:186:02, 03:190-03:191, 03:193-03:194, 03:197-03:201N, 03:208N-03:213, 03:215-03:216, 03:218-03:219, 03:221-03:222, 03:224N-03:226, 03:232-03:236, 03:238-03:240, 03:244Q-03:250, 03:252, 03:255-03:261, 03:263, 03:265N-03:266, 03:269-03:270, 03:277N, 03:278w, 03:279-03:283, 03:286-03:287, 03:292-03:294, 03:296-03:303, 03:305-03:306, 03:309-03:311, 04:03:01-04:03:02, 04:06, 04:16, 04:80, 04:107, 04:147, 04:160, 04:171, 05:58:01, 06:03:01, 06:132:01-06:132:02, 07:96:01-07:96:02, 07:314:02, 12:03:23, 15:02:01:01-15:02:20, 15:02:22-15:05:08, 15:05:10-15:09, 15:10:02-15:11, 15:13, 15:15-15:19, 15:21-15:22, 15:24-15:35, 15:37-15:60, 15:62, 15:64-15:106, 15:108-15:121, 16:34 and 16:70 and in the B*40:164, B*51:01:28 and B*58:74 alleles.

Primer mix 16: Specific PCR fragment of 110 bp in the C*07:104N and 12:46N and in the B*15:272N and B*40:286N alleles. Specific PCR fragment of 240 bp in the C*03:53, 12:07 and 12:17 and in the B*35:310 alleles. Specific PCR fragment of 275 bp in the C*01:65, 03:05, 03:13:01-03:13:02, 03:17, 03:25, 03:27, 03:35, 03:71, 03:167, 03:178, 03:198, 03:249, 03:267, 03:271, 03:292, 03:296, 03:302, 04:212, 05:86, 07:133, 07:242 and 07:330 and in the B*15:96 alleles. Specific PCR fragment of 370 bp in the C*15:22, 15:37, 15:55, 15:58, 15:65, 15:72 and 15:102 alleles.

Primer mix 17: Specific PCR fragment of 110 bp in the C*01:118, 03:08, 03:29, 03:246, 06:44, 07:01:01:01-07:01:24, 07:01:26-07:01:52, 07:06-07:07, 07:09, 07:16, 07:18-07:20, 07:22, 07:24, 07:26:01-07:26:02, 07:28, 07:30, 07:35-07:36, 07:40, 07:44, 07:52-07:53, 07:55N, 07:57-07:59, 07:65, 07:70-07:71, 07:73, 07:77-07:78:02, 07:81-07:83, 07:86, 07:89, 07:91-07:96:02, 07:98N, 07:103-07:104N, 07:106, 07:108:01-07:108:02,  07:110-07:113, 07:115-07:116, 07:118-07:120, 07:122, 07:124, 07:128-07:129, 07:131:01-07:132, 07:134, 07:140-07:141:02, 07:148, 07:150Q-07:151, 07:153, 07:162, 07:164N-07:166, 07:170, 07:173, 07:176, 07:179-07:180, 07:182, 07:188-07:191N, 07:196-07:197, 07:200-07:201, 07:203-07:207, 07:210, 07:212, 07:214-07:215, 07:219, 07:222-07:224, 07:227N-07:228, 07:230-07:231, 07:235Q-07:237, 07:246-07:250, 07:253-07:257:04, 07:263, 07:266-07:269, 07:276-07:282, 07:292-07:293, 07:296-07:301, 07:303-07:304, 07:310-07:311, 07:313, 07:317, 07:330-07:332, 07:337, 07:343, 07:353, 07:362, 07:366, 07:369-07:371, 07:377, 07:385-07:387, 07:404, 07:407-07:408, 07:410, 07:412, 07:417, 07:419, 07:421-07:422, 07:424, 07:427, 07:430-07:431, 07:435, 07:437N, 07:440-07:444, 07:448, 07:450, 07:453-07:455, 07:458, 07:460-07:463, 07:468-07:471, 07:475-07:476N, 07:479, 07:481, 14:73, 15:02:10, 15:02:17, 15:99 and 18:05 and in the B*07:259 alleles. Specific PCR fragment of 180 bp in the C*01:51, 03:03:01-03:03:25, 03:03:27-03:03:29, 03:11:01-03:11:02, 03:13:01-03:13:02, 03:18:01-03:18:02, 03:20N-03:22Q, 03:30, 03:39, 03:43:01-03:43:02, 03:49-03:50, 03:52-03:53, 03:55-03:56, 03:58-03:59, 03:61-03:62, 03:66-03:69, 03:75-03:76, 03:79, 03:81, 03:83, 03:85-03:86, 03:88, 03:96-03:97, 03:102-03:103, 03:112, 03:116:01-03:116:02, 03:119-03:120, 03:122, 03:124, 03:126-03:127, 03:132-03:133, 03:141-03:142, 03:144, 03:150-03:152, 03:158, 03:160-03:161, 03:165, 03:167-03:168, 03:171, 03:175-03:177, 03:182, 03:185, 03:187-03:189N, 03:192, 03:195-03:196, 03:202-03:207, 03:214, 03:217, 03:220, 03:223, 03:227-03:231, 03:237, 03:241-03:243, 03:251, 03:253-03:254, 03:262, 03:267-03:268, 03:271-03:276, 03:284-03:285, 03:288-03:291, 03:295, 03:304, 03:307-03:308, 03:312, 06:53:01, 07:56:01, 07:79,14:46:02 and 17:26 and in the B*07:93, B*08:39, B*35:101:02, 35:195 and 35:211, B*38:27, B*51:76 and B*54:34 alleles.

Specific PCR fragment of 565 bp in the C*01:02:18, 06:02:38, 07:447, 12:03:04, 12:03:09 and 14:02:08 and in the B*27:05:27 alleles. Specific PCR fragment of 110 bp and 180 bp in the C*03:31, 07:60 and 07:109 alleles.

Primer mix 18: Specific PCR fragment of 90 bp in the C*03:21, 03:142 and 03:287 and in the B*15:96 alleles. Specific PCR fragment of 120 bp in the C*03:14, 03:162, 03:220, 03:261, 07:133 and 07:330 alleles.  Specific PCR fragment of 185 bp in the C*07:21 allele.

Specific PCR fragment of 225 bp in the C*03:23 allele. Specific PCR fragment of 255 bp in the C*03:09, 03:22Q, 03:42, 03:71, 03:89, 03:116:01, 03:119, 03:139, 03:169Q, 03:199, 03:240, 03:249, 03:267, 03:292, 03:300, 04:212, 05:86, 07:36, 07:133, 07:242, 07:330 and 07:355 and in the B*56:37 alleles.

Primer mix 19: Specific PCR fragment of 115 bp in the C*03:14, 03:36, 03:77, 03:162, 03:187, 03:220, 03:261, 03:296, 04:212, 07:133 and 07:330 alleles. Specific PCR fragment of 265 bp in the C*03:13:01-03:13:02, 03:35, 03:178, 03:198, 03:267, 03:292, 04:212, 05:86, 07:133, 07:242 and 07:330 and in the B*15:96 alleles.

Specific PCR fragment of 305 bp in the C*04:80, 04:100, 04:178, 06:14, 06:143, 07:10, 07:43, 07:196, 15:25 and 15:62 and in the B*15:357, B*35:178 and 35:282 alleles. Specific PCR fragment of 540 bp in the C*03:15, 03:60, 04:212, 07:133, 07:242 and 07:330 alleles. Specific PCR fragment of 305 bp and 540 bp in the C*03:32, 03:45, 03:136 and 03:297 alleles.

Primer mix 20: Specific PCR fragment of 240 bp in the C*02:02:01-02:02:03, 02:02:05-02:02:29, 02:03, 02:05:01-02:13, 02:15-02:25Q, 02:26:02-02:26:03, 02:28-02:40:02, 02:42-02:64, 02:66-02:86, 02:88-02:106, 02:108-02:112, 03:07, 03:10, 03:15, 03:29, 03:45, 03:163, 03:268, 03:297, 04:03:01-04:03:02, 04:06, 04:16, 04:80, 04:107, 04:147, 04:160, 04:171, 05:13, 05:58:01, 06:03:01, 06:132:01-06:132:02, 15:02:01:01-15:02:20, 15:02:22-15:05:08, 15:05:10-15:06:03, 15:08-15:09, 15:10:02-15:13, 15:15-15:19, 15:21-15:22, 15:24, 15:26-15:35, 15:37-15:42, 15:44-15:60, 15:62, 15:64-15:84Q, 15:86-15:103, 15:105Q-15:106, 15:108-15:121 and 16:70 alleles. Specific PCR fragment of 265 bp in the C*01:14 allele. Specific PCR fragment of 240 bp and 265 bp in the C*15:104 allele.

Primer mix 21: Specific PCR fragment of 195 bp in the C*03:30, 03:65, 04:18, 04:52, 04:55, 05:55, 12:58, 14:10 and 15:12 alleles. Specific PCR fragment of 225 bp in the C*04:19 allele. Specific PCR fragment of 250 bp in the C*03:50 allele. Specific PCR fragment of 280 bp in the C*03:251, 04:120, 05:10, 05:64:01-05:64:02, 08:19, 08:44, 08:61, 08:82, 08:101 and 08:126 and in the B*44:148 and B*58:76 alleles.

Primer mix 22: Specific PCR fragment of 80 bp in the C*03:02:01-03:02:14, 03:15-03:16, 03:33, 03:36, 03:40:01, 03:40:03-03:40:04, 03:42-03:43:01, 03:60, 03:84, 03:89, 03:95, 03:108, 03:110, 03:121N, 03:132, 03:139, 03:146, 03:169Q, 03:175, 03:194, 03:197-03:201N, 03:216, 03:221-03:222, 03:224N-03:226, 03:238, 03:240, 03:245, 03:248, 03:258, 03:264, 03:279, 03:298-03:301, 04:54, 07:02:32, 07:242, 14:08-14:09, 14:28:01-14:28:02 and 15:77 and in the B*15:78:03 alleles.

Specific PCR fragment of 140 bp in the C*03:17, 03:249, 03:302, 04:37, 04:230, 06:34:02, 14:24:02, 14:26 and 14:55 alleles.  Specific PCR fragment of 80 bp and 140 bp in the C*01:97, 03:71, 03:271, 06:02:08, 12:03:20, 14:02:01, 14:02:03-14:07N, 14:10-14:16, 14:18-14:24:01, 14:27, 14:29-14:53, 14:56-14:73 and 16:01:06 alleles.

Primer mix 23: Specific PCR fragment of 150 bp in the C*01:14, 02:02:13, 02:02:29, 03:07, 03:10, 03:15, 03:29, 03:45, 03:163, 03:268, 03:297, 04:01:23, 05:01:12, 05:29:02, 06:02:01:01-06:02:01:03, 06:02:03-06:04:02, 06:06-06:10, 06:12-06:76:01, 06:77-06:81, 06:83-06:146, 06:148-06:173, 07:07, 07:09, 07:76:01-07:76:02, 07:315, 07:328, 07:406, 12:04:01, 14:12, 14:49, 15:02:10, 15:02:17, 15:99, 17:01:01:01-17:01:03, 17:01:05-17:21, 17:23-17:31 and 18:01-18:10 and in the B*15:277, B*35:222 and B*42:22 alleles. Specific PCR fragment of 240 bp in the C*05:13 allele.

Primer mix 26: Specific PCR fragment of 145 bp in the C*04:07, 04:27, 04:89, 04:153, 05:78, 07:64, 07:402 and 15:36 alleles. Specific PCR fragment of 170 bp in the C*02:09 allele.

Primer mix 27: Specific PCR fragment of 135 bp in the C*01:35, 01:107, 04:08, 04:34, 04:147, 04:212, 14:20 and 18:08 alleles. Specific PCR fragment of 175 bp in the C*04:30, 15:15 and 15:77 alleles. Specific PCR fragment of 220 bp in the C*04:09N allele.

Primer mix 29: Specific PCR fragment of 130 bp in the C*03:05, 03:13:02, 03:25, 03:27, 03:135, 03:167, 03:296, 04:15:01, 04:15:03 and 14:09 and in the B*15:78:03 alleles. Specific PCR fragment of 190 bp in the C*02:15 and in the B*07:221 alleles. Specific PCR fragment of 290 bp in the C*07:14 allele.

Primer mix 31: Specific PCR fragment of 150 bp in the C*01:13, 02:51, 03:87:01-03:87:02, 04:223, 05:09:01-05:09:03, 05:17, 05:52, 07:130, 08:15:01-08:15:02, 08:51, 08:73, 12:12, 12:144 and 16:27 and in the B*15:33, 15:248, B*35:08:02 and 35:08:05 alleles. Specific PCR fragment of 285 bp in the C*02:57, 04:40, 06:123 and 07:45 and in the B*38:23 and B*39:73 alleles.

Primer mix 32: Specific PCR fragment of 120 bp in the C*05:03 and 07:52 alleles. Specific PCR fragment of 180 bp in the C*03:48 allele. Specific PCR fragment of 210 bp in the C*16:07:01 allele.

Primer mix 33: Specific PCR fragment of 190 bp in the C*04:28 allele. Specific PCR fragment of 225 bp in the C*05:04:01-05:04:02, 05:103, 06:129, 07:68, 07:260, 07:302, 08:09, 08:11, 08:59 and 08:113 and in the B*15:337, B*18:83 and B*58:76 alleles.

Primer mix 34: Specific PCR fragment of 220 bp in the C*03:231w, 04:10, 04:11w, 04:36w, 04:55w, 04:153, 04:169, 04:210, 04:214w, 04:215, 05:78 and 15:36 alleles. Specific PCR fragment of 275 bp in the C*03:251, 04:27, 05:05, 05:07N, 08:08:01-08:08:02, 08:62 and 08:82 and in the A*02:425, 02:519, A*29:10, A*68:69, B*14:32, B*15:337, B*18:83 and B*44:148 alleles.

Primer mix 36: Specific PCR fragment of 85 bp in the C*01:65, 03:28, 03:49, 03:117, 05:06 and 05:86 alleles. Specific PCR fragment of 120 bp in the C*05:15 allele.

Specific PCR fragment of 160 bp in the C*03:25, 04:129, 05:42, 05:46, 06:67, 07:101, 07:148, 07:161, 08:05, 08:21, 12:127 and 17:05 alleles. Specific PCR fragment of 210 bp in the C*03:06:01-03:06:02 and C*03:26 alleles. Specific PCR fragment of 235 bp in the C*03:88 allele.

Primer mix 37: Specific PCR fragment of 80 bp in the C*07:50 allele. Specific PCR fragment of 130 bp in the C*05:23, 05:62, 07:04:01-07:04:10, 07:11, 07:63, 07:68, 07:101, 07:139, 07:142, 07:181, 07:199:01-07:199:02, 07:260, 07:272, 07:302, 07:323-07:324, 07:328-07:329N, 07:336, 07:338, 07:354-07:355, 07:357-07:358, 07:361, 07:365, 07:378, 07:394-07:395, 07:403, 07:406, 07:420, 07:426, 07:428, 07:459, 07:466-07:467, 07:480, 07:487, 08:07, 08:47 and 08:104 alleles. Specific PCR fragment of 195 bp in the C*05:14 allele. Specific PCR fragment of 235 bp in the C*05:19 allele. Specific PCR fragment of 280 bp in the C*05:16, 05:85, 05:107, 06:129 and 08:12 and in the B*14:32 allele. Specific PCR fragment of 130 bp and 280 bp in the C*07:364 allele.

Primer mix 43: Specific PCR fragment of 100 bp in the C*01:02:34, 01:21, 02:42, 02:107, 04:140, 04:166, 04:220, 05:98, 06:05, 07:02:09, 08:14, 08:80, 08:103, 12:16, 12:147, 15:63, 15:113 and 16:80 and in the B*67:02 alleles. Specific PCR fragment of 155 bp in the C*06:13 allele. Specific PCR fragment of 240 bp in the C*01:10, 01:51, 02:05:01-02:05:03, 02:17, 03:11:01-03:11:02, 03:39, 06:08, 06:53:01-06:53:02, 07:56:01-07:56:02, 07:60, 07:79, 07:109, 12:119, 14:25, 16:29, 16:50, 17:21, B*07:93, B*07:239, B*08:39, B*14:46, B*38:27, B*40:243 alleles. Specific PCR fragment of 275 bp in the C*06:27 and 16:47 alleles.

Primer mix 44: Specific PCR fragment of 90 bp in the C*03:20N allele. Specific PCR fragment of 135 bp in the C*16:08 allele. Specific PCR fragment of 165 bp in the C*02:22, 04:14, 04:68, 04:94:01-04:94:02, 05:08, 05:52, 05:89, 05:106, 05:112, 06:09, 06:144, 08:27, 08:29, 08:31, 12:31, 12:144, 16:45, 17:04 and 18:03 and in the A*24:96, 24:146 and B*15:137 alleles.

Primer mix 45: Specific PCR fragment of 140 bp in the C*06:11, 06:122w, 06:124w, 06:147, 07:01:13, 07:04:01-07:04:10, 07:11-07:12, 07:45, 07:63, 07:68, 07:101, 07:139, 07:142, 07:181, 07:199:01-07:199:02, 07:272, 07:302, 07:323-07:324, 07:329N, 07:338, 07:354-07:355, 07:358, 07:364-07:365, 07:378, 07:394-07:395, 07:420, 07:426, 07:428, 07:459, 07:466-07:467, 07:480, 07:487, 12:02:11, 12:11, 12:118 and 16:01:16 alleles. Specific PCR fragment of 255 bp in the C*03:47 and 12:126 alleles. Specific PCR fragment of 550 bp in the C*06:02:38 allele. Specific PCR fragment of 140 bp and 550 in the C*07:447 and 12:03:09 allele.

Primer mix 50: Specific PCR fragment of 235 bp in the C*03:03:29, 03:04:19, 03:113, 07:29:01-07:29:02, 07:177, 07:194 and 07:248 alleles. Specific PCR fragment of 305 bp in the C*03:03:29, 03:04:19, 03:113, 07:10, 07:43, 07:196 and 07:367 and in the B*35:01:38, 35:03:15, 35:08:04, 35:15:02, B*37:04:02, B*57:01:18 and B*58:01:17 alleles. Specific PCR fragment of 415 bp in the C*06:17, 07:07, 07:09, 07:49, 07:61N, 07:76:01-07:76:02, 07:210, 07:238, 07:247, 07:315, 07:328, 07:403, 07:406 and 18:01-18:10 alleles. Specific PCR fragment of 235 and 405 bp in the C*07:184 allele.

Primer mix 51: Specific PCR fragment of 100 bp in the C*07:32N and 07:67 alleles.  Specific PCR fragment of 235 bp in the C*01:16-01:18, 01:21, 01:23, 07:316 and 07:338 alleles. Specific PCR fragment of 265 bp in the C*01:19, 07:24 and 07:218 alleles.

Primer mix 53: Specific PCR fragment of 80 bp in the C*07:11 allele. Specific PCR fragment of 115 bp in the C*05:12 allele. Specific PCR fragment of 265 bp in the C*05:29:01-05:29:02, 08:13, 08:16:01, 08:25 and 08:94 allele.

Primer mix 55: Specific PCR fragment of 165 bp in the C*02:104, 03:248, 04:11, 04:33, 04:169, 04:172, 15:100 and 16:62 alleles. Specific PCR fragment of 220 bp in the C*07:26:01-07:26:02, 07:33N, 07:92, 07:96:01-07:96:02, 07:194, 07:314:01-07:314:02, 07:317, 07:351 and 18:09 alleles. Specific PCR fragment of 335 bp in the C*01:12:01-01:12:02, 01:34, 01:49:01-01:50, 01:72, 01:79:01-01:79:02 and 01:101-01:102 and in the A*01:21 alleles. Specific PCR fragment of 405 bp in the C*07:61N allele.  Specific PCR fragment of 465 bp in the C*07:07, 07:16, 07:51, 07:72, 07:177, 07:181, 07:328 and 07:367 alleles. Specific PCR fragment of 335 bp and 465 bp in the C*01:21 allele.

Primer mix 56: Specific PCR fragment of 100 bp in the C*07:32N allele. Specific PCR fragment of 130 bp in the C*02:02:28, 06:02:23, 07:22, 07:25, 07:137:01, 07:138, 07:404 and 12:03:10 and in the A*03:01:46, 03:154:02, A*11:32:02, A*24:02:62, A*25:27:02, A*33:03:09 and A*68:01:12 alleles. Specific PCR fragment of 250 bp in the C*07:35 and 07:352 alleles. Specific PCR fragment of 305 bp in the C*03:03:29, 03:04:19, 03:113, 07:10, 07:43, 07:184, 07:196, 07:367, B*35:01:38, 35:03:15, 35:08:04 and 35:15:02 and in the B*37:04:02, B*57:01:18 and B*58:01:17 alleles.

Primer mix 65: Specific PCR fragment of 150 bp in the C*12:06 allele. Specific PCR fragment of 415 bp in the C*01:118, 03:08, 03:29, 03:31, 03:246, 04:112, 04:169, 05:36, 06:44, 12:08, 12:81, 14:73, 16:64, 16:70 and 16:87 alleles.

Primer mix 73: Specific PCR fragment of 205 bp in the C*06:12 allele. Specific PCR fragment of 325 bp in the C*01:90, 01:101-01:102, 01:113, 01:116, 02:06:01-02:06:02, 02:23, 02:36, 02:68, 03:81, 03:175, 03:199, 03:245, 04:108, 04:178, 06:89, 07:123, 07:173, 07:294, 08:113, 12:15, 12:113, 15:02:01:01-15:02:17, 15:02:19-15:02:21, 15:02:23-15:03, 15:05:01-15:13, 15:15-15:19, 15:21-15:24, 15:26-15:29, 15:31-15:39, 15:41-15:63, 15:67-15:75, 15:78-15:101, 15:103-15:104, 15:106-15:109, 15:111-15:121 and 16:20 alleles.

Primer mix 74: Specific PCR fragment of 95 bp in the C*01:03, 01:24, 01:27, 01:78, 03:58, 04:37, 05:85 and 07:364 allele. Specific PCR fragment of 175 bp in the C*03:08, 03:29, 03:31, 03:246, 07:20, 07:96:01-07:96:02, 07:263, 15:02:01:01-15:09, 15:12-15:13, 15:15, 15:18-15:19, 15:21-15:24, 15:26, 15:28-15:37, 15:39-15:42, 15:44-15:62, 15:64-15:106, 15:108-15:112, 15:114-15:121 and 16:70 alleles.

Primer mix 75: Specific PCR fragment of 145 bp in the C*07:57 allele. Specific PCR fragment of 195 bp in the C*02:12, 02:49, 02:55:01-02:55:02, 03:15, 03:27, 03:38:01-03:38:02, 03:69, 03:130, 03:136, 03:163, 03:246w, 03:297, 04:03:01, 04:06, 04:16, 04:80, 04:107, 04:147, 04:160, 04:171, 04:190, 06:70:01-06:70:02, 06:132:01-06:132:02, 07:20, 07:96:01-07:96:02, 07:127, 07:263, 07:390w, 12:140w, 15:03, 15:16 and 15:25 alleles.

Primer mix 78: Specific PCR fragment of 195 bp in the C*06:12 and 07:349 alleles. Specific PCR fragment of 235 bp in the C*06:16N allele. Specific PCR fragment of 310 bp in the C*01:02:01-01:02:36, 01:05-01:20, 01:22-01:23, 01:25-01:33, 01:35-01:53, 01:56N-01:77, 01:79:01-01:89N, 01:91-01:96, 01:98N-01:101, 01:104-01:114, 01:117N-01:119, 02:19, 02:23, 03:03:01-03:11:02, 03:13:01-03:13:02, 03:17-03:18:02, 03:20N-03:25, 03:27-03:32, 03:34-03:35, 03:37:01-03:39, 03:41:01-03:41:02, 03:44-03:57:02, 03:59, 03:61-03:70, 03:72-03:83, 03:85-03:88, 03:90-03:94, 03:96-03:101, 03:103-03:107, 03:109, 03:111-03:118, 03:120, 03:122-03:131, 03:133-03:138, 03:140-03:145, 03:147-03:168, 03:170-03:174, 03:176-03:189N, 03:191-03:193, 03:195-03:196, 03:202-03:215, 03:217-03:220, 03:223, 03:227-03:237, 03:239, 03:241-03:244Q, 03:246-03:247, 03:249, 03:251-03:257, 03:259-03:263, 03:265N-03:266, 03:268-03:270, 03:272-03:278, 03:280-03:297, 03:302-03:306, 03:308-03:312, 04:89, 04:135, 04:210, 05:47, 05:107, 06:107, 07:29:01-07:29:02, 07:177, 07:184, 07:248, 14:54-14:55, 15:06:01-15:06:03, 15:18, 15:40, 15:55, 15:102 allele. Specific PCR fragment of 260 bp in the C*06:15 allele.

Primer mix 80: Specific PCR fragment of 70 bp in the C*03:34, 03:142, 03:261 and 03:272 alleles. Specific PCR fragment of 170 bp in the C*01:90w, 02:06:01-02:06:02, 02:47, 06:168, 12:15, 14:19, 15:08, 15:74 and in the A*26:41, B*07:78w, B*13:18w, 13:31w, 13:41w,13:73w, B*15:73w, 15:303w, B*54:10w, 54:20w, 54:33w, B*55:09w, 55:21w, 55:37w, 55:52w, B*56:43w alleles. Specific PCR fragment of 210 bp in the C*03:19, 03:44, 03:102 and 03:307 and in the B*56:37w alleles. Specific PCR fragment of 235 bp in the C*03:88 allele.

Primer mix 81: Specific PCR fragment of 65 bp in the C*07:06, 07:18-07:19 and 08:102 allele. Specific PCR fragment of 180 bp in the C*04:30 allele. Specific PCR fragment of 180 bp in the C*01:90, 02:06:01-02:06:02, 02:47, 03:19, 03:102, 03:307, 04:178, 06:168, 07:289, 12:15, 15:11, 15:16-15:17, 15:43 and in the B*46:11, 46:18, B*56:08 and 56:14 alleles.

Primer mix 83: Specific PCR fragment of 235 bp in the C*06:16N allele. Specific PCR fragment of 315 bp in the C*02:06:01-02:06:02, 02:47, 03:19, 03:102, 03:307, 04:178, 06:168, 07:289, 12:15, 15:02:01:01-15:02:08, 15:02:10-15:03, 15:07-15:08, 15:10:01-15:13, 15:15-15:18, 15:21, 15:26, 15:28, 15:31-15:35, 15:37-15:39, 15:41-15:45, 15:47-15:53, 15:56-15:58, 15:60, 15:62-15:63, 15:67-15:68, 15:71-15:75, 15:78-15:89, 15:91-15:101, 15:103, 15:106, 15:109, 15:112-15:114 and 15:118-15:121 alleles. Specific PCR fragment of 360 bp in the C*06:15 allele.

Primer mix 86: Specific PCR fragment of 150 bp in the C*07:30 allele. Specific PCR fragment of 320 bp in the C*07:72 allele. Specific PCR fragment of 415 bp in the C*07:39-07:40, 07:177, 07:210w, 07:238w, 07:328w and 07:335 alleles.

Primer mix 88: Specific PCR fragment of 85 bp in the C*04:23 and 04:108 alleles. Specific PCR fragment of 195 bp in the C*04:31 allele. Specific PCR fragment of 545 bp in the C*01:85, 08:22, 08:56, 08:102, 15:29 and 15:87 alleles.

Primer mix 90: Specific PCR fragment of 205 bp in the C*07:66 allele. Specific PCR fragment of 350 bp in the C*17:01:01:01-17:01:03, 17:01:05-17:01:11, 17:03, 17:04?, 17:05, 17:06?-17:16?, 17:17, 17:18?-17:29?, 17:30 and 17:31? alleles.

Primer mix 91: Specific PCR fragment of 70 bp in the C*17:02, 17:04?, 17:06?-17:16?, 17:18?-17:29? and 17:31? alleles. Specific PCR fragment of 105 bp in the C*07:67 allele. Specific PCR fragment of 170 bp in the C*07:48 allele. Specific PCR fragment of 245 bp in the C*04:01:43, 05:01:17, 05:09:02, 05:17, 07:13, 07:28, 07:55N, 07:289, 08:15:02 and 08:72:02 and in the B*35:03:15 and B*37:04:02 alleles. Specific PCR fragment of 270 bp in the C*07:47 and 07:461 alleles.

Primer mix 92: Specific PCR fragment of 155 bp in the C*03:135 and 04:20 alleles. Specific PCR fragment of 300 bp in the C*17:03, 17:04?, 17:06?-17:16?, 17:17, 17:18?-17:29? and 17:31? alleles. Specific PCR fragment of 485 bp in the C*07:419 allele.

Primer mix 93: Specific PCR fragment of 90 bp in the C*17:01:01:01-17:01:11, 17:04?, 17:05, 17:06?-17:16?, 17:18?-17:29?, 17:30, 17:31? alleles. Specific PCR fragment of 210 bp in the C*03:271, 14:06, 14:15 and 14:53 alleles. Specific PCR fragment of 290 bp in the C*14:11 allele.

4The following HLA-C alleles can be distinguished by the different sizes of the specific PCR product:

| Alleles | Primer mix | Alleles | Primer mix |
| --- | --- | --- | --- |
| C*01:06, 01:11 | 3 | C*06:08, 06:13 | 43 |
| C*01:15, 01:20, 01:60 | 6 | C*06:15-06:16N | 78, 83 |
| C*02:51, 02:57 | 31 | C*07:22, 07:35, 07:404 | 56 |
| C*03:06:01-03:06:02, 03:117 | 36 | C*07:25, 07:137:01, 07:138, 07:352 | 56 |
| C*03:09, 03:23, 03:287 | 18 | C*07:30, 07:40 | 86 |
| C*03:21-03:22Q, 03:116:01 | 18 | C*07:33N, 07:51 | 55 |
| C*03:34, 03:44 | 80 | C*07:37, 07:54 | 2 |
| C*03:37:01-03:37:02, 03:46 | 11 | C*07:47-07:48 | 91 |
| C*04:23, 04:31 | 88 | C*12:06,12:81 | 65 |
| C*05:10, 05:55 | 21 | C*12:07 ,12:46N | 16 |
| C*05:14, 05:16, 05:19 | 37 |

The HLA-C typing kit cannot distinguish the following silent mutations: the C*01:02:01-01:02:17, 01:02:19-01:02:31, 01:02:33 and 01:02:35-01:02:36 alleles, the C*01:12:01-01:12:02, 01:49:01-01:49:02 and 01:79:01-01:79:02 alleles, the C*02:02:01-02:02:03, 02:02:05-02:02:07, 02:02:09-02:02:11 and 02:02:14-02:02:20, 02:02:22-02:02:27 and 02:26:02-02:26:03 alleles, the C*02:02:30 and 02:26:01 alleles, the C*02:05:01-02:05:03 alleles, the C*02:06:01-02:06:02 alleles, the C*02:10:01:01-02:10:01:02 alleles, the C*02:27:01-02:27:02 alleles, the C*02:55:01-02:55:02 alleles, the C*03:02:01-03:02:07, 03:02:09 and 03:02:11-03:02:14 alleles, the C*03:03:01-03:03:20, 03:03:23-03:03:25 and 03:03:27-03:03:28 alleles, the C*03:04:01:01-03:04:14, 03:04:16-03:04:18, 03:04:20-03:04:24, 03:04:28-03:04:29, 03:04:31-03:04:33 and 03:04:35-03:04:49 alleles, the C*03:04:27 and 03:04:30 alleles, the C*03:06:01-03:06:02 alleles, the C*03:11:01-03:11:02 alleles, the C*03:37:01-03:37:02 alleles, the C*03:38:01-03:38:02 alleles, the C*03:40:01 and 03:40:03-03:40:04 alleles, the C*04:01:01:01-04:01:02, 04:01:04, 04:01:06, 04:01:08, 04:01:10-04:01:17, 04:01:19-04:01:21, 04:01:24, 04:01:27-04:01:28, 04:01:30-04:01:33, 04:01:35-04:01:42, 04:01:44-04:01:51, 04:01:53-04:01:67, 04:01:69-04:01:70 and 04:98:01-04:98:02 alleles, the C*04:01:05, 04:01:07, 04:01:09, 04:01:18, 04:01:22, 04:01:26, 04:01:68, 04:64:01-04:64:02 and 04:180:01 alleles, the C*04:15:01 and 04:15:03 alleles, the C*04:42:01-04:42:02 alleles,  the C*05:01:01:01-05:01:11, 05:01:13-05:01:15, 05:01:19-05:01:21, 05:01:23, 05:01:25-05:01:26, 05:01:29-05:01:31 05:22:01-05:22:02 and 05:44:01-05:44:02 alleles, the C*05:01:16, 05:01:18 and 05:18:02-05:18:03 alleles, the C*05:04:01-05:04:02 alleles, the C*05:09:01 and 05:09:03 alleles, the C*05:64:01-05:64:02 alleles, the C*06:02:01:01-06:02:01:03, 06:02:03-06:02:07, 06:02:09-06:02:11, 06:02:13-06:02:15, 06:02:17-06:02:22, 06:02:24-06:02:37, 06:02:39-06:02:42, 06:02:44, 06:42:02, 06:43:02 and 06:106:01-06:106:02 alleles, the C*06:02:16, 06:02:43 and 06:127:01-06:127:02 alleles, the C*06:04:01-06:04:02 alleles, the C*06:53:01-06:53:02 alleles, the C*06:70:01-06:70:02 alleles, the C*06:132:01-06:132:02 alleles, the C*07:01:01:01-07:01:04, 07:01:06-07:01:12, 07:01:14-07:01:24, 07:01:26-07:01:35, 07:01:37-07:01:45, 07:01:47-07:01:52 and 07:78:01-07:78:02 alleles, the C*07:01:05 and 07:01:36 alleles, the C*07:02:01:01-07:02:08, 07:02:10-07:02:31, 07:02:33-07:02:41, 07:02:43-07:02:62, 07:02:64, 07:02:66-07:02:71 and 07:38:01-07:38:02 alleles, the 07:17:01-07:17:02 alleles, the C*07:04:01-07:04:10 and 07:199:01-07:199:02 alleles, the C*07:26:01-07:26:02 alleles, the C*07:27:01-07:27:02 alleles, the C*07:29:01-07:29:02 alleles, the C*07:31:01-07:31:02 alleles, the C*07:56:01-07:56:02 alleles, the C*07:76:01-07:76:02 alleles, the C*07:96:01-07:96:02 alleles, the C*07:108:01-07:108:02 alleles, the C*07:141:01-07:141:02 alleles, the C*07:172:01-07:172:02 alleles, the C*08:01:01-08:01:06, 08:01:08-08:01:13, 08:01:15-08:01:18 and 08:16:02 alleles, the C*08:02:01:01-08:02:02, 08:02:04-08:02:06, 08:02:08-08:02:10 and 08:02:12 alleles, the C*08:03:01-08:03:03 alleles, the C*08:04:01-08:04:03 alleles, the C*12:02:01-12:02:05, 12:02:07, 12:02:09-12:02:10 and 12:02:13 alleles, the C*12:02:06, 12:02:08 and 12:02:12 alleles, the C*12:03:01:01-12:03:01:03, 12:03:03, 12:03:05-12:03:07, 12:03:11-12:03:12, 12:03:14-12:03:15, 12:03:17-12:03:18, 12:03:21-12:03:22, 12:03:25-12:03:29, 12:03:31, 12:03:33 and 12:03:35-12:03:37 alleles, the C*12:03:02, 12:03:08, 12:03:16 and 12:03:30 alleles, the C*12:03:19 and 12:03:32 alleles, the C*12:03:34:01-12:03:34:02 alleles,  the C*12:10:01-12:10:02 alleles, the C*12:14:01 and 12:18:02, the C*12:14:02 and 12:18:01 alleles, the C*14:02:01, 14:02:04-14:02:07, 14:02:09 and 14:02:12-14:02:19 alleles, the C*15:02:01:01-15:02:08, 15:02:11-15:02:16, 15:02:19-15:02:20 and 15:02:23-15:02:24 alleles, the C*15:02:10 and 15:02:17 alleles, the C*15:02:18 and 15:02:22, the C*15:04:01-15:04:02 alleles, the C*15:05:01-15:05:08 and 15:05:10 alleles, the C*15:06:01-15:06:03 alleles, the C*15:10:02-15:10:03 alleles, the C*16:01:01:01-16:01:01:02, 16:01:03-16:01:04, 16:01:07-16:01:15 and 16:01:18-16:01:19 alleles, the C*16:01:05, 16:07:02 alleles, the C*16:02:01-16:02:10 and 16:02:12 alleles, C*16:04:01 and 16:04:03 alleles,  and the C*17:01:01:01-17:01:03, 17:01:05-17:01:07 and 17:01:09-17:01:11 alleles.

‘?’, nucleotide sequence information not available for the primer matching sequence.

‘w’, might be weakly amplified.

Changes in revision R01 compared to R00:

- The C*14:17, 16:68 and C*16:88 alleles are amplified by primer mix 84. This has been corrected in the Specificity and Interpretation Tables.

Changes in revision R02 compared to R01:

- Primer mix 18 does not amplify the C*03:28 and C*03:49 alleles. This has been corrected in the Specificity and Interpretation Tables.