PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® HLA-C*03

Prod. No:	101.611-12/12u

Lot No:    	57N

Expiry Date:    	2014-September-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 48 primer mixes used for HLA-C*03 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified HLA-C*03 alleles3 | Other amplified HLA Class I alleles4 |
| --- | --- | --- | --- | --- |
| 1 | 280 bp | 800 bp | *03:02:01-03:02:08, 03:04:01:01-03:10, 03:14-03:17, 03:19, 03:23-03:29, 03:32-03:38:02, 03:40:01-03:42, 03:44-03:48, 03:51, 03:54, 03:57, 03:60, 03:63-03:65, 03:70-03:74, 03:77-03:78, 03:80, 03:82, 03:84, 03:87, 03:89-03:95, 03:98, 03:100-03:101, 03:104-03:111, 03:113-03:115, 03:117-03:118, 03:121N, 03:123, 03:125, 03:128-03:131, 03:134-03:136, 03:138-03:139 | *02:02:01-02:02:03, 02:02:05-02:03, 02:04w, 02:05-02:13, 02:14w, 02:15-02:25Q, 02:26:02-02:32, 02:34-02:40, 02:42-02:55, 04:03, 04:06, 04:16, 04:80, 04:107, 05:58:01, 06:03, 07:02:10w, 07:96, 07:127w, 15:02:01-15:09, 15:10:02-15:11, 15:13, 15:15-15:22, 15:24-15:35, 15:37-15:56, 16:34 |
| 2 | 210 bp | 1070 bp | *03:02:01-03:02:08, 03:14-03:15, 03:33, 03:36, 03:40:01-03:40:02, 03:42-03:43:02, 03:60, 03:71, 03:84, 03:89, 03:95, 03:108, 03:110, 03:119, 03:121N, 03:132, 03:139 | *07:133, B*15:96 |
| 3 | 280 bp | 800 bp | *03:03:01-03:03:18, 03:11:01-03:13, 03:20N-03:22Q, 03:30-03:31, 03:43:01-03:43:02, 03:49-03:50, 03:52-03:53, 03:55-03:56, 03:58-03:59, 03:61-03:62, 03:66, 03:67w, 03:68-03:69, 03:75-03:76, 03:79, 03:81, 03:83, 03:85-03:86, 03:88, 03:96-03:97, 03:102-03:103, 03:112, 03:116, 03:119-03:120, 03:122, 03:124, 03:126-03:127, 03:132-03:133 | *15:12 |
| 4 | 275 bp | 1070 bp | *03:03:01-03:04:25, 03:06-03:12, 03:14, 03:18-03:24, 03:26, 03:28-03:32, 03:34, 03:37-03:40:02, 03:42-03:57, 03:59, 03:61-03:70, 03:72-03:83, 03:85, 03:87-03:88, 03:90-03:93, 03:96, 03:98, 03:100-03:107, 03:109, 03:111-03:112, 03:114-03:120, 03:122-03:131, 03:133-03:134, 03:136-03:138 |  |
| 5 | 275 bp | 1070 bp | *03:05, 03:13, 03:17, 03:25, 03:27, 03:35, 03:71 | *07:133, B*15:96 |
| 65,9 | 90 bp, 215 bp | 800 bp | *03:06, 03:21 | B*15:96 |
| 75,6,10 | 110 bp, 185 bp | 1070 bp | *03:03:01-03:03:18, 03:08, 03:11:01-03:13, 03:18, 03:20N-03:22Q, 03:29-03:31, 03:39, 03:43:01-03:43:02, 03:49-03:50, 03:52-03:53, 03:55-03:56, 03:58-03:59, 03:61-03:62, 03:66-03:69, 03:75-03:76, 03:79, 03:81, 03:83, 03:85-03:86, 03:88, 03:96-03:97, 03:102-03:103, 03:112, 03:116, 03:119-03:120, 03:122, 03:124, 03:126-03:127, 03:132-03:133 | *01:51, 06:44, 06:53, 07:01:01-07:01:24, 07:06-07:07, 07:09, 07:16, 07:18-07:20, 07:22, 07:24, 07:26, 07:28, 07:30, 07:35-07:36, 07:40, 07:44, 07:52-07:53, 07:55N-07:56:01, 07:57-07:60, 07:65, 07:70-07:71, 07:73, 07:77-07:79, 07:81-07:83, 07:86, 07:89, 07:91-07:96, 07:98N, 07:103-07:104N, 07:106, 07:108-07:113, 07:115-07:116, 07:118-07:120, 07:122, 07:124, 07:128-07:129, 07:131-07:132, 07:134, 07:140-07:141, 07:148, 07:150Q-07:151, 07:153, 07:162, 07:164N-07:166, 07:170, 07:173, 07:176, 07:179-07:180, 07:182, 07:188-07:191N, 07:196-07:197, 07:200-07:201, 07:203-07:207, 07:210, 07:212, 07:214-07:215, 07:219, 18:05, B*07:93, B*08:39, B*35:101:02, B*38:27, B*51:76 |
| 8 | 255 bp | 1070 bp | *03:09, 03:22Q, 03:42, 03:71, 03:89, 03:116, 03:119, 03:139 | *07:133 |
| 9 | 240 bp | 1070 bp | *03:07, 03:15, 03:45, 03:130 | *02:02:01-02:02:03, 02:02:05-02:02:11, 02:02:13-02:03, 02:04w, 02:05-02:11, 02:13, 02:14w, 02:15-02:25Q, 02:26:02-02:26:03, 02:28-02:40, 02:42-02:55, 04:03, 04:06, 04:16, 04:80, 04:107, 05:58:01, 06:03, 15:02:01-15:06:03, 15:08-15:09, 15:10:02-15:13, 15:15-15:20, 15:22, 15:24, 15:26-15:35, 15:37-15:42, 15:44-15:56 |
| 10 | 250 bp | 800 bp | *03:07, 03:10, 03:15, 03:29, 03:45 | *02:02:01w, 02:02:02-02:02:03, 02:02:05-02:03, 02:04w, 02:05-02:13, 02:14w, 02:15-02:25Q, 02:26:02-02:26:03, 02:28-02:40, 02:42-02:55, 04:03, 04:06, 04:16, 04:80, 04:107, 05:58:01, 06:03, 15:02:01-15:06:03, 15:08-15:09, 15:10:02-15:13, 15:15-15:22, 15:24, 15:26-15:35, 15:37-15:42, 15:44-15:56 |
| 1111 | 145 bp, 415 bp | 1070 bp | *03:08, 03:29, 03:31, 03:87 | *05:36, 06:44, 12:08 |
| 1212 | 260 bp, 305 bp | 1070 bp | *03:03:12, 03:03:15, 03:04:19, 03:11:01, 03:18, 03:39, 03:113 | *01:51, 05:55, 06:53, 07:10, 07:43, 07:56:01, 07:60, 07:79, 07:109, 07:184, 07:196, 12:58, 15:12, B*35:08:04, B*37:04:02 |
| 135,13 | 80 bp, 140 bp | 1070 bp | *03:02:01-03:02:08, 03:15-03:17, 03:33, 03:36, 03:40:01, 03:42-03:43:01, 03:60, 03:71, 03:84, 03:89, 03:95, 03:108, 03:110, 03:121N, 03:132, 03:139 | *04:37, 04:54, 06:02:08, 12:03:20, 14:02:01, 14:02:03-14:16, 14:18-14:34, 16:01:06, B*15:78:03 |
| 14 | 210 bp | 1070 bp | *03:12, 03:19, 03:102 |  |
| 157 | 135 bp | 1070 bp | *03:02:01-03:04:24, 03:05-03:17, 03:19-03:38:02, 03:40:01-03:66, 03:67w, 03:68-03:98, 03:100-03:117, 03:119-03:136, 03:138-03:139 | *02:02:13, 06:03w, 07:96, 15:43 |
| 16 | 150 bp | 1070 bp | *03:07, 03:10, 03:15, 03:29, 03:45 | *01:14, 02:02:13, 04:01:23, 05:01:12, 06:02:01:01-06:02:01:02, 06:02:03-06:04, 06:06-06:10, 06:12-06:69, 07:07, 07:09, 07:76, 12:04:01, 14:12, 17:01:01:01-17:01:03, 17:01:05-17:11, 18:01-18:05 |
| 175 | 100 bp | 1070 bp | *03:20N, 03:101 |  |
| 1814 | 165 bp, 195 bp, 225 bp | 1070 bp | *03:23, 03:83, 03:102 | B*15:96 |
| 195,15 | 105 bp, 260 bp | 800 bp | *03:24, 03:33 |  |
| 2016 | 165 bp, 255 bp | 1070 bp | *03:25, 03:47 |  |
| 2117 | 210 bp, 240 bp | 1070 bp | *03:26, 03:57, 03:88 |  |
| 22 | 195 bp | 800 bp | *03:15, 03:27, 03:38:01-03:38:02, 03:69, 03:130, 03:136 | *02:12, 02:49, 02:55, 04:03, 04:06, 04:16, 04:80, 04:107, 07:20, 07:96, 07:127, 15:03, 15:16, 15:25 |
| 235,18 | 90 bp, 260 bp | 1070 bp | *03:22Q, 03:28, 03:49, 03:90, 03:103, 03:117 |  |
| 2419 | 160 bp, 200 bp | 1070 bp | *03:30, 03:65, 03:79, 03:120-03:121N |  |
| 2520 | 305 bp, 355 bp | 1070 bp | *03:32, 03:38:01-03:38:02, 03:45, 03:69, 03:130, 03:136 | *04:80, 04:100, 06:14, 07:10, 07:43, 07:196, 12:44, 15:03, 15:16, 15:25, B*35:178 |
| 265,21 | 70 bp, 105 bp, 265 bp | 1070 bp | *03:13, 03:34-03:36, 03:77 | *07:133, B*15:96 |
| 275,22 | 80 bp, 275 bp | 1070 bp | *03:11:01-03:11:02, 03:15, 03:32, 03:37, 03:45, 03:60, 03:136 | *02:42, 02:49, 04:03, 04:06, 04:16, 04:80, 06:03, 07:96, 15:25 |
| 28 | 170 bp | 1070 bp | *03:16, 03:41, 03:113 | *04:01:01:01-04:01:04, 04:01:06, 04:01:08, 04:01:10-04:01:17, 04:01:19-04:01:21, 04:01:23-04:01:25, 04:01:27-04:01:36, 04:03-04:04:01, 04:05-04:20, 04:23-04:49, 04:51-04:63, 04:65-04:76, 04:78-04:94:01, 04:95N-04:98, 04:100-04:107, 06:02:08, 12:03:20, 14:02:01, 14:02:03-14:16, 14:18-14:31, 14:33-14:34, 16:01:06, 18:01-18:05, B*15:01:17, B*15:78:03 |
| 2923 | 140 bp, 170 bp, 285 bp | 1070 bp | *03:39, 03:51, 03:67, 03:96, 03:98 | *04:64, 06:53, 07:56:01, 07:60, 07:79, 07:109, B*08:39 |
| 3024 | 150 bp, 170 bp | 1070 bp | *03:46, 03:61, 03:96 | *04:64 |
| 315,25 | 100 bp, 210 bp | 800 bp | *03:44, 03:59, 03:123 |  |
| 325,8,26 | 90 bp, 180 bp | 1070 bp | *03:48, 03:58, 03:84 | *01:03, 01:24, 04:37 |
| 335,27 | 110 bp, 365 bp | 1070 bp | *03:43:01-03:43:02, 03:62, 03:119, 03:132 | *07:60, 07:79, 12:58, 14:10, B*35:101:02 |
| 3428 | 130 bp, 240 bp | 1070 bp | *03:53-03:54, 03:98 | *12:17 |
| 355,29 | 70 bp, 120 bp, 210 bp | 1070 bp | *03:56, 03:64, 03:79, 03:82, 03:85 | *02:08, 15:10:02 |
| 3630 | 165 bp, 245 bp | 1070 bp | *03:50, 03:72, 03:122 |  |
| 37 | 140 bp | 1070 bp | *03:16, 03:21, 03:55, 03:80, 03:92, 03:113, 03:117 | *04:01:01:01-04:01:04, 04:01:08, 04:01:10-04:01:21, 04:01:23-04:01:25, 04:01:27-04:01:36, 04:03-04:04:01, 04:05-04:07, 04:09N-04:20, 04:23-04:31, 04:33, 04:35-04:49, 04:51-04:58, 04:60-04:76, 04:78-04:84, 04:86-04:94:01, 04:95N-04:98, 04:100-04:107, 06:02:08, 08:02:07, 08:33:02, 12:03:20, 14:02:01, 14:02:03-14:16, 14:18-14:19, 14:21N-14:24, 14:26-14:31, 14:33-14:34, 16:01:06, 18:01-18:02, 18:04-18:05 |
| 385,31 | 100 bp, 265 bp | 1070 bp | *03:52, 03:73, 03:95 | *04:43 |
| 395,32 | 110 bp, 170 bp | 1070 bp | *03:36, 03:63, 03:77 | *04:34 |
| 405,33 | 70 bp, 260 bp | 1070 bp | *03:66, 03:80 |  |
| 415,34 | 120 bp, 195 bp | 1070 bp | *03:68, 03:100 |  |
| 425,35 | 120 bp, 205 bp | 1070 bp | *03:70, 03:78, 03:127 | B*55:30 |
| 437,36 | 220 bp, 275 bp | 800 bp | *03:74, 03:81, 03:97 | B*15:96 |
| 445,37 | 70 bp, 160 bp | 1070 bp | *03:75, 03:91, 03:120 |  |
| 4538 | 140 bp, 295 bp | 1070 bp | *03:76, 03:93 |  |
| 46 | 210 bp | 1070 bp | *03:58, 03:86, 03:94, 03:99 | *01:02:01-01:03, 01:06-01:08, 01:10-01:11, 01:13-01:20, 01:23-01:33, 01:38-01:48, 01:51-01:54, 01:56N-01:61, 04:37 |
| 475,39 | 70 bp, 195 bp, 240 bp | 1070 bp | *03:14, 03:21, 03:55, 03:57, 03:92, 03:109, 03:117 | *07:133 |
| 48 | 555 bp | 1070 bp | *03:58, 03:86, 03:94, 03:99 | *15:37, 16:18 |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-C*03 SSP typings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective lengths of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most wells, or a band of 800 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-C*03 subtyping.

In addition, wells number 3, 6, 10, 19, 22, 31 and 43 contain the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3For several HLA-C*03 alleles 1st and 4th exon or intron nucleotide sequences are not available. In these instances it is not known whether some of the primers of the SSP sets are completely matched with the target sequences or not. We assume that unknown sequences in the 1st and 4th exons or the introns are conserved within allelic groups.

4Due to the sharing of sequence motifs between HLA-C alleles the primer pairs in wells 1 to 3, 5, 7 to 13, 15, 16, 22, 25  to 30, 32 to 35, 37 to 39 and 46 to 48 will amplify non-HLA-C*03 alleles. In addition, the B*07:93, 38:27 and 51:76 alleles are amplified by primer mix 7, the B*08:39 allele is amplified by primer mix 7 and 29, the B*15:01:17 allele is amplified by primer mix 28, the B*15:78:03 allele is amplified by primer mixes 13 and 28, the B*15:96 allele is amplified by primer mixes 2, 5, 6, 18, 26 and 43, the B*35:08:04 and 37:04:02 alleles are amplified by primer mix 12, the B*35:101:02 allele is amplified by primer mix 7 and 33, the B*35:178 allele is amplified by primer mix 25 and the B*55:30 allele is amplified by primer mix 42.

5Short specific PCR fragments are less intense and not as sharp as longer specific bands.

6Primer mix 7 may yield less specific PCR product than the other HLA-C*03 primer mixes.

7Primer mixes 15 and 43 may give rise to primer oligomer formation.

8Primer mix 32 may give rise to nonspecific amplifications.

9Primer mix 6: Specific PCR fragment of 90 bp in the C*03:21 and the B*15:96 alleles. Specific PCR fragment of 215 bp in the C*03:06 allele.

10Primer mix 7: Specific PCR fragment of 110 bp in the C*03:08 and 03:29 and the C*06:44, 07:01:01-07:01:24, 07:06-07:07, 07:09, 07:16, 07:18-07:20, 07:22, 07:24, 07:26, 07:28, 07:30, 07:35-07:36, 07:40, 07:44, 07:52-07:53, 07:55N, 07:57-07:59, 07:65, 07:70-07:71, 07:73, 07:77-07:78, 07:81-07:83, 07:86, 07:89, 07:91-07:96, 07:98N, 07:103-07:104N, 07:106, 07:108, 07:110-07:113, 07:115-07:116, 07:118-07:120, 07:122, 07:124, 07:128-07:129, 07:131-07:132, 07:134, 07:140-07:141, 07:148, 07:150Q-07:151, 07:153, 07:162, 07:164N-07:166, 07:170, 07:173, 07:176, 07:179-07:180, 07:182, 07:188-07:191N, 07:196-07:197, 07:200-07:201, 07:203-07:207, 07:210, 07:212, 07:214-07:215, 07:219 and 18:05 alleles. Specific PCR fragment of 185 bp in the C*03:03:01-03:03:18, 03:11:01-03:13, 03:18, 03:20N-03:22Q, 03:30, 03:39, 03:43:01-03:43:02, 03:49-03:50, 03:52-03:53, 03:55-03:56, 03:58-03:59, 03:61-03:62, 03:66-03:69, 03:75-03:76, 03:79, 03:81, 03:83, 03:85-03:86, 03:88, 03:96-03:97, 03:102-03:103, 03:112, 03:116, 03:119-03:120, 03:122, 03:124, 03:126-03:127 and 03:132-03:133 and the C*01:51, 06:53, 07:56:01 and 07:79 and in the  B*07:93, B*08:39, B*35:101:02, B*38:27, B*51:76 alleles. Specific PCR fragments of 110 and 185 bp in the C*03:31 and in the C*07:60 and 07:109 alleles.

11Primer mix 11: Specific PCR fragment of 145 bp in the C*03:87 allele. Specific PCR fragment of 415 bp in the C*03:08, 03:29 and 03:31 and in the C*05:36, 06:44 and 12:08 alleles.

12Primer mix 12: Specific PCR fragment of 260 bp in the C*03:03:12, 03:03:15, 03:11:01, 03:18 and 03:39 and the C*01:51, 05:55, 06:53, 07:56:01, 07:60, 07:79, 07:109, 12:58 and 15:12 alleles. Specific PCR fragment of 305 bp in the C*03:04:19 and 03:113 and the C*07:10, 07:43, 07:184 and 07:196 and in the B*35:08:04 and B*37:04:02 alleles.

13Primer mix 13: Specific PCR fragment of 80 bp in the C*03:02:01-03:02:08, 03:15-03:16, 03:33, 03:36, 03:40:01, 03:42-03:43:01, 03:60, 03:84, 03:89, 03:95, 03:108, 03:110, 03:121N, 03:132 and 03:139 and the C*04:54, 14:08-14:09 and 14:28 and in the B*15:78:03 alleles. Specific PCR fragment of 140 bp in the C*03:17 and the C*04:37 and 14:26 alleles.  Specific PCR fragments of 80 and 140 bp in the C*03:71 and in the C*06:02:08, 12:03:20, 14:02:01, 14:02:03-14:07N, 14:10-14:16, 14:18-14:25, 14:27, 14:29-14:34 and 16:01:06 alleles.

14Primer mix 18: Specific PCR fragment of 165 bp in the C*03:102 and in the B*15:96 alleles. Specific PCR fragment of 195 bp in the C*03:83 allele. Specific PCR fragment of 225 bp in the C*03:23 allele.

15Primer mix 19: Specific PCR fragment of 105 bp in the C*03:24 allele. Specific PCR fragment of 260 bp in the C*03:33 allele.

16Primer mix 20: Specific PCR fragment of 165 bp in the C*03:25 allele. Specific PCR fragment of 255 bp in the C*03:47 allele.

17Primer mix 21: Specific PCR fragment of 210 bp in the C*03:26 allele. Specific PCR fragment of 240 bp in the C03:57 and 03:88 alleles.

18Primer mix 23: Specific PCR fragment of 90 bp in the C*03:28, 03:49 and 03:117 alleles. Specific PCR fragment of 260 bp in the C*03:22Q, 03:90 and 03:103 alleles.

19Primer mix 24: Specific PCR fragment of 160 bp in the C*03:120 allele. Specific PCR fragment of 200 bp in the C*03:30, 03:65, 03:79 and 03:121N alleles.

20Primer mix 25: Specific PCR fragment of 305 bp in the C*03:32 and 03:45 and in the B*35:178 alleles. Specific PCR fragment of 355 bp in the C*03:38:01-03:38:02, 03:69 and 03:130 and the C*12:44, 15:03 and 15:16 alleles. Specific PCR fragments of 305 and 355 bp in the C*03:136 and the C*04:80, 04:100, 06:14, 07:10, 07:43, 07:196 and 15:25 alleles.

21Primer mix 26: Specific PCR fragment of 70 bp in the C*03:34 allele. Specific PCR fragment of 105 bp in the C*03:36 and 03:77 alleles. Specific PCR fragment of 265 bp in the C*03:13 and 03:35 and in the C*07:133 and the B*15:96 alleles.

22Primer mix 27: Specific PCR fragment of 80 bp in the C*03:11:01-03:11:02 and 03:37 and the C*02:42 alleles. Specific PCR fragment of 275 bp in the C*03:15, 03:32, 03:45, 03:60 and 03:136 and the C*02:49, 04:03, 04:06, 04:16, 04:80, 07:96 and 15:25 alleles. Specific PCR fragment of 80 and 275 bp in the C*06:03 allele.

23Primer mix 29: Specific PCR fragment of 140 bp in the C*03:98 allele. Specific PCR fragment of 170 bp in the C*03:51 and 03:96 and in the C*04:64 alleles. Specific PCR fragment of 285 bp in the C*03:39 and 03:67 and the C*06:53, 07:56:01, 07:60, 07:79 and 07:109 and in the B*08:39 alleles.

24Primer mix 30: Specific PCR fragment of 150 bp in the C*03:61 allele. Specific PCR fragment of 170 bp in the C*03:46 and 03:96 and the C*04:64 alleles.

25Primer mix 31: Specific PCR fragment of 100 bp in the C*03:59 and 03:123 alleles. Specific PCR fragment of 210 bp in the C*03:44 allele.

26Primer mix 32: Specific PCR fragment of 90 bp in the C*03:58 and in the C*01:03, 01:24 and 04:37 alleles. Specific PCR fragment of 180 bp in the C*03:48 and 03:84 alleles.

27Primer mix 33: Specific PCR fragment of 110 bp in the C*03:62 allele. Specific PCR fragment of 365 bp in the C*03:43:01-03:43:02, 03:119 and 03:132 and the C*07:60, 07:79, 12:58 and 14:10 and in the B*35:101:02 alleles.

28Primer mix 34: Specific PCR fragment of 130 bp in the C*03:54 and 03:98 alleles. Specific PCR fragment of 240 bp in the C* 03:53 and the C*12:17 alleles.

29Primer mix 35: Specific PCR fragment of 70 bp in the C*03:64 and the C*02:08 and 15:10:02 alleles. Specific PCR fragment of 120 bp in the C*03:56 allele. Specific PCR fragment of 210 bp in the C*03:79, 03:82 and 03:85 alleles.

30Primer mix 36: Specific PCR fragment of 165 bp in the C*03:122 allele. Specific PCR fragment of 245 bp in the C*03:50 and 03:72 alleles.

31Primer mix 38: Specific PCR fragment of 100 bp in the C*03:52 and 03:95 and in the C*04:43 alleles. Specific PCR fragment of 265 bp in the C*03:73 allele.

32Primer mix 39: Specific PCR fragment of 110 bp in the C*03:36 and 03:77 and in the C*04:34 alleles. Specific PCR fragment of 170 bp in the C*03:63 allele.

33Primer mix 40: Specific PCR fragment of 70 bp in the C*03:80 alleles. Specific PCR fragment of 260 bp in the C*03:66 allele.

34Primer mix 41: Specific PCR fragment of 120 bp in the C*03:68 alleles. Specific PCR fragment of 195 bp in the C*03:100 allele.

35Primer mix 42: Specific PCR fragment of 120 bp in the C*03:78 alleles. Specific PCR fragment of 205 bp in the C*03:70 and 03:127 and in the B*55:30 alleles.

36Primer mix 43: Specific PCR fragment of 220 bp in the C*03:74 and 03:81 and in the B*15:96 alleles. Specific PCR fragment of 275 bp in the C*03:97 allele.

37Primer mix 44: Specific PCR fragment of 70 bp in the C*03:91 allele. Specific PCR fragment of 160 bp in the C*03:75 and 03:120 alleles.

38Primer mix 45: Specific PCR fragment of 140 bp in the C*03:93 alleles. Specific PCR fragment of 295 bp in the C*03:76 allele.

39Primer mix 47: Specific PCR fragment of 70 bp in the C*03:14, 03:21, 03:55, 03:92 and 03:117 and the C*07:133 alleles. Specific PCR fragment of 195 bp in the C*03:109 allele. Specific PCR fragment of 240 bp in the C*03:57 allele.

‘w’, might be weakly amplified.

1The nucleotide sequence of the HLA-Cw*0301 allele has been shown to be identical to C*03:04:01:01.

2HLA-C*03 alleles in bold lettering are listed as confirmed alleles on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.7.0, January 2012.

3The C*03:28 and 03:90 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 23.

4The C*03:44 and 03:123 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 31.

5The C*03:49 and 03:103 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 23.

6The C*03:50 and 03:122 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 36.

7The C*03:56 and 03:85 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 35.

8The C*03:64 and 03:82 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 35.

9The C*03:70 and 03:78 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 42.

10The C*03:135 and C*15:43 give rise to identical amplification patterns with the HLA-C*03  subtyping kit. These two alleles can be distinguished by the HLA-C low resolution kit and/or the HLA-C*15 subtyping kits.

11Primer mix 6: Specific PCR fragment of 90 bp in the C*03:21 and the B*15:96 alleles. Specific PCR fragment of 215 bp in the C*03:06 allele.

Primer mix 7: Specific PCR fragment of 110 bp in the C*03:08 and 03:29 and the C*06:44, 07:01:01-07:01:24, 07:06-07:07, 07:09, 07:16, 07:18-07:20, 07:22, 07:24, 07:26, 07:28, 07:30, 07:35-07:36, 07:40, 07:44, 07:52-07:53, 07:55N, 07:57-07:59, 07:65, 07:70-07:71, 07:73, 07:77-07:78, 07:81-07:83, 07:86, 07:89, 07:91-07:96, 07:98N, 07:103-07:104N, 07:106, 07:108, 07:110-07:113, 07:115-07:116, 07:118-07:120, 07:122, 07:124, 07:128-07:129, 07:131-07:132, 07:134, 07:140-07:141, 07:148, 07:150Q-07:151, 07:153, 07:162, 07:164N-07:166, 07:170, 07:173, 07:176, 07:179-07:180, 07:182, 07:188-07:191N, 07:196-07:197, 07:200-07:201, 07:203-07:207, 07:210, 07:212, 07:214-07:215, 07:219 and 18:05 alleles. Specific PCR fragment of 185 bp in the C*03:03:01-03:03:18, 03:11:01-03:13, 03:18, 03:20N-03:22Q, 03:30, 03:39, 03:43:01-03:43:02, 03:49-03:50, 03:52-03:53, 03:55-03:56, 03:58-03:59, 03:61-03:62, 03:66-03:69, 03:75-03:76, 03:79, 03:81, 03:83, 03:85-03:86, 03:88, 03:96-03:97, 03:102-03:103, 03:112, 03:116, 03:119-03:120, 03:122, 03:124, 03:126-03:127 and 03:132-03:133 and the C*01:51, 06:53, 07:56:01 and 07:79 and in the  B*07:93, B*08:39, B*35:101:02, B*38:27, B*51:76 alleles. Specific PCR fragments of 110 and 185 bp in the C*03:31 and in the C*07:60 and 07:109 alleles.

Primer mix 11: Specific PCR fragment of 145 bp in the C*03:87 allele. Specific PCR fragment of 415 bp in the C*03:08, 03:29 and 03:31 and in the C*05:36, 06:44 and 12:08 alleles.

Primer mix 12: Specific PCR fragment of 260 bp in the C*03:03:12, 03:03:15, 03:11:01, 03:18 and 03:39 and the C*01:51, 05:55, 06:53, 07:56:01, 07:60, 07:79, 07:109, 12:58 and 15:12 alleles. Specific PCR fragment of 305 bp in the C*03:04:19 and 03:113 and the C*07:10, 07:43, 07:184 and 07:196 and in the B*35:08:04 and B*37:04:02 alleles.

Primer mix 13: Specific PCR fragment of 80 bp in the C*03:02:01-03:02:08, 03:15-03:16, 03:33, 03:36, 03:40:01, 03:42-03:43:01, 03:60, 03:84, 03:89, 03:95, 03:108, 03:110, 03:121N, 03:132 and 03:139 and the C*04:54, 14:08-14:09 and 14:28 and in the B*15:78:03 alleles. Specific PCR fragment of 140 bp in the C*03:17 and the C*04:37 and 14:26 alleles.  Specific PCR fragments of 80 and 140 bp in the C*03:71 and in the C*06:02:08, 12:03:20, 14:02:01, 14:02:03-14:07N, 14:10-14:16, 14:18-14:25, 14:27, 14:29-14:34 and 16:01:06 alleles.

Primer mix 18: Specific PCR fragment of 165 bp in the C*03:102 and in the B*15:96 alleles. Specific PCR fragment of 195 bp in the C*03:83 allele. Specific PCR fragment of 225 bp in the C*03:23 allele.

Primer mix 19: Specific PCR fragment of 105 bp in the C*03:24 allele. Specific PCR fragment of 260 bp in the C*03:33 allele.

Primer mix 20: Specific PCR fragment of 165 bp in the C*03:25 allele. Specific PCR fragment of 255 bp in the C*03:47 allele.

Primer mix 21: Specific PCR fragment of 210 bp in the C*03:26 allele. Specific PCR fragment of 240 bp in the C03:57 and 03:88 alleles.

Primer mix 23: Specific PCR fragment of 90 bp in the C*03:28, 03:49 and 03:117 alleles. Specific PCR fragment of 260 bp in the C*03:22Q, 03:90 and 03:103 alleles.

Primer mix 24: Specific PCR fragment of 160 bp in the C*03:120 allele. Specific PCR fragment of 200 bp in the C*03:30, 03:65, 03:79 and 03:121N alleles.

Primer mix 25: Specific PCR fragment of 305 bp in the C*03:32 and 03:45 and in the B*35:178 alleles. Specific PCR fragment of 355 bp in the C*03:38:01-03:38:02, 03:69 and 03:130 and the C*12:44, 15:03 and 15:16 alleles. Specific PCR fragments of 305 and 355 bp in the C*03:136 and the C*04:80, 04:100, 06:14, 07:10, 07:43, 07:196 and 15:25 alleles.

Primer mix 26: Specific PCR fragment of 70 bp in the C*03:34 allele. Specific PCR fragment of 105 bp in the C*03:36 and 03:77 alleles. Specific PCR fragment of 265 bp in the C*03:13 and 03:35 and in the C*07:133 and the B*15:96 alleles.

Primer mix 27: Specific PCR fragment of 80 bp in the C*03:11:01-03:11:02 and 03:37 and the C*02:42 alleles. Specific PCR fragment of 275 bp in the C*03:15, 03:32, 03:45, 03:60 and 03:136 and the C*02:49, 04:03, 04:06, 04:16, 04:80, 07:96 and 15:25 alleles. Specific PCR fragment of 80 and 275 bp in the C*06:03 allele.

Primer mix 29: Specific PCR fragment of 140 bp in the C*03:98 allele. Specific PCR fragment of 170 bp in the C*03:51 and 03:96 and in the C*04:64 alleles. Specific PCR fragment of 285 bp in the C*03:39 and 03:67 and the C*06:53, 07:56:01, 07:60, 07:79 and 07:109 and in the B*08:39 alleles.

Primer mix 30: Specific PCR fragment of 150 bp in the C*03:61 allele. Specific PCR fragment of 170 bp in the C*03:46 and 03:96 and the C*04:64 alleles.

Primer mix 31: Specific PCR fragment of 100 bp in the C*03:59 and 03:123 alleles. Specific PCR fragment of 210 bp in the C*03:44 allele.

Primer mix 32: Specific PCR fragment of 90 bp in the C*03:58 and in the C*01:03, 01:24 and 04:37 alleles. Specific PCR fragment of 180 bp in the C*03:48 and 03:84 alleles.

Primer mix 33: Specific PCR fragment of 110 bp in the C*03:62 allele. Specific PCR fragment of 365 bp in the C*03:43:01-03:43:02, 03:119 and 03:132 and the C*07:60, 07:79, 12:58 and 14:10 and in the B*35:101:02 alleles.

Primer mix 34: Specific PCR fragment of 130 bp in the C*03:54 and 03:98 alleles. Specific PCR fragment of 240 bp in the C* 03:53 and the C*12:17 alleles.

Primer mix 35: Specific PCR fragment of 70 bp in the C*03:64 and the C*02:08 and 15:10:02 alleles. Specific PCR fragment of 120 bp in the C*03:56 allele. Specific PCR fragment of 210 bp in the C*03:79, 03:82 and 03:85 alleles.

Primer mix 36: Specific PCR fragment of 165 bp in the C*03:122 allele. Specific PCR fragment of 245 bp in the C*03:50 and 03:72 alleles.

Primer mix 38: Specific PCR fragment of 100 bp in the C*03:52 and 03:95 and in the C*04:43 alleles. Specific PCR fragment of 265 bp in the C*03:73 allele.

Primer mix 39: Specific PCR fragment of 110 bp in the C*03:36 and 03:77 and in the C*04:34 alleles. Specific PCR fragment of 170 bp in the C*03:63 allele.

Primer mix 40: Specific PCR fragment of 70 bp in the C*03:80 alleles. Specific PCR fragment of 260 bp in the C*03:66 allele.

Primer mix 41: Specific PCR fragment of 120 bp in the C*03:68 alleles. Specific PCR fragment of 195 bp in the C*03:100 allele.

Primer mix 42: Specific PCR fragment of 120 bp in the C*03:78 alleles. Specific PCR fragment of 205 bp in the C*03:70 and 03:127 and in the B*55:30 alleles.

Primer mix 43: Specific PCR fragment of 220 bp in the C*03:74 and 03:81 and in the B*15:96 alleles. Specific PCR fragment of 275 bp in the C*03:97 allele.

Primer mix 44: Specific PCR fragment of 70 bp in the C*03:91 allele. Specific PCR fragment of 160 bp in the C*03:75 and 03:120 alleles.

Primer mix 45: Specific PCR fragment of 140 bp in the C*03:93 alleles. Specific PCR fragment of 295 bp in the C*03:76 allele.

Primer mix 47: Specific PCR fragment of 70 bp in the C*03:14, 03:21, 03:55, 03:92 and 03:117 and the C*07:133 alleles. Specific PCR fragment of 195 bp in the C*03:109 allele. Specific PCR fragment of 240 bp in the C*03:57 allele.

‘w’, might be weakly amplified.