PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® HLA-C*04

Prod. No:	101.612-12/12u

Lot No:    	12M

Expiry Date:    	2013-November-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 47 primer mixes used for HLA-C*04 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified HLA-C*04 alleles3 | Other amplified HLA Class I alleles4 |
| --- | --- | --- | --- | --- |
| 1 | 250 bp | 800 bp | *04:01:01:01-04:01:26, 04:03, 04:05, 04:07-04:12, 04:15:01-04:18, 04:19w, 04:20, 04:23-04:33, 04:35-04:57, 04:59Q-04:67, 04:69-04:84 | *01:02:01-01:21, 01:23-01:34, 01:36-01:51, 02:02:01-02:02:03, 02:02:05-02:02:17, 02:04-02:15, 02:17, 02:19-02:40, 02:42-02:47, 05:01:01:01-05:01:15, 05:03-05:10, 05:12-05:16, 05:18-05:26, 05:28-05:50, 05:52-05:53, 06:02:01:01-06:02:01:02, 06:02:03-06:03, 06:05-06:39, 06:41-06:58, 08:02:01-08:02:05, 08:05, 08:07, 08:12, 08:15:01-08:15:02, 08:17-08:19, 08:23, 08:25, 08:27-08:35, 08:37, 08:43, 08:45, 12:02:01-12:13, 12:15-12:17, 12:21-12:51, 14:02:01-14:05, 14:07N-14:14, 14:16-14:29, 15:08, 16:04:01, 16:29, 18:01-18:04 |
| 2 | 220 bp | 1070 bp | *04:01:01:01-04:01:26, 04:04:01-04:05, 04:07-04:09N, 04:12w, 04:13-04:15:03, 04:17-04:20, 04:23-04:35, 04:37-04:41, 04:43-04:54, 04:56-04:70, 04:72-04:79, 04:81-04:84 | *07:64 |
| 3 | 145 bp | 1070 bp | *04:01:01:01-04:01:26, 04:04:01-04:05, 04:08-04:15:03, 04:17-04:20, 04:23-04:26, 04:28-04:41, 04:43-04:79, 04:81-04:84 |  |
| 4 | 210 bp | 1070 bp | *04:03, 04:06, 04:80 | *02:12 |
| 5 | 250 bp | 1070 bp | *04:04:01-04:04:02, 04:06, 04:13, 04:34, 04:58 | *01:22, 01:35, 02:03, 02:16:01-02:16:02, 02:18, 05:11, 05:17, 05:27, 06:04, 08:01:01-08:01:04, 08:03:01-08:04, 08:06, 08:08-08:11, 08:13-08:14, 08:16, 08:20-08:22, 08:24, 08:26N, 08:36N, 08:38-08:42, 08:44, 08:46, 12:14:01-12:14:02, 12:18, 12:20, 14:06, 14:15, 15:02:01-15:07, 15:09-15:13, 15:15-15:24, 15:26-15:48, 17:01:01:01-17:08, B*58:02 |
| 65 | 95 bp | 1070 bp | *04:05 |  |
| 76 | 145 bp | 1070 bp | *04:07, 04:27, 04:32, 04:77 | *03:34, 07:64, 14:25, 15:36, 18:03 |
| 8 | 270 bp | 1070 bp | *04:08, 04:34 | *01:35, 05:27, 05:39, 08:41, 14:20, 15:15, 17:07, B*58:02 |
| 95,8 | 110 bp, 180 bp, 220 bp | 1070 bp | *04:09N, 04:30, 04:35 |  |
| 109 | 195 bp, 220 bp | 1070 bp | *04:10-04:11, 04:31, 04:36, 04:55 | *15:36 |
| 11 | 180 bp | 1070 bp | *04:01:01:01-04:01:09, 04:01:11-04:01:22, 04:01:24-04:01:26, 04:03-04:10, 04:12-04:20, 04:23-04:26, 04:28-04:32, 04:34-04:51, 04:53-04:54, 04:56-04:84 | *05:25, 05:42, 06:05, 07:02:09, 08:28, 12:28, 15:25, 16:26 |
| 125,10 | 125 bp, 165 bp | 1070 bp | *04:11, 04:29, 04:33, 04:36, 04:55 | *07:125 |
| 1311 | 225 bp, 270 bp | 1070 bp | *04:12, 04:52, 04:55 |  |
| 1412 | 155 bp, 190 bp | 1070 bp | *04:13, 04:16, 04:18, 04:58 | *07:08, 07:108 |
| 15 | 180 bp | 1070 bp | *04:14, 04:28, 04:68 | A*24:96, 24:146 |
| 165 | 130 bp | 1070 bp | *04:15:01, 04:15:03, 04:37 | *03:05, 03:25, 03:27, 14:09 |
| 17 | 320 bp | 1070 bp | *04:17, 04:80 | *01:50 |
| 185,13 | 125 bp, 225 bp | 1070 bp | *04:19, 04:70 |  |
| 195,14 | 120 bp, 155 bp | 1070 bp | *04:20, 04:35, 04:37 |  |
| 2015 | 165 bp, 250 bp, 545 bp | 1070 bp | *04:15:02, 04:17, 04:44, 04:47 | *15:36 |
| 215,6,16 | 85 bp, 145 bp, 240 bp | 1070 bp | *04:23, 04:38-04:39 |  |
| 225 | 130 bp | 1070 bp | *04:24 |  |
| 235,17 | 85 bp, 145 bp, 215 bp | 1070 bp | *04:25, 04:40-04:41 |  |
| 246 | 170 bp | 1070 bp | *04:26, 04:30, 04:42 | *07:125 |
| 25 | 190 bp | 1070 bp | *04:43 |  |
| 267 | 215 bp | 800 bp | *04:45 |  |
| 275,18 | 125 bp, 285 bp | 800 bp | *04:46, 04:50 |  |
| 285,19 | 120 bp, 215 bp | 1070 bp | *04:48, 04:75 |  |
| 29 | 185 bp | 1070 bp | *04:49 |  |
| 305 | 75 bp | 1070 bp | *04:53 | *05:49, B*07:90 |
| 31 | 180 bp | 1070 bp | *04:51 |  |
| 325,20 | 105 bp, 235 bp, 275 bp | 1070 bp | *04:59Q, 04:77-04:78 | B*40:100 |
| 335,21 | 80 bp, 180 bp | 1070 bp | *04:58, 04:65, 04:72 | *07:08, 07:108, B*40:100 |
| 34 | 200 bp | 1070 bp | *04:13, 04:58, 04:61, 04:68 | *07:08, 07:108 |
| 355,22 | 85 bp, 175 bp | 1070 bp | *04:62, 04:76 |  |
| 365 | 115 bp | 1070 bp | *04:57, 04:63 |  |
| 375,23 | 110 bp, 135 bp | 1070 bp | *04:63, 04:73-04:74 |  |
| 385,24 | 105 bp, 135 bp | 1070 bp | *04:74, 04:83 |  |
| 3925 | 145 bp, 330 bp | 1070 bp | *04:71, 04:79 | *01:21 |
| 4026 | 140 bp, 270 bp | 800 bp | *04:56, 04:64 |  |
| 41 | 280 bp | 1070 bp | *04:54 | *01:04, 14:02:01-14:25, 14:27-14:29 |
| 425,7 | 85 bp | 1070 bp | *04:69 | *16:12 |
| 437 | 170 bp | 1070 bp | *04:81 |  |
| 44 | 135 bp | 1070 bp | *04:60 |  |
| 45 | 135 bp | 1070 bp | *04:66 |  |
| 46 | 300 bp | 1070 bp | *04:67 |  |
| 477 | 185 bp | 1070 bp | *04:84 |  |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-C*04 SSP typings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective lengths of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most wells, or a band of 800 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-C*04 subtyping.

In addition, wells number 26, 27 and 40 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3The C*04:15:02, C*04:44 and C*04:47alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 20.

The C*04:23, C*04:38 and C*04:39 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 21.

The C*04:25, C*04:40 and C*04:41 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 23.

The C*04:46 and C*04:50 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 27.

The C*04:48 and C*04:75 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 28.

The C*04:56 and C*04:64 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 40.

The C*04:59Q and C*04:78 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 32.

The C*04:62 and C*04:76 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 35.

The C*04:65 and C*04:72 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 33.

4Due to the sharing of sequence motifs between HLA-C alleles some non-HLA-C*04 alleles will be amplified by primer mixes 1, 2, 4, 5, 7, 8, 10 to 12, 14, 16, 17, 20, 24, 30, 33, 34, 39, 41 and 42. In addition, the A*24:96 and A*24:146 alleles will be amplified by primer mix 15, the B*58:02 allele will be amplified by primer mixes 5 and 8, the B*07:90 allele will be amplified by primer mix 30 and the B*40:100 allele will be amplified by primer mixes 32 and 33.

5Short specific PCR fragments are less intense and not as sharp as longer specific bands.

6The primers in primer mixes 7, 21 and 24 may give rise to primer oligomer formation.

7Primer mixes 26, 42, 43 and 47 may have tendencies of unspecific amplification.

8Primer mix 9: Specific PCR fragment of 110 bp in the C*04:35 allele. Specific PCR fragment of 180 bp in the C*04:30 allele. Specific PCR fragment of 220 bp in the C*04:09N allele.

9Primer mix 10: Specific PCR fragment of 195 bp in the C*04:31 allele. Specific PCR fragment of 220 bp in the C*04:10, 04:11, 04:36 and 04:55 and in the C*15:36 alleles.

10Primer mix 12: Specific PCR fragment of 125 bp in the C*04:29, 04:36 and 04:55 and the C*07:125 alleles. Specific PCR fragment of 165 bp in the C*04:33 allele. Specific PCR fragment of 125 and 165 bp in the C*04:11 allele.

11Primer mix 13: Specific PCR fragment of 225 bp in the C*04:12 allele. Specific PCR fragment of 270 bp in the C*04:52 and 04:55 alleles.

12Primer mix 14: Specific PCR fragment of 155 bp in the C*04:16 allele. Specific PCR fragment of 190 bp in the C*04:13, 04:18 and 04:58 and the C*07:08 and 07:108 alleles.

13Primer mix 18: Specific PCR fragment of 125 bp in the C*04:70 allele. Specific PCR fragment of 225 bp in the C*04:19 allele.

14Primer mix 19: Specific PCR fragment of 120 bp in the C*04:35 and 04:37 alleles. Specific PCR fragment of 155 bp in the C*04:20 allele.

15Primer mix 20: Specific PCR fragment of 165 bp in the C*04:44 allele. Specific PCR fragment of 250 bp in the C*04:47 allele. Specific PCR fragment of 545 bp in the C*04:15:02 and 04:17 and the C*15:36 alleles.

16Primer mix 21: Specific PCR fragment of 85 bp in the C*04:23 allele. Specific PCR fragment of 145 bp in the C*04:38 allele. Specific PCR fragment of 240 bp in the C*04:39 allele.

17Primer mix 23: Specific PCR fragment of 85 bp in the C*04:25 allele. Specific PCR fragment of 145 bp in the C*04:40 allele. Specific PCR fragment of 215 bp in the C*04:41 allele.

18Primer mix 27: Specific PCR fragment of 125 bp in the C*04:50 allele. Specific PCR fragment of 285 bp in the C*04:46 allele.

19Primer mix 28: Specific PCR fragment of 120 bp in the C*04:75 allele. Specific PCR fragment of 215 bp in the C*04:48 allele.

20Primer mix 32: Specific PCR fragment of 105 bp in the C*04:78 allele. Specific PCR fragment of 235 bp in the C*04:59Q allele. Specific PCR fragment of 275 bp in the C*04:77 and in the B*40:100 alleles.

21Primer mix 33: Specific PCR fragment of 80 bp in the C*04:72 allele. Specific PCR fragment of 180 bp in the C*04:58 and 04:65 and the C*07:08 and 07:108 and in the B*40:100 alleles.

22Primer mix 35: Specific PCR fragment of 85 bp in the C*04:62 allele. Specific PCR fragment of 175 bp in the C*04:76 allele.

23Primer mix 37: Specific PCR fragment of 110 bp in the C*04:63 and 04:73 alleles. Specific PCR fragment of 135 bp in the C*04:74 allele.

24Primer mix 38: Specific PCR fragment of 105 bp in the C*04:83 allele. Specific PCR fragment of 135 bp in the C*04:74 allele.

25Primer mix 39: Specific PCR fragment of 145 bp in the C*04:71 and the C*01:21 alleles. Specific PCR fragment of 330 bp in the C*04:79 allele.

26Primer mix 40: Specific PCR fragment of 140 bp in the C*04:56 allele. Specific PCR fragment of 270 bp in the C*04:64 allele.

‘w’, may be weakly amplified.

1The nucleotide sequence of the Cw*0402 allele has been shown to be identical to C*04:01:01:01.

The nucleotide sequence of the Cw*0421 allele has been renamed C*04:15:02.

The nucleotide sequence of the Cw*0422 allele has been shown to be identical to C*04:21.

2The C*04:15:02, C*04:44 and C*04:47alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 20.

3The C*04:23, C*04:38 and C*04:39 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 21.

4The C*04:25, C*04:40 and C*04:41 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 23.

5The C*04:46 and C*04:50 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 27.

6The C*04:48 and C*04:75 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 28.

7The C*04:56 and C*04:64 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 40.

8The C*04:59Q and C*04:78 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 32.

9The C*04:62 and C*04:76 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 35.

10The C*04:65 and C*04:72 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 33.

11Primer mix 9: Specific PCR fragment of 110 bp in the C*04:35 allele. Specific PCR fragment of 180 bp in the C*04:30 allele. Specific PCR fragment of 220 bp in the C*04:09N allele.

Primer mix 10: Specific PCR fragment of 195 bp in the C*04:31 allele. Specific PCR fragment of 220 bp in the C*04:10, 04:11, 04:36 and 04:55 and in the C*15:36 alleles.

Primer mix 12: Specific PCR fragment of 125 bp in the C*04:29, 04:36 and 04:55 and the C*07:125 alleles. Specific PCR fragment of 165 bp in the C*04:33 allele. Specific PCR fragment of 125 and 165 bp in the C*04:11 allele.

Primer mix 13: Specific PCR fragment of 225 bp in the C*04:12 allele. Specific PCR fragment of 270 bp in the C*04:52 and 04:55 alleles.

Primer mix 14: Specific PCR fragment of 155 bp in the C*04:16 allele. Specific PCR fragment of 190 bp in the C*04:13, 04:18 and 04:58 and the C*07:08 and 07:108 alleles.

Primer mix 18: Specific PCR fragment of 125 bp in the C*04:70 allele. Specific PCR fragment of 225 bp in the C*04:19 allele.

Primer mix 19: Specific PCR fragment of 120 bp in the C*04:35 and 04:37 alleles. Specific PCR fragment of 155 bp in the C*04:20 allele.

Primer mix 20: Specific PCR fragment of 165 bp in the C*04:44 allele. Specific PCR fragment of 250 bp in the C*04:47 allele. Specific PCR fragment of 545 bp in the C*04:15:02 and 04:17 and the C*15:36 alleles.

Primer mix 21: Specific PCR fragment of 85 bp in the C*04:23 allele. Specific PCR fragment of 145 bp in the C*04:38 allele. Specific PCR fragment of 240 bp in the C*04:39 allele.

Primer mix 23: Specific PCR fragment of 85 bp in the C*04:25 allele. Specific PCR fragment of 145 bp in the C*04:40 allele. Specific PCR fragment of 215 bp in the C*04:41 allele.

Primer mix 27: Specific PCR fragment of 125 bp in the C*04:50 allele. Specific PCR fragment of 285 bp in the C*04:46 allele.

Primer mix 28: Specific PCR fragment of 120 bp in the C*04:75 allele. Specific PCR fragment of 215 bp in the C*04:48 allele.

Primer mix 32: Specific PCR fragment of 105 bp in the C*04:78 allele. Specific PCR fragment of 235 bp in the C*04:59Q allele. Specific PCR fragment of 275 bp in the C*04:77 and in the B*40:100 alleles.

Primer mix 33: Specific PCR fragment of 80 bp in the C*04:72 allele. Specific PCR fragment of 180 bp in the C*04:58 and 04:65 and the C*07:08 and 07:108 and in the B*40:100 alleles.

Primer mix 35: Specific PCR fragment of 85 bp in the C*04:62 allele. Specific PCR fragment of 175 bp in the C*04:76 allele.

Primer mix 37: Specific PCR fragment of 110 bp in the C*04:63 and 04:73 alleles. Specific PCR fragment of 135 bp in the C*04:74 allele.

Primer mix 38: Specific PCR fragment of 105 bp in the C*04:83 allele. Specific PCR fragment of 135 bp in the C*04:74 allele.

Primer mix 39: Specific PCR fragment of 145 bp in the C*04:71 and the C*01:21 alleles. Specific PCR fragment of 330 bp in the C*04:79 allele.

Primer mix 40: Specific PCR fragment of 140 bp in the C*04:56 allele. Specific PCR fragment of 270 bp in the C*04:64 allele.

12Primer mix 42 does not amplify the C*04:82 allele. Thus, this lot of the HLA-C*04 subtyping kit cannot distinguish the C*04:01:01:01-04:01:09, 04:01:11-04:01:22, 04:01:24-04:01:27 and 04:01:29-04:01:37 alleles and the C*04:82 allele.

‘w’, may be weakly amplified.

Change in revision R01 compared to R00:

- Primer mix 26 may have tendencies of unspecific amplification.

Change in revision R02 compared  to R01:

- Primer mix 42 does not amplify the C*04:82 allele. This has been corrected in the specificity and amplification tables. Thus, this lot of the HLA-C*04 subtyping kit cannot distinguish the C*04:01:01:01-04:01:09, 04:01:11-04:01:22, 04:01:24-04:01:27 and 04:01:29-04:01:37 alleles and the C*04:82 allele.