PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® HLA-C*04

Prod. No:	101.612-12/12u

Lot No:    	86S

Expiry Date:    	2016-May-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 48 primer mixes used for HLA-C*04 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified HLA-C*04 alleles3 | Other amplified HLA Class I alleles4 |
| --- | --- | --- | --- | --- |
| 17 | 250 bp | 800 bp | *04:01:01:01-04:01:33, 04:01:34w, 04:01:35-04:01:52, 04:03:01-04:03:02, 04:05, 04:07-04:12, 04:15:01-04:18, 04:19w, 04:20, 04:23-04:33, 04:35-04:57, 04:59Q-04:67, 04:69-04:94:02, 04:95Nw, 04:96-04:109, 04:110?, 04:112-04:121, 04:123N-04:152 | *01:02:01-01:21, 01:23-01:34, 01:36-01:78, 02:02:01-02:02:03, 02:02:05-02:02:26, 02:04-02:15, 02:17, 02:19-02:40, 02:42-02:68, 05:01:01:01-05:01:27, 05:03-05:10, 05:12-05:16, 05:18:01-05:26, 05:28-05:50, 05:52-05:61, 05:63-05:67, 05:69, 05:71-05:78, 05:80-05:94, 06:02:01:01-06:02:01:02, 06:02:03-06:03:02, 06:05-06:39, 06:41-06:78, 06:80-06:102, 08:02:01-08:02:10, 08:05, 08:07, 08:12, 08:15:01-08:15:02, 08:17-08:19, 08:23, 08:25, 08:27-08:35, 08:37, 08:43, 08:45, 08:47-08:49, 08:51-08:53, 08:55N, 08:62-08:63, 08:67-08:71, 08:73-08:77, 12:02:01-12:13, 12:15-12:17, 12:21-12:82, 12:84N-12:99, 14:02:01-14:05, 14:07N-14:14, 14:16-14:52, 15:08, 16:04:01, 16:29, 16:33, 16:42, 16:53, 16:55, 17:17, 18:01-18:07N |
| 2 | 220 bp | 1070 bp | *04:01:01:01-04:01:27, 04:01:29-04:01:52, 04:04:01-04:05, 04:07-04:09N, 04:12w, 04:13-04:15:03, 04:17-04:20, 04:23-04:35, 04:37-04:41, 04:43-04:54, 04:56-04:70, 04:72-04:79, 04:81-04:102, 04:104-04:106, 04:108-04:139, 04:141-04:146, 04:148-04:152 | *07:64 |
| 3 | 145 bp | 1070 bp | *04:01:01:01-04:01:52, 04:04:01-04:05, 04:08-04:15:03, 04:17-04:20, 04:23-04:26, 04:28-04:41, 04:43-04:79, 04:81-04:86, 04:88N, 04:90-04:102, 04:104-04:106, 04:108-04:139, 04:141-04:146, 04:148-04:152 |  |
| 47 | 210 bp | 1070 bp | *04:03:01-04:03:02, 04:06, 04:80, 04:107, 04:147 | *02:12, 02:49, 02:55 |
| 57 | 250 bp | 1070 bp | *04:04:01-04:04:02, 04:06, 04:13, 04:34, 04:58, 04:122 | *01:22, 01:35, 02:03, 02:16:01-02:16:02, 02:18, 05:11, 05:17, 05:27, 05:68, 05:79, 06:04, 08:01:01-08:01:10, 08:03:01-08:04:03, 08:06, 08:08-08:11, 08:13-08:14, 08:16:01-08:16:02, 08:20-08:22, 08:24, 08:26N, 08:36N, 08:38-08:42, 08:44, 08:46, 08:50, 08:54, 08:56-08:61, 08:65-08:66, 08:72:01-08:72:02, 08:78, 12:14:01-12:14:02, 12:18:01-12:18:02, 12:20, 12:83, 14:06, 14:15, 15:02:01-15:07, 15:09-15:13, 15:15-15:19, 15:21-15:24, 15:26-15:50, 15:52-15:71, 16:35, 16:40, 16:48, 17:01:01:01-17:16, 17:18-17:19 |
| 65,9 | 95 bp, 215 bp | 1070 bp | *04:05, 04:112 | *15:36 |
| 76 | 145 bp | 1070 bp | *04:07, 04:27, 04:32, 04:77, 04:89 | *03:34, 03:142, 05:78, 07:64, 14:25, 15:36, 18:03 |
| 8 | 270 bp | 1070 bp | *04:08, 04:34, 04:147 | *01:35, 02:58, 05:27, 05:39, 06:96, 08:41, 12:83, 14:20, 15:15, 17:07 |
| 95,10 | 110 bp, 180 bp, 220 bp | 800 bp | *04:09N, 04:30, 04:35 |  |
| 1011 | 195 bp, 220 bp | 1070 bp | *04:10-04:11, 04:31, 04:36, 04:55 | *05:78, 15:36 |
| 117 | 180 bp | 1070 bp | *04:01:01:01-04:01:09, 04:01:11-04:01:22, 04:01:24-04:01:52, 04:03:01-04:10, 04:12-04:20, 04:23-04:26, 04:28-04:32, 04:34-04:51, 04:53-04:54, 04:56-04:106, 04:108-04:115N, 04:117-04:129, 04:131-04:152 | *02:49, 05:25, 05:42, 06:05, 06:76:02, 07:02:09, 08:28, 12:28, 15:25, 15:62, 16:26, 16:46, 16:55 |
| 125,12 | 125 bp, 165 bp | 1070 bp | *04:11, 04:29, 04:33, 04:36, 04:55 | *07:125 |
| 1313 | 225 bp, 270 bp | 1070 bp | *04:12, 04:52, 04:55 |  |
| 1414 | 155 bp, 185 bp | 1070 bp | *04:16, 04:18 |  |
| 1515 | 180 bp, 235 bp | 1070 bp | *04:14, 04:28, 04:68, 04:144 | *03:171, 05:93, 06:73, 08:20, 08:40, 16:45, A*24:96, A*24:146 |
| 165,16 | 85 bp, 130 bp | 1070 bp | *04:15:01-04:15:03, 04:17, 04:37, 04:123N | *03:05, 03:13:01, 03:25, 03:27, 03:35, 03:135, 03:167, 08:01:07, 08:02:07, 08:33:02, 14:09, 14:45, 18:07N, B*15:78:03 |
| 17 | 320 bp | 1070 bp | *04:17, 04:80, 04:100 | *01:50 |
| 185,17 | 125 bp, 220 bp | 1070 bp | *04:19, 04:70, 04:94:01-04:94:02 | *06:101, 12:10:01-12:10:02, 18:03, B*15:27:01-15:27:03, B*15:109 |
| 195,18 | 120 bp, 155 bp | 1070 bp | *04:20, 04:35, 04:37 | *03:135, 14:45 |
| 2019 | 165 bp, 250 bp, 545 bp | 1070 bp | *04:15:02, 04:17, 04:44, 04:47, 04:100 | *05:78, 15:36, A*01:118, A*02:109, A*33:52 |
| 215,20 | 85 bp, 145 bp, 240 bp | 1070 bp | *04:23, 04:38-04:39, 04:108 |  |
| 225,21 | 120 bp, 170 bp | 1070 bp | *04:24, 04:26, 04:139-04:140 | *07:125 |
| 235,22 | 85 bp, 145 bp, 225 bp | 1070 bp | *04:25, 04:40-04:41, 04:144 | *03:171, 05:93, 06:73, 08:20, 08:40 |
| 246 | 170 bp | 1070 bp | *04:30, 04:42 | *07:125 |
| 25 | 200 bp | 1070 bp | *04:43, 04:94:01-04:94:02 | *06:101, 12:10:01-12:10:02, 18:03, B*15:27:01-15:27:03, B*15:109 |
| 267 | 215 bp | 800 bp | *04:45 |  |
| 275,23 | 125 bp, 280 bp | 800 bp | *04:46, 04:50, 04:120 | *05:64:01-05:64:02, 08:19 |
| 285,8,24 | 120 bp, 215 bp | 1070 bp | *04:48, 04:75 |  |
| 295,25 | 105 bp, 185 bp | 1070 bp | *04:49, 04:82 |  |
| 305 | 75 bp | 1070 bp | *04:53 | *05:49, B*07:90 |
| 316,26 | 130 bp, 180 bp | 1070 bp | *04:51, 04:95N, 04:139 |  |
| 325,27 | 105 bp, 235 bp, 275 bp | 1070 bp | *04:59Q, 04:77-04:78 | B*40:100 |
| 335,28 | 80 bp, 180 bp | 1070 bp | *04:58, 04:65, 04:72 | *07:08, 07:108, B*40:100 |
| 345,29 | 75 bp, 200 bp, 275 bp | 1070 bp | *04:13, 04:58, 04:61, 04:68, 04:96, 04:120 | *05:64:01-05:64:02, 07:08, 07:108, 08:19 |
| 355,30 | 85 bp, 145 bp, 175 bp | 1070 bp | *04:62, 04:76, 04:115N |  |
| 365 | 115 bp | 1070 bp | *04:57, 04:63 |  |
| 375,31 | 110 bp, 135 bp | 1070 bp | *04:63, 04:73-04:74 |  |
| 385,32 | 95 bp, 135 bp | 1070 bp | *04:74, 04:83, 04:123N | *18:07N |
| 3933 | 135 bp, 330 bp | 1070 bp | *04:71, 04:79, 04:95N | *01:21 |
| 4034 | 140 bp, 270 bp | 800 bp | *04:56, 04:64:01-04:64:02 |  |
| 416,35 | 135 bp, 280 bp | 1070 bp | *04:54, 04:105N | *01:04, 01:54, 14:02:01-14:24:01, 14:25, 14:27-14:52 |
| 425 | 95 bp | 1070 bp | *04:69, 04:82 | *16:12 |
| 437,36 | 145 bp, 170 bp, 260 bp | 1070 bp | *04:81, 04:88N, 04:115N |  |
| 445,37 | 90 bp, 135 bp | 1070 bp | *04:60, 04:108 | *03:81, 03:175, B*15:78:03 |
| 4538 | 135 bp, 165 bp | 1070 bp | *04:16, 04:26, 04:66, 04:103 |  |
| 465,39 | 125 bp, 300 bp | 1070 bp | *04:67, 04:93N |  |
| 475,7,40 | 50 bp, 295 bp | 800 bp | *04:114, 04:146 | *01:59, 02:65, 03:130, 03:140, 05:20, 06:82, 07:49, 07:210, 07:238, 07:247, 12:54, 14:04, 16:57, A*68:46 |
| 4841 | 155 bp, 240 bp | 800 bp | *04:84, 04:106 | *07:167, 08:78 |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-C*04 SSP typings.

When the primers in a primer mix can give rise to HLA-specific PCR products of more than one length this is indicated if the size difference is more than 20 base pairs. Size differences of 20 base pairs or less are not given. For high resolution SSP kits the respective lengths of the HLA-specific PCR product(s) are given for the alleles amplified by these primer mixes.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most wells, or a band of 800 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-C*04 subtyping.

In addition, wells number 9, 26, 27, 40, 47 and 48  contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3For several HLA Class I alleles 1st and/or 4th exon(s) and beyond, as well as intron nucleotide sequences, are not available. In these instances it is not known whether some of the primers of the SSP sets are completely matched with the target sequences or not. We assume that unknown sequences in these regions are conserved within allelic groups.

4Due to the sharing of sequence motifs between HLA-C alleles non-HLA-C*04 alleles will be amplified by primer mixes 1, 2, 4 to 8, 10 to 12, 15 to 20, 22 to 25, 27, 30, 33, 34, 38, 39, 41, 42, 44, 47 and 48. In addition, the A*01:118, A*02:109 and A*33:52 alleles will be amplified by primer mix 20, the A*24:96 and A*24:146 alleles will be amplified by primer mix 15, the A*68:46 allele will be amplified by primer mix 47, the B*07:90 allele will be amplified by primer mix 30, the B*15:27:01-15:27:03 and 15:109 alleles will be amplified by primer mixes 18 and 25, the B*15:78:03 allele will be amplified by primer mixes 16 and 44, the B*40:100 allele will be amplified by primer mixes 32 and 33.

5HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

6Primer mixes 7, 24, 31 and 41 have a tendency to giving rise to primer oligomer formation.

7Primer mixes 1, 4, 5, 11, 26, 43 and 47 may have tendencies of unspecific amplification.

8Primer mix, 28 may give rise to a long unspecific amplification product of approximately 500 bp. This should be disregarded when interpreting the C*04 typings.

9Primer mix 6: Specific PCR fragment of 95 bp in the C*04:05 allele. Specific PCR fragment of 215 bp in the C*04:112 and the C*15:36 allele.

10Primer mix 9: Specific PCR fragment of 110 bp in the C*04:35 allele. Specific PCR fragment of 180 bp in the C*04:30 allele. Specific PCR fragment of 220 bp in the C*04:09N allele.

11Primer mix 10: Specific PCR fragment of 195 bp in the C*04:31 allele. Specific PCR fragment of 220 bp in the C*04:10-04:11, 04:36 and 04:55 and the C*05:78 and 15:36 alleles.

12Primer mix 12: Specific PCR fragment of 125 bp in the C*04:29, 04:36 and 04:55 and the C*07:125 alleles. Specific PCR fragment of 165 bp in the C*04:33. Specific PCR fragment of 125 and 165 bp in the C*04:11 allele.

13Primer mix 13: Specific PCR fragment of 225 bp in the C*04:12 allele. Specific PCR fragment of 270 bp in the C*04:52 and 04:55 alleles.

14Primer mix 14: Specific PCR fragment of 155 bp in the C*04:16 allele. Specific PCR fragment of 185 bp in the C*04:18 allele.

15Primer mix 15: Specific PCR fragment of 180 bp in the C*04:14, 04:28 and 04:68 and the C*16:45 and in the A*24:96 and 24:146 alleles. Specific PCR fragment of 235 bp in the C*04:144 and the C*03:171, 05:93, 06:73, 08:20 and 08:40 alleles.

16Primer mix 16: Specific PCR fragment of 85 bp in the C*04:123N and the C*18:07N allele. Specific PCR fragment of 130 bp in the C*04:15:01-04:15:03, 04:17 and 04:37 and the C*03:05, 03:13:01, 03:25, 03:27, 03:35, 03:135, 03:167, 08:01:07, 08:02:07, 08:33:02, 14:09 and 14:45 and in the B*15:78:03 alleles.

17Primer mix 18: Specific PCR fragment of 125 bp in the C*04:70 allele. Specific PCR fragment of 220 bp in the C*04:19 and 04:94:01-04:94:02 and the C* 06:101, 12:10:01-12:10:02 and 18:03 and in the B*15:27:01-15:27:03 and B*15:109 alleles.

18Primer mix 19: Specific PCR fragment of 120 bp in the C*04:35 and 04:37 and the C*14:45 alleles. Specific PCR fragment of 155 bp in the C*04:20 and the C*03:135 alleles.

19Primer mix 20: Specific PCR fragment of 165 bp in the C*04:44 allele. Specific PCR fragment of 250 bp in the C*04:47 allele. Specific PCR fragment of 545 bp in the C*04:15:02, 04:17 and 04:100 and in the C*05:78 and 15:36 and in the A*01:118, 02:109 and 33:52 alleles.

20Primer mix 21: Specific PCR fragment of 85 bp in the C*04:23 and 04:108 alleles. Specific PCR fragment of 145 bp in the C*04:38 allele. Specific PCR fragment of 240 bp in the C*04:39 allele.

21Primer mix 22: Specific PCR fragment of 120 bp in the C*04:24, 04:139 and 04:140 and the C*07:125 alleles. Specific PCR fragment of 170 bp in the C*04:26 allele.

22Primer mix 23: Specific PCR fragment of 85 bp in the C*04:25 allele. Specific PCR fragment of 145 bp in the C*04:40 allele. Specific PCR fragment of 225 bp in the C*04:41, 04:144 and the C*03:171, 05:93, 06:73, 08:20 and 08:40 alleles.

23Primer mix 27: Specific PCR fragment of 125 bp in the C*04:50 allele. Specific PCR fragment of 280 bp in the C*04:46 and 04:120 and the C*05:64:01-05:64:02 and 08:19 alleles.

24Primer mix 28: Specific PCR fragment of 120 bp in the C*04:75 allele. Specific PCR fragment of 215 bp in the C*04:48 allele.

25Primer mix 29: Specific PCR fragment of 105 bp in the C*04:82 allele. Specific PCR fragment of 185 bp in the C*04:49 allele.

26Primer mix 31: Specific PCR fragment of 130 bp in the C*04:95N and 04:139 alleles. Specific PCR fragment of 180 bp in the C*04:51 allele.

27Primer mix 32: Specific PCR fragment of 105 bp in the C*04:78 allele. Specific PCR fragment of 235 bp in the C*04:59Q allele. Specific PCR fragment of 275 bp in the C*04:77 and in the B*40:100 alleles.

28Primer mix 33: Specific PCR fragment of 80 bp in the C*04:72 allele. Specific PCR fragment of 180 bp in the C*04:58 and 04:65 and the C*07:08 and 07:108 and in the B*40:100 alleles.

29Primer mix 34: Specific PCR fragment of 75 bp in the C*04:96 allele. Specific PCR fragment of 200 bp in the C*04:13, 04:58, 04:61 and 04:68 and the C*07:08 and 07:108 alleles. Specific PCR fragment of 275 bp in the C*04:120 and the C*05:64:01-05:64:02 and 08:19 alleles.

30Primer mix 35: Specific PCR fragment of 85 bp in the C*04:62 allele. Specific PCR fragment of 145 bp in the C*04:115N allele. Specific PCR fragment of 175 bp in the C*04:76 allele.

31Primer mix 37: Specific PCR fragment of 110 bp in the C*04:63 and 04:73 alleles. Specific PCR fragment of 135 bp in the C*04:74 allele.

32Primer mix 38: Specific PCR fragment of 95 bp in the C*04:83 and 04:123N and the C*18:07N alleles. Specific PCR fragment of 135 bp in the C*04:74 allele.

33Primer mix 39: Specific PCR fragment of 135 bp in the C*04:71 and 04:95N and the C*01:21 alleles. Specific PCR fragment of 330 bp in the C*04:79 allele.

34Primer mix 40: Specific PCR fragment of 140 bp in the C*04:56 allele. Specific PCR fragment of 270 bp in the C*04:64:01-04:64:02 alleles.

35Primer mix 41: Specific PCR fragment of 135 bp in the C*04:105N allele. Specific PCR fragment of 280 bp in the C*04:54 and the C*01:04, 01:54, 14:02:01-14:24:01, 14:25 and 14:27-14:52 alleles.

36Primer mix 43: Specific PCR fragment of 145 bp in the C*04:115N allele. Specific PCR fragment of 170 bp in the C*04:81 allele. Specific PCR fragment of 260 bp in the C*04:88N allele.

37Primer mix 44: Specific PCR fragment of 90 bp in the C*04:108 and in the C*03:81, 03:175 and in the  B*15:78:03 alleles. Specific PCR fragment of 135 bp in the C*04:60 allele.

38Primer mix 45: Specific PCR fragment of 135 bp in the C*04:66 allele. Specific PCR fragment of 165 bp in the C*04:16, 04:26 and 04:103 alleles.

39Primer mix 46: Specific PCR fragment of 125 bp in the C*04:93N allele. Specific PCR fragment of 300 bp in the C*04:67 allele.

40Primer mix 47: Specific PCR fragment of 50 bp in the C*04:114 and the C*01:59, 02:65, 03:130, 03:140, 05:20, 06:82, 07:49, 07:210, 07:238, 07:247, 12:54, 14:04, 16:57 and in the A*68:46 alleles. Specific PCR fragment of 295 bp in the C *04:146 allele.

41Primer mix 48: Specific PCR fragment of 155 bp in the C*04:84 and the C*07:167 alleles. Specific PCR fragment of 240 bp in the C*04:106 and the C*08:78 allele.

‘w’, may be weakly amplified.

1The nucleotide sequence of the Cw*0402 allele has been shown to be identical to C*04:01:01:01.

The nucleotide sequence of the Cw*0421 allele has been renamed C*04:15:02.

The nucleotide sequence of the Cw*0422 allele has been shown to be identical to C*04:21.

2HLA-C*04 alleles in bold lettering are listed as confirmed alleles on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.13.1, July 2013.

3The HLA-C*04 primer set cannot separate the C*04:03:01-04:03:02 and 02:49 alleles. These alleles can be distinguished by the HLA-C low resolution kit and/or the HLA-C*02 high resolution kit.

4The HLA-C*04 primer set cannot separate the C*04:107 and the C*02:12 and 02:55 alleles. These alleles can be distinguished by the HLA-C low resolution kit and/or the HLA-C*02 high resolution kit.

5The C*04:05 and C*04:112 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 6.

6The C*04:44 and C*04:47alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 20.

7The C*04:23, C*04:38 and C*04:39 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 21.

8The C*04:25, C*04:40 and C*04:41 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 23.

9The C*04:46 and C*04:50 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 27.

10The C*04:48 and C*04:75 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 28.

11The C*04:54 and C*04:105N alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 41.

12The C*04:56 and C* 04:64:01-04:64:02 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 40.

13The C*04:59Q and C*04:78 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 32.

14The C*04:61 and C*04:96 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 34.

15The C*04:62 and C*04:76 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 35.

16The C*04:65 and C*04:72 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 33.

17The C*04:67 and C*04:93N alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 46.

18The C*04:81 and C*04:88N alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 43.

19The C*04:84 and C*04:106 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 48.

20The C*04:114 and 04:146 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 47.

21Primer mix 6: Specific PCR fragment of 95 bp in the C*04:05 allele. Specific PCR fragment of 215 bp in the C*04:112 and the C*15:36 allele.

Primer mix 9: Specific PCR fragment of 110 bp in the C*04:35 allele. Specific PCR fragment of 180 bp in the C*04:30 allele. Specific PCR fragment of 220 bp in the C*04:09N allele.

Primer mix 10: Specific PCR fragment of 195 bp in the C*04:31 allele. Specific PCR fragment of 220 bp in the C*04:10-04:11, 04:36 and 04:55 and the C*05:78 and 15:36 alleles.

Primer mix 12: Specific PCR fragment of 125 bp in the C*04:29, 04:36 and 04:55 and the C*07:125 alleles. Specific PCR fragment of 165 bp in the C*04:33. Specific PCR fragment of 125 and 165 bp in the C*04:11 allele.

Primer mix 13: Specific PCR fragment of 225 bp in the C*04:12 allele. Specific PCR fragment of 270 bp in the C*04:52 and 04:55 alleles.

Primer mix 14: Specific PCR fragment of 155 bp in the C*04:16 allele. Specific PCR fragment of 185 bp in the C*04:18 allele.

Primer mix 15: Specific PCR fragment of 180 bp in the C*04:14, 04:28 and 04:68 and the C*16:45 and in the A*24:96 and 24:146 alleles. Specific PCR fragment of 235 bp in the C*04:144 and the C*03:171, 05:93, 06:73, 08:20 and 08:40 alleles.

Primer mix 16: Specific PCR fragment of 85 bp in the C*04:123N and the C*18:07N allele. Specific PCR fragment of 130 bp in the C*04:15:01-04:15:03, 04:17 and 04:37 and the C*03:05, 03:13:01, 03:25, 03:27, 03:35, 03:135, 03:167, 08:01:07, 08:02:07, 08:33:02, 14:09 and 14:45 and in the B*15:78:03 alleles.

Primer mix 18: Specific PCR fragment of 125 bp in the C*04:70 allele. Specific PCR fragment of 220 bp in the C*04:19 and 04:94:01-04:94:02 and the C* 06:101, 12:10:01-12:10:02 and 18:03 and in the B*15:27:01-15:27:03 and B*15:109 alleles.

Primer mix 19: Specific PCR fragment of 120 bp in the C*04:35 and 04:37 and the C*14:45 alleles. Specific PCR fragment of 155 bp in the C*04:20 and the C*03:135 alleles.

Primer mix 20: Specific PCR fragment of 165 bp in the C*04:44 allele. Specific PCR fragment of 250 bp in the C*04:47 allele. Specific PCR fragment of 545 bp in the C*04:15:02, 04:17 and 04:100 and in the C*05:78 and 15:36 and in the A*01:118, 02:109 and 33:52 alleles.

Primer mix 21: Specific PCR fragment of 85 bp in the C*04:23 and 04:108 alleles. Specific PCR fragment of 145 bp in the C*04:38 allele. Specific PCR fragment of 240 bp in the C*04:39 allele.

Primer mix 22: Specific PCR fragment of 120 bp in the C*04:24, 04:139 and 04:140 and the C*07:125 alleles. Specific PCR fragment of 170 bp in the C*04:26 allele.

Primer mix 23: Specific PCR fragment of 85 bp in the C*04:25 allele. Specific PCR fragment of 145 bp in the C*04:40 allele. Specific PCR fragment of 225 bp in the C*04:41, 04:144 and the C*03:171, 05:93, 06:73, 08:20 and 08:40 alleles.

Primer mix 27: Specific PCR fragment of 125 bp in the C*04:50 allele. Specific PCR fragment of 280 bp in the C*04:46 and 04:120 and the C*05:64:01-05:64:02 and 08:19 alleles.

Primer mix 28: Specific PCR fragment of 120 bp in the C*04:75 allele. Specific PCR fragment of 215 bp in the C*04:48 allele.

Primer mix 29: Specific PCR fragment of 105 bp in the C*04:82 allele. Specific PCR fragment of 185 bp in the C*04:49 allele.

Primer mix 31: Specific PCR fragment of 130 bp in the C*04:95N and 04:139 alleles. Specific PCR fragment of 180 bp in the C*04:51 allele.

Primer mix 32: Specific PCR fragment of 105 bp in the C*04:78 allele. Specific PCR fragment of 235 bp in the C*04:59Q allele. Specific PCR fragment of 275 bp in the C*04:77 and in the B*40:100 alleles.

Primer mix 33: Specific PCR fragment of 80 bp in the C*04:72 allele. Specific PCR fragment of 180 bp in the C*04:58 and 04:65 and the C*07:08 and 07:108 and in the B*40:100 alleles.

Primer mix 34: Specific PCR fragment of 75 bp in the C*04:96 allele. Specific PCR fragment of 200 bp in the C*04:13, 04:58, 04:61 and 04:68 and the C*07:08 and 07:108 alleles. Specific PCR fragment of 275 bp in the C*04:120 and the C*05:64:01-05:64:02 and 08:19 alleles.

Primer mix 35: Specific PCR fragment of 85 bp in the C*04:62 allele. Specific PCR fragment of 145 bp in the C*04:115N allele. Specific PCR fragment of 175 bp in the C*04:76 allele.

Primer mix 37: Specific PCR fragment of 110 bp in the C*04:63 and 04:73 alleles. Specific PCR fragment of 135 bp in the C*04:74 allele.

Primer mix 38: Specific PCR fragment of 95 bp in the C*04:83 and 04:123N and the C*18:07N alleles. Specific PCR fragment of 135 bp in the C*04:74 allele.

Primer mix 39: Specific PCR fragment of 135 bp in the C*04:71 and 04:95N and the C*01:21 alleles. Specific PCR fragment of 330 bp in the C*04:79 allele.

Primer mix 40: Specific PCR fragment of 140 bp in the C*04:56 allele. Specific PCR fragment of 270 bp in the C*04:64:01-04:64:02 alleles.

Primer mix 41: Specific PCR fragment of 135 bp in the C*04:105N allele. Specific PCR fragment of 280 bp in the C*04:54 and the C*01:04, 01:54, 14:02:01-14:24:01, 14:25 and 14:27-14:52 alleles.

Primer mix 43: Specific PCR fragment of 145 bp in the C*04:115N allele. Specific PCR fragment of 170 bp in the C*04:81 allele. Specific PCR fragment of 260 bp in the C*04:88N allele.

Primer mix 44: Specific PCR fragment of 90 bp in the C*04:108 and in the C*03:81, 03:175 and in the  B*15:78:03 alleles. Specific PCR fragment of 135 bp in the C*04:60 allele.

Primer mix 45: Specific PCR fragment of 135 bp in the C*04:66 allele. Specific PCR fragment of 165 bp in the C*04:16, 04:26 and 04:103 alleles.

Primer mix 46: Specific PCR fragment of 125 bp in the C*04:93N allele. Specific PCR fragment of 300 bp in the C*04:67 allele.

Primer mix 47: Specific PCR fragment of 50 bp in the C*04:114 and the C*01:59, 02:65, 03:130, 03:140, 05:20, 06:82, 07:49, 07:210, 07:238, 07:247, 12:54, 14:04, 16:57 and in the A*68:46 alleles. Specific PCR fragment of 295 bp in the C *04:146 allele.

Primer mix 48: Specific PCR fragment of 155 bp in the C*04:84 and the C*07:167 alleles. Specific PCR fragment of 240 bp in the C*04:106 and the C*08:78 allele.

‘w’, may be weakly amplified.

Change in revision R01 compared to R00:

- The HLA-C*04:15:02 and 04:17 and the C*03:13:01, 03:35, 08:01:07, 08:02:07 and 08:33:02 alleles are amplified by primer mix 16. This has been corrected in the Specificity and Interpretation Tables.

Change in revision R02 compared to R01:

- Primer mixes 5 and 8 do not amplify the B*58:02 allele. This has been corrected in the Specificity and Interpretation Tables.