Gel Documentation Form and Worksheet

HLA-C*04 (101.612-12/12u) 	Lot No: 67X	Expiry Date: 2017-May-01

Name:__________________

Sample ID:______________

DNA Conc.(ng/ul):_________

Test  Date:_______________

Tested By:_______________

Review Date:_______________

Reviewed By:_______________

Interpretation:___________	       Failed lanes: ____________	   Comments:_______________

Gel Picture

PHOTO DOCUMENT

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 7, 24, 31 and 41 have a tendency to giving rise to primer oligomer formation.

Primer mixes 1, 4, 5, 11, 15, 23, 26, 43, 47 and 48 may have tendencies of unspecific amplification, most pronounced in primer mix 15.

Primer mix, 28 may give rise to a long unspecific amplification product of approximately 500 bp. This should be disregarded when interpreting the C*04 typings.

Primer mix 53 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 430 base pairs.

1HLA-C*04 alleles in bold lettering are listed as confirmed alleles on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.17.0, July 2014.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3Primer mix 6: Specific PCR fragment of 95 bp in the C*04:05 allele. Specific PCR fragment of 215 bp in the C*04:112 and 04:169 and the C*15:36 alleles.

Primer mix 9: Specific PCR fragment of 110 bp in the C*04:35 allele. Specific PCR fragment of 180 bp in the C*04:30 allele. Specific PCR fragment of 220 bp in the C*04:09N allele.

Primer mix 10: Specific PCR fragment of 190 bp in the C*04:31 and 04:91 alleles. Specific PCR fragment of 220 bp in the C*04:10-04:11, 04:36, 04:55, 04:153 and 04:169 and the C*03:231, 05:78 and 15:36 alleles.

Primer mix 12: Specific PCR fragment of 125 bp in the C*04:29, 04:36 and 04:55 and the C*03:231, 07:125 and 07:356 alleles. Specific PCR fragment of 165 bp in the C*04:33 and 04:169 alleles. Specific PCR fragment of 125 and 165 bp in the C*04:11 and 04:172 and the C*03:248 and 16:62 alleles.

Primer mix 13: Specific PCR fragment of 215 bp in the C*04:12 and 04:132 allele. Specific PCR fragment of 270 bp in the C*04:52 and 04:55 and the C*03:231 alleles.

Primer mix 14: Specific PCR fragment of 155 bp in the C*04:16 and 04:163 and the C*03:248 and in the A*24:52 alleles. Specific PCR fragment of 185 bp in the C*04:18 allele.

Primer mix 15: Specific PCR fragment of 180 bp in the C*04:14, 04:28 and 04:68 and the C*16:45 and in the A*24:96 and 24:146 alleles. Specific PCR fragment of 235 bp in the C*04:144 and the C*03:171, 03:211:01, 05:93, 06:73, 08:20, 08:40 and 12:109 alleles.

Primer mix 16: Specific PCR fragment of 85 bp in the C*04:123N and the C*18:07N allele. Specific PCR fragment of 130 bp in the C*04:15:01-04:15:03, 14:17 and 04:37 and the C*03:05, 03:25, 03:13:01-03:13:02, 03:25, 03:27, 03:35, 03:135, 03:167, 08:01:07, 08:02:07, 08:33:02, 14:09, 14:45 and in the B*15:78:03 alleles.

Primer mix 18: Specific PCR fragment of 125 bp in the C*04:70 allele. Specific PCR fragment of 220 bp in the C*04:19 and 04:94:01-04:94:02 and the C*06:101, 12:10:01-12:10:02 and 18:03 and in the B*15:27:01-15:27:03, B*15:109 and 15:237 alleles.

Primer mix 19: Specific PCR fragment of 120 bp in the C*04:35 and 04:37 and the C*14:45 alleles. Specific PCR fragment of 155 bp in the C*04:20 and the C*03:135 alleles.

Primer mix 20: Specific PCR fragment of 165 bp in the C*04:44 allele. Specific PCR fragment of 250 bp in the C*04:47 and 04:170N alleles. Specific PCR fragment of 545 bp in the C*04:15:02, 04:17 and 04:100 and in the C*03:231, 05:78 and 15:36 and in the A*01:118, 02:109 and 33:52 alleles.

Primer mix 21: Specific PCR fragment of 85 bp in the C*04:23 and 04:108 alleles. Specific PCR fragment of 145 bp in the C*04:38 allele. Specific PCR fragment of 240 bp in the C*04:39 and 04:121 alleles.

Primer mix 22: Specific PCR fragment of 120 bp in the C*04:24, 04:139-04:140 and 04:166 and the C*07:125 and 07:356 alleles. Specific PCR fragment of 170 bp in the C*04:26 allele.

Primer mix 23: Specific PCR fragment of 85 bp in the C*04:25 allele. Specific PCR fragment of 145 bp in the C*04:40 allele. Specific PCR fragment of 225 bp in the C*04:41, 04:144 and the C*03:171, 03:211:01 05:93, 06:73, 08:20, 08:40 and 12:109 alleles.

Primer mix 25: Specific PCR fragment of 160 bp in the C*04:163 and in the A*24:52 alleles. Specific PCR fragment of 200 bp in the C*04:43, 04:94:01-04:94:02 and 04:171 and the C*06:101, 12:10:01-12:10:02 and 18:03 and in the B*15:27:01-15:27:03, 15:109 and 15:327 alleles.

Primer mix 27: Specific PCR fragment of 125 bp in the C*04:50 allele. Specific PCR fragment of 280 bp in the C*04:46 and 04:120 and the C*05:64:01-05:64:02, 08:19 and 08:101 alleles.

Primer mix 28: Specific PCR fragment of 120 bp in the C*04:75 allele. Specific PCR fragment of 215 bp in the C*04:48 allele.

Primer mix 29: Specific PCR fragment of 105 bp in the C*04:82 and 04:159 alleles. Specific PCR fragment of 195 bp in the C*04:49 and 04:132 alleles. Specific PCR fragment of 255 bp in the C*04:170N allele.

Primer mix 31: Specific PCR fragment of 125 bp in the C*04:95N and 04:139 alleles. Specific PCR fragment of 180 bp in the C*04:51 allele.

Primer mix 32: Specific PCR fragment of 105 bp in the C*04:78 allele. Specific PCR fragment of 240 bp in the C*04:59Q and 04:121 alleles. Specific PCR fragment of 275 bp in the C*04:77 and in the B*40:100 alleles.

Primer mix 33: Specific PCR fragment of 80 bp in the C*04:72 allele. Specific PCR fragment of 180 bp in the C*04:58, 04:65 and 04:160 and the C*07:08 and 07:108:01-07:108:02 and in the B*40:100 alleles.

Primer mix 34: Specific PCR fragment of 75 bp in the C*04:96 allele. Specific PCR fragment of 200 bp in the C*04:13, 04:58, 04:61, 04:68 and 04:160 and the C*07:08 and 07:108:01-07:108:02 and in the B*47:09 alleles. Specific PCR fragment of 275 bp in the C*04:120 and the C*05:64:01-05:64:02 and 08:19 alleles. Specific PCR fragment of 200 bp and 275 bp in the C*08:101 alleles.

Primer mix 35: Specific PCR fragment of 85 bp in the C*04:62 allele. Specific PCR fragment of 145 bp in the C*04:115N allele. Specific PCR fragment of 175 bp in the C*04:76 and 04:137 alleles.

Primer mix 36: Specific PCR fragment of 115 bp in the C*04:57 and 04:63 alleles. Specific PCR fragment of 150 bp in the C*04:117 allele.

Primer mix 37: Specific PCR fragment of 110 bp in the C*04:63 and 04:73 alleles. Specific PCR fragment of 135 bp in the C*04:74 and 04:125 alleles.

Primer mix 38: Specific PCR fragment of 95 bp in the C*04:83 and 04:123N and the C*03:232 and 18:07N alleles. Specific PCR fragment of 140 bp in the C*04:74 and 04:117 alleles.

Primer mix 39: Specific PCR fragment of 110 bp in the C*04:113 allele. Specific PCR fragment of 135 bp in the C*04:71 and 04:95N and the C*01:21 alleles. Specific PCR fragment of 330 bp in the C*04:79 allele.

Primer mix 40: Specific PCR fragment of 140 bp in the C*04:56 allele. Specific PCR fragment of 205 bp in the C*04:86 allele. Specific PCR fragment of 270 bp in the C*04:64:01-04:64:02 alleles.

Primer mix 41: Specific PCR fragment of 135 bp in the C*04:105N allele. Specific PCR fragment of 165 bp in the C*04:131 allele. Specific PCR fragment of 280 bp in the C*04:54 and the C*01:04, 01:54, 01:97, 14:02:01-14:24:01, 14:25, 14:27-14:53 and 14:56-14:67 alleles.

Primer mix 43: Specific PCR fragment of 145 bp in the C*04:115N allele. Specific PCR fragment of 170 bp in the C*04:81 and 04:137 alleles. Specific PCR fragment of 260 bp in the C*04:88N allele.

Primer mix 44: Specific PCR fragment of 90 bp in the C*04:108 and in the C*03:81, 03:175, 03:199 and 03:245 and in the  B*15:78:03 alleles. Specific PCR fragment of 135 bp in the C*04:60 and 04:125 alleles.

Primer mix 45: Specific PCR fragment of 135 bp in the C*04:66 allele. Specific PCR fragment of 165 bp in the C*04:16, 04:26 and 04:103 and the C*03:248 and 16:62 alleles.

Primer mix 46: Specific PCR fragment of 125 bp in the C*04:93N allele. Specific PCR fragment of 165 bp in the C*04:131 allele. Specific PCR fragment of 300 bp in the C*04:67 allele.

Primer mix 47: Specific PCR fragment of 50 bp in the C*04:114 and the C*01:59, 02:65, 03:130, 03:140, 03:243, 05:20, 06:82, 07:49, 07:210, 07:238, 07:247, 12:54, 14:04, 14:64, 15:85 and 16:57 and in the A*68:46 alleles. Specific PCR fragment of 295 bp in the C*04:146 and the C*03:205w alleles.

Primer mix 50: Specific PCR fragment of 295 bp in the C*03:205w allele. Specific PCR fragment of 335 bp in the C*04:162 allele.

Primer mix 52: Specific PCR fragment of 110 bp in the C*04:150 allele. Specific PCR fragment of 240 bp in the C*04:155 allele.

4The HLA-C*04 primer set cannot separate the C*04:03:01-04:03:02 and 02:49 alleles or the C*04:107 and the C*02:12 and 02:55 alleles. These alleles can be distinguished by the HLA-C low resolution kit and/or the HLA-C*02 high resolution kit.

The C*04:115N and C*04:137 alleles cannot be separated by this lot of the HLA-C*04 high resolution kit.

The C*04:106 and the C*04:01:01:01-04:01:09, 04:01:11-04:01:22, 04:01:24-04:01:27, 04:01:29-04:01:33, 04:01:35-04:01:62, 04:85, 04:90, 04:92, 04:97-04:99, 04:101-04:102, 04:104, 04:106, 04:109, 04:118-04:119, 04:124, 04:126-04:129, 04:133-04:136, 04:138, 04:141-04:143, 04:145, 04:148-04:149, 04:151-04:152, 04:154, 04:156, 04:158, 04:164, 04:167-04:168 and 04:174-04:177 alleles cannot be separated by this lot of HLA-C*04 high resolution kit.

5The following HLA-C*04 alleles can be distinguished by the different sizes of the HLA-specific PCR product:

| Alleles | Primer mix | Alleles | Primer mix |
| --- | --- | --- | --- |
| C*04:05, 04:112 | 6 | C*04:61, 04:96 | 34 |
| C*04:44, 04:47 | 20 | C*04:62, 04:76 | 35 |
| C*04:23, 04:38, 04:39 | 21 | C*04:65, 04:72 | 33 |
| C*04:25, 04:40, 04:41 | 23 | C*04:67, 04:93N | 46 |
| C*04:46, 04:50 | 27 | C*04:79, 04:113 | 39 |
| C*04:48, 04:75 | 28 | C*04:81, 04:88N | 43 |
| C*04:54, 04:105N | 41 | C*04:114, 04:146 | 47 |
| C*04:56, 04:64:01-04:64:02 | 40 | C*04:150, 04:155 | 52 |
| C*04:59Q, 04:78 | 32 |

The HLA-C*04 primer set cannot distinguish the following silent mutations: the C*04:01:01:01-04:01:09, 04:01:11-04:01:22, 04:01:24-04:01:27,  04:01:29-04:01:33 and 04:01:35-04:01:62 alleles, the C*04:01:10 and 04:01:23 alleles, the C*04:03:01-04:03:02 alleles, the C*04:04:01-04:04:02 alleles, the C*04:15:01 and 04:15:03 alleles, the C*04:42:01-04:42:02 alleles, the C*04:64:01-04:64:02 and the C*04:94:01-04:94:02 alleles.

‘w’, may be weakly amplified.

‘?’, nucleotide sequence information not available for the primer matching sequence.

Change in revision R01 compared to R00:

- The HLA-C*04:15:02 and 04:17 and the C*03:13:01, 03:35, 08:01:07, 08:02:07 and 08:33 alleles are amplified by primer mix 16. This has been corrected in the Specificity and Interpretation Tables.

Change in revision R02 compared to R01:

- Primer mixes 5 and 8 do not amplify the B*58:02 allele. This has been corrected in the Specificity and Interpretation Tables.

Change in revision R03 compared to R02:

- Primer mixes 47 and 50 do not amplify the C*04:161 allele. This has been corrected in the Interpretation and Specificity Tables.