Gel Documentation Form and Worksheet

HLA-C*05 (101.613-12/12u) 	Lot No: 02Y	Expiry Date: 2017-October-01

Name:__________________

Sample ID:______________

DNA Conc.(ng/ul):_________

Test  Date:_______________

Tested By:_______________

Review Date:_______________

Reviewed By:_______________

Interpretation:___________                           Failed lanes: ____________                               Comments:________________

Gel Picture

PHOTO DOCUMENT

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mix 12 may have tendencies of unspecific amplifications.

Primer mix 28 may have a tendency to giving rise to primer oligomer formation.

Primer mix 32 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 430 base pairs.

1HLA-C*05 alleles in bold lettering are listed as confirmed alleles on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.19.0, January 2015.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3Primer mix 4: Specific PCR fragment of 120 bp in the C*05:03 and in the C*07:52 alleles. Specific PCR fragment of 285 bp in the C*05:07N allele.

Primer mix 5: Specific PCR fragment of 225 bp in the C*05:04:01-05:04:02 and 05:103 and the C*06:129, 07:68, 07:260, 07:302, 08:09, 08:11, 08:59 and 08:113 and in the B*15:337 and B*18:83 alleles. Specific PCR fragment of 285 bp in the C*05:31 allele.

Primer mix 6: Specific PCR fragment of 255 bp in the C*05:05 and 05:99N and the C*03:251, 08:62 and 08:82 and in the A*02:425, A*02:519, A*29:10, A*68:69, B*14:32, B*15:337, B*18:83 and B*44:148 alleles. Specific PCR fragment of 280 bp in the C*05:16, 05:85 and 05:107 and the C*06:129, 07:364 and 08:12 and in the B*14:32 allele.

Primer mix 8: Specific PCR fragment of 85 bp in the C*05:06 allele. Specific PCR fragment of 210 bp in the C*05:55 allele.

Primer mix 9: Specific PCR fragment of 105 bp in the C*05:51Q allele. Specific PCR fragment of 175 bp in the C*05:48N allele.  Specific PCR fragment of 245 bp in the C*05:113N allele.

Primer mix 10: Specific PCR fragment of 95 bp in the C*05:08, 05:52 and 05:89 and the C*02:51, 08:29, 08:31 and 12:144 and in the B*15:33 and B*15:248 alleles. Specific PCR fragment of 250 bp in the C*05:30 and the C*03:247 and 06:125 alleles. Specific PCR fragment of 320 bp in the C*05:92N and the C*08:55N alleles.

Primer mix 11: Specific PCR fragment of 115 bp in the C*05:33 allele. Specific PCR fragment of 205 bp in the C*05:25 and 05:42 and the C*04:129, 06:05, 06:67, 07:101, 07:148, 07:161 and 08:28 and in the A*11:166 and A*80:01:01:01w-80:03w alleles.

Primer mix 12: Specific PCR fragment of 155 bp in the C*05:97 allele. Specific PCR fragment of 225 bp in the C*05:38 and the C*08:97 alleles. Specific PCR fragment of 285 bp in the C*05:10 and the C*03:251, 08:44, 08:61 and 08:82 and the B*44:148 alleles.

Primer mix 14: Specific PCR fragment of 120 bp in the C*05:12 and 05:15 alleles. Specific PCR fragment of 200 bp in the C*05:80 and 05:91N alleles.

Primer mix 15: Specific PCR fragment of 115 bp in the C*05:65 and the C*04:96 allele. Specific PCR fragment of 185 bp in the C*05:34 and in the C*02:93 and 06:13 alleles. Specific PCR fragment of 240 bp in the C*05:13 allele.

Primer mix 16: Specific PCR fragment of 195 bp in the C*05:14 allele. Specific PCR fragment of 470 bp in the C*05:93 and in the C*03:171, 03:211:01, 04:144, 06:73, 08:20, 08:40 and 12:109 alleles.

Primer mix 18: Specific PCR fragment of 160 bp in the C*05:87 allele. Specific PCR fragment of 245 bp in the C*05:19 and 05:99N alleles.

Primer mix 20: Specific PCR fragment of 260 bp in the C*05:26 allele. Specific PCR fragment of 390 bp in the C*05:21 allele.

Primer mix 23: Specific PCR fragment of 85 bp in the C*05:24 allele. Specific PCR fragment of 135 bp in the C*05:36 and in the C*07:148 alleles.

Primer mix 24: Specific PCR fragment of 185 bp in the C*05:43 and the C*08:37 alleles. Specific PCR fragment of 265 bp in the C*05:29:01-05:29:02 and in the C*08:13, 08:16:01, 08:25 and 08:94 alleles.

Primer mix 25: Specific PCR fragment of 105 bp in the C*05:40 allele. Specific PCR fragment of 205 bp in the C*05:35 and 05:80 alleles.

Primer mix 26: Specific PCR fragment of 115 bp in the C*05:27 and 05:39 and in the C*03:87:01-03:87:02 and 08:115 and in the B*15:33 and B*15:248 alleles. Specific PCR fragment of 185 bp in the C*05:28 and the C*06:64 alleles. Specific PCR fragment of 245 bp in the C*05:113N allele.

Primer mix 28: Specific PCR fragment of 115 bp in the C*05:65 and the C*04:96 alleles. Specific PCR fragment of 200 bp in the C*05:45 allele.

Primer mix 29: Specific PCR fragment of 140 bp in the C*05:41 and the C*04:78 alleles. Specific PCR fragment of 185 bp in the C*05:34 and in the C*02:93 and 06:13 alleles. Specific PCR fragment of 260 bp in the C*05:37 allele.

Primer mix 30: Specific PCR fragment of 70 bp in the C*05:53 allele. Specific PCR fragment of 175 bp in the C*05:32 allele.

4The C*05:18:01 and C*08:10 alleles, the C*05:29:01-05:29:02 and C*08:25 and 08:94 alleles and the C*05:92N and C*08:55N alleles will give rise to identical amplification patterns with the HLA-C*05 subtyping kit. These alleles can be distinguished by the HLA-C low resolution kit and/or the C*08 subtyping kit.

5The following HLA-C*05 alleles can be distinguished by the different sizes of the HLA-specific PCR product:

| Alleles | Primer mix | Alleles | Primer mix |
| --- | --- | --- | --- |
| C*05:03, 05:07N | 4 | C*05:24, 05:36 | 23 |
| C*05:06, 05:55 | 8 | C*05:25, 05:33 | 11 |
| C*05:08, 05:30 | 10 | C*05:28, 05:39 | 26 |
| C*05:14, 05:93 | 16 | C*05:32, 05:53 | 30 |
| C*05:15, 05:91N | 14 | C*05:35, 05:40 | 25 |
| C*05:21, 05:26 | 20 | C*05:37, 05:41 | 29 |

The HLA-C*05 subtyping kit cannot distinguish the following silent mutations: the C*05:01:01:01-05:01:31, the C*05:04:01-05:04:02, the C*05:09:01-05:09:03, the C*05:18:02-05:18:03, the C*05:22:01-05:22:02, the C*05:29:01-05:29:02 and the C*05:44:01-05:44:02 alleles.

‘w’, might be weakly amplified.

Changes in revision R01 compared to R00:

- Primer mixes 4 and 30 do not amplify the C*05:108 and the C*06:111, C*07:316 and C*07:359  alleles. This has been corrected in the Specificity and Interpretation Tables.