PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® HLA-C*07

Prod. No:	101.615-24/24u

101.615-04/04u

Lot No:    	03M

Expiry Date:    	2013-October-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 48 primer mixes used for HLA-C*07 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified HLA-C*07 alleles3 | Other amplified HLA-Class I alleles4 |
| --- | --- | --- | --- | --- |
| 110 | 245 bp, 425 bp | 800 bp | *07:01:01-07:33N, 07:35-07:99 |  |
| 2 | 290 bp | 800 bp | *07:01:01-07:01:04, 07:01:06-07:01:12, 07:05-07:07, 07:09, 07:14, 07:16, 07:18, 07:20-07:22, 07:24, 07:26-07:28, 07:30-07:31, 07:35-07:36, 07:40, 07:43-07:44, 07:52-07:53, 07:55N, 07:58-07:60, 07:65, 07:69-07:71, 07:73, 07:77-07:78, 07:81-07:83, 07:85-07:86, 07:89, 07:91-07:96, 07:98N | *03:04:19, 05:17, 12:03:10, B*35:08:04 |
| 35 | 110 bp | 800 bp | *07:01:01-07:01:12, 07:06, 07:09, 07:18-07:22, 07:24, 07:26, 07:28, 07:30, 07:35-07:36, 07:40, 07:44, 07:52-07:53, 07:55N, 07:57-07:59, 07:65, 07:69-07:71, 07:73, 07:77-07:78, 07:81-07:83, 07:85-07:86, 07:89, 07:91-07:96, 07:98N |  |
| 4 | 185 bp | 1070 bp | *07:02:01:01-07:02:14, 07:10, 07:13, 07:15, 07:19, 07:23, 07:25, 07:29, 07:32N-07:33N, 07:37-07:39, 07:42, 07:46-07:51, 07:54, 07:56:01-07:56:02, 07:61N-07:62, 07:64, 07:66-07:68, 07:72, 07:74-07:76, 07:79-07:80, 07:84, 07:87-07:88, 07:90, 07:97, 07:99 | B*40:60 |
| 55,11 | 70 bp, 145 bp, 245 bp | 800 bp | *07:03, 07:53, 07:57 |  |
| 65 | 100 bp | 1070 bp | *07:04:01-07:04:03, 07:11-07:12, 07:63, 07:68 |  |
| 7 | 190 bp | 1070 bp | *07:05 | B*27:36, B*48:21 |
| 812 | 285 bp, 505 bp | 1070 bp | *07:06, 07:45 | *04:40 |
| 96,13 | 215 bp, 240 bp, 425 bp | 800 bp | *07:07, 07:09, 07:29, 07:33N, 07:49, 07:76 | *03:04:19, 06:17, 18:01-18:03 |
| 10 | 185 bp | 1070 bp | *07:08 | *04:58 |
| 1114 | 275 bp, 305 bp | 800 bp | *07:10, 07:24, 07:43 | *03:04:19, B*35:08:04, B*37:04:02 |
| 125 | 130 bp | 1070 bp | *07:12, 07:41 | *05:01:01:01-05:01:13, 05:03-05:08, 05:10-05:16, 05:18-05:22, 05:24-05:42, 08:01:01-08:06, 08:08-08:14, 08:16-08:31, A*29:10 |
| 135,7 | 90 bp | 1070 bp | *07:11, 07:56:01-07:56:02, 07:60, 07:79 | *03:69, 04:27, 04:52 |
| 14 | 240 bp | 800 bp | *07:13, 07:28 | *05:17, B*37:04:02 |
| 155,15 | 80 bp, 290 bp, 460 bp | 800 bp | *07:14, 07:50, 07:89 |  |
| 16 | 535 bp | 1070 bp | *07:02:01:01-07:02:14, 07:04:01-07:05, 07:08, 07:10-07:11, 07:13-07:15, 07:23, 07:25, 07:27:01-07:27:02, 07:29, 07:31-07:33N, 07:37-07:39, 07:42-07:43, 07:45-07:51, 07:54, 07:56:01-07:56:02, 07:61N-07:64, 07:66-07:68, 07:72, 07:74-07:76, 07:79-07:80, 07:84, 07:87-07:88, 07:90, 07:97, 07:99 | *03:14, 04:58, 05:23, 08:07, 12:14:01-12:14:02, 14:17, 17:01:01:01-17:06, B*46:06 |
| 175,7 | 130 bp | 1070 bp | *07:04:01-07:04:03, 07:11, 07:63, 07:68 | *05:23, 08:07 |
| 185,16 | 120 bp, 230 bp | 800 bp | *07:15, 07:52 | *05:03, B*07:77, B*40:60 |
| 1917 | 220 bp, 355 bp, 465 bp | 800 bp | *07:07, 07:16, 07:26, 07:33N, 07:51, 07:80, 07:92, 07:96 |  |
| 2018 | 145 bp, 170 bp, 195 bp | 1070 bp | *07:03, 07:17, 07:44, 07:81 | *03:04:19, 05:17, 12:03:10, A*30:01:03, B*35:08:04, B*37:04:02 |
| 215,19 | 100 bp, 225 bp | 800 bp | *07:06, 07:18-07:19, 07:54 |  |
| 2220 | 205 bp, 250 bp | 800 bp | *07:20, 07:84, 07:96 | *03:15, 03:32, 03:45, 03:60, 04:03, 04:06, 04:16, 15:25 |
| 2321 | 155 bp, 185 bp | 1070 bp | *07:21, 07:23 |  |
| 245 | 130 bp | 1070 bp | *07:22, 07:25 | *12:03:10 |
| 255,22 | 105 bp, 150 bp | 1070 bp | *07:30, 07:46 | B*56:08w, B*56:14 |
| 265,23 | 100 bp, 250 bp, 300 bp | 1070 bp | *07:31-07:32N, 07:55N | *12:03:10 |
| 275 | 130 bp | 1070 bp | *07:37, 07:69 | B*35:08:04 |
| 2824 | 175 bp, 260 bp | 800 bp | *07:35, 07:44, 07:47 |  |
| 29 | 255 bp | 1070 bp | *07:36, 07:42 |  |
| 305 | 60 bp | 1070 bp | *07:75 |  |
| 318 | 415 bp | 1070 bp | *07:39-07:40 |  |
| 3225 | 245 bp, 285 bp | 1070 bp | *07:07, 07:09, 07:49, 07:60, 07:76, 07:79 | *02:14, 03:39, 03:67, 04:42, 06:02:01:01-06:02:01:02, 06:02:03-06:02:05, 06:04-06:10, 06:12-06:37, 15:23, 18:01-18:03, B*08:06, B*08:17, B*08:39, B*08:47 |
| 3326 | 205 bp, 230 bp | 1070 bp | *07:66, 07:82 |  |
| 345,27 | 105 bp, 270 bp | 1070 bp | *07:67, 07:70 |  |
| 356,28 | 320 bp, 360 bp | 800 bp | *07:71-07:72, 07:80 |  |
| 366 | 265 bp | 1070 bp | *07:20, 07:64, 07:73, 07:92, 07:96 |  |
| 37 | 265 bp | 1070 bp | *07:38, 07:63, 07:68 | B*07:77, B*40:60 |
| 389,29 | 135 bp, 195 bp, 405 bp | 1070 bp | *07:58, 07:61N, 07:65 |  |
| 395,30 | 100 bp, 255 bp | 800 bp | *07:59, 07:74 |  |
| 4031 | 205 bp, 440 bp | 1070 bp | *07:07, 07:09, 07:62, 07:76, 07:78 | *06:17, 18:01-18:03 |
| 4132 | 140 bp, 255 bp | 1070 bp | *07:88, 07:91 |  |
| 42 | 205 bp | 1070 bp | *07:77, 07:93 |  |
| 435,33 | 105 bp, 140 bp | 1070 bp | *07:94, 07:99 |  |
| 446,34 | 165 bp, 195 bp | 1070 bp | *07:65, 07:77, 07:98N |  |
| 45 | 475 bp | 1070 bp | *07:95 | *04:34, B*73:01 |
| 4635 | 200 bp, 350 bp | 1070 bp | *07:83, 07:90 |  |
| 4736 | 200 bp, 455 bp | 1070 bp | *07:86, 07:97 |  |
| 4837 | 220 bp, 400 bp | 1070 bp | *07:85, 07:87 |  |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-C*07 SSP typings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective lengths of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most wells, or a band of 800 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-C*07 subtyping.

In addition, wells number 2, 3, 5, 9, 11, 14, 15, 18, 19, 21, 22, 28, 35 and 39 contain the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3For several HLA-C*07 alleles 1st, 5th, 6th and 7th exon nucleotide sequences are not available. In these instances it is not known whether some of the primers of the SSP sets are completely matched with the target sequences or not. We assume that unknown sequences in the 1st, 5th, 6th and 7th exons are conserved within allelic groups.

4Due to the sharing of sequence motifs between HLA-C alleles non-HLA-C*07 alleles will be amplified by primer mixes 2, 8 to 14, 16 to 18, 20, 22, 24, 26, 32, 40 and 45. In addition, the A*29:10 allele will be amplified by primer mix 12, the A*30:01:03 allele by primer mix 20, the B*07:77 allele will be amplified by primer mixes 18 and 37, the B*08:06, B*08:17, B*08:39 and B*08:47 alleles by primer mix 32, the B*27:36 and B*48:21 alleles by primer mix 7, the B*35:08:04 allele by primer mixes 2, 11, 20 and 27, the B*37:04:02 allele by primer mixes 11, 14 and 20, the B*40:60 allele by primer mixes 4, 18 and 37, the B*46:06 allele by primer mix 16, the B*56:08 and B*56:14 alleles by primer mix 25 and the B*73:01 allele will be amplified by primer mix 45.

5Short specific PCR fragments are less intense and not as sharp as longer specific bands.

6Primer mixes 9, 35, 36 and 44 have a tendency of primer oligomer formation.

7Primer mixes 13 and 17 have a tendency of giving rise to nonspecific amplifications.

8Primer mix 31 may give a lower yield of specific PCR product than the other C*07 primer mixes.

9Primer mix 38 may generate a false band of about 900 base pairs. This band should be disregarded when interpreting HLA-C*07 SSP typings.

10Primer mix 1: Specific PCR fragment of 245 bp in the C*07:07, 07:09, 07:49 and 07:76 alleles. Specific PCR fragment of 425 bp in the C*07:41 allele. Specific PCR fragment of 245 and 425 bp in the C*07:01:01-07:06, 07:08, 07:10-07:33N, *07:35-07:40, 07:42-07:48, 07:50-07:75 and 07:77-07:99 alleles.

11Primer mix 5: Specific PCR fragment of 70 bp in the C*07:53 allele. Specific PCR fragment of 145 bp in the C*07:57 allele. Specific PCR fragment of 245 bp in the C*07:03 allele.

12Primer mix 8: Specific PCR fragment of 285 bp in the C*07:45 and the C*04:40 alleles. Specific PCR fragment of 505 bp in the C*07:06 allele.

13Primer mix 9: Specific PCR fragment of 215 bp in the C*07:33N allele. Specific PCR fragment of 240 bp in the C*07:29 and the C*03:04:19 alleles. Specific PCR fragment of 425 bp in the C*07:07, 07:09, 07:49 and 07:76 and the C*06:17 and 18:01-18:03 alleles.

14Primer mix 11: Specific PCR fragment of 275 bp in the C*07:24 allele. Specific PCR fragment of 305 bp in the C*07:10 and 07:43 and the C*03:04:19 and in the B*35:08:04 and B*37:04:02 alleles.

15Primer mix 15: Specific PCR fragment of 80 bp in the C* 07:50 allele. Specific PCR fragment of 290 bp in the C*07:14 allele.  Specific PCR fragment of 460 bp in the C*07:89 allele.

16Primer mix 18: Specific PCR fragment of 120 bp in the C*07:52 and the C*05:03 alleles. Specific PCR fragment of 230 bp in the C*07:15 and in the B*07:77 and B*40:60 alleles.

17Primer mix 19: Specific PCR fragment of 220 bp in the C*07:26, 07:33N, 07:92 and 07:96 alleles. Specific PCR fragment of 355 bp in the C*07:80 allele. Specific PCR fragment of 465 bp in the C*07:07, 07:16 and 07:51 alleles.

18Primer mix 20: Specific PCR fragment of 145 bp in the C*07:03 and 07:17 and the C*03:04:19, 05:17 and 12:03:10 and in the A*30:01:03 alleles. Specific PCR fragment of 170 bp in the 07:44 allele. Specific PCR fragment of 195 bp in the C*07:81 allele. Specific PCR fragment of 145 bp and 195 bp in the B*35:08:04 and B*37:04:02 alleles.

19Primer mix 21:	 Specific PCR fragment of 100 bp in the C*07:54 allele. Specific PCR fragment of 225 bp in the C*07:06, 07:18 and 07:19 alleles.

20Primer mix 22:	 Specific PCR fragment of 205 bp in the C*07:84 allele. Specific PCR fragment of 250 bp in the C*07:20 and 07:96 and the C*03:15, 03:32, 03:45, 03:60, 04:03, 04:06, 04:16 and 15:25 alleles.

21Primer mix 23:	 Specific PCR fragment of 155 bp in the C*07:23 allele. Specific PCR fragment of 185 bp in the C*07:21 allele.

22Primer mix 25: Specific PCR fragment of 105 bp in the C*07:46 and B*56:06w and B*56:14 alleles. Specific PCR fragment of 150 bp in the C*07:30 allele.

23Primer mix 26: Specific PCR fragment of 100 bp in the C*07:32N allele. Specific PCR fragment of 250 bp in the C*07:55N allele. Specific PCR fragment of 300 bp in the C*07:31 and the C*12:03:10 alleles.

24Primer mix 28: Specific PCR fragment of 175 bp in the C*07:44 allele. Specific PCR fragment of 260 bp in the C*07:35 and 07:47 alleles.

25Primer mix 32: Specific PCR fragment of 245 bp in the C*07:07, 07:09, 07:49 and 07:76 and the C*02:14, 04:42, 06:02:01:01-06:02:01:02, 06:02:03-06:02:05, 06:04-06:10, 06:12-06:37, 15:23 and 18:01-18:03 and in the B*08:17 and B*08:47 alleles. Specific PCR fragment of 285 bp in the C*07:60 and 07:79 and the C*03:39 and 03:67 and the B*08:39 alleles.

26Primer mix 33: Specific PCR fragment of 205 bp in the C*07:66 allele. Specific PCR fragment of 230 bp in the C*07:82 allele.

27Primer mix 34: Specific PCR fragment of 105 bp in the C*07:67 allele. Specific PCR fragment of 270 bp in the C*07:70 allele.

28Primer mix 35: Specific PCR fragment of 320 bp in the C*07:72 allele. Specific PCR fragment of 360 bp in the C*07:71 and 07:80 alleles.

29Primer mix 38: Specific PCR fragment of 135 bp in the C*07:58 allele. Specific PCR fragment of 195 bp in the C*07:65 allele. Specific PCR fragment of 405 bp in the C*07:61N allele.

30Primer mix 39: Specific PCR fragment of 100 bp in the C*07:74 allele. Specific PCR fragment of 255 bp in the C*07:59 allele.

31Primer mix 40: Specific PCR fragment of 205 bp in the C*07:62 allele. Specific PCR fragment of 440 bp in the C*07:07, 07:09, 07:76 and 07:78 and the C*06:17 and 18:01-18:03 alleles.

32Primer mix 41: Specific PCR fragment of 140 bp in the C*07:91 allele. Specific PCR fragment of 255 bp in the C*07:88 allele.

33Primer mix 43: Specific PCR fragment of 105 bp in the C*07:94 allele. Specific PCR fragment of 140 bp in the C*07:99 allele.

34Primer mix 44: Specific PCR fragment of 165 bp in the C*07:98N allele. Specific PCR fragment of 195 bp in the C*07:65 and 07:77 alleles.

35Primer mix 46: Specific PCR fragment of 200 bp in the C*07:83 allele. Specific PCR fragment of 350 bp in the C*07:90 allele.

36Primer mix 47: Specific PCR fragment of 200 bp in the C*07:86 allele. Specific PCR fragment of 455 bp in the C*07:97 allele.

37Primer mix 48: Specific PCR fragment of 220 bp in the C*07:87 allele. Specific PCR fragment of 400 bp in the C*07:85 allele.

‘w’, may be weakly amplified.

1The Cw*0734 allele has been renamed to C*07:27:02.

2Primer mix 1: Specific PCR fragment of 245 bp in the C*07:07, 07:09, 07:49 and 07:76 alleles. Specific PCR fragment of 425 bp in the C*07:41 allele. Specific PCR fragment of 245 and 425 bp in the C*07:01:01-07:06, 07:08, 07:10-07:33N, *07:35-07:40, 07:42-07:48, 07:50-07:75 and 07:77-07:99 alleles.

Primer mix 5: Specific PCR fragment of 70 bp in the C*07:53 allele. Specific PCR fragment of 145 bp in the C*07:57 allele. Specific PCR fragment of 245 bp in the C*07:03 allele.

Primer mix 8: Specific PCR fragment of 285 bp in the C*07:45 and the C*04:40 alleles. Specific PCR fragment of 505 bp in the C*07:06 allele.

Primer mix 9: Specific PCR fragment of 215 bp in the C*07:33N allele. Specific PCR fragment of 240 bp in the C*07:29 and the C*03:04:19 alleles. Specific PCR fragment of 425 bp in the C*07:07, 07:09, 07:49 and 07:76 and the C*06:17 and 18:01-18:03 alleles.

Primer mix 11: Specific PCR fragment of 275 bp in the C*07:24 allele. Specific PCR fragment of 305 bp in the C*07:10 and 07:43 and the C*03:04:19 and in the B*35:08:04 and B*37:04:02 alleles.

Primer mix 15: Specific PCR fragment of 80 bp in the C* 07:50 allele. Specific PCR fragment of 290 bp in the C*07:14 allele.  Specific PCR fragment of 460 bp in the C*07:89 allele.

Primer mix 18: Specific PCR fragment of 120 bp in the C*07:52 and the C*05:03 alleles. Specific PCR fragment of 230 bp in the C*07:15 and in the B*07:77 and B*40:60 alleles.

Primer mix 19: Specific PCR fragment of 220 bp in the C*07:26, 07:33N, 07:92 and 07:96 alleles. Specific PCR fragment of 355 bp in the C*07:80 allele. Specific PCR fragment of 465 bp in the C*07:07, 07:16 and 07:51 alleles.

Primer mix 20: Specific PCR fragment of 145 bp in the C*07:03 and 07:17 and the C*03:04:19, 05:17 and 12:03:10 and in the A*30:01:03 alleles. Specific PCR fragment of 170 bp in the 07:44 allele. Specific PCR fragment of 195 bp in the C*07:81 allele. Specific PCR fragment of 145 bp and 195 bp in the B*35:08:04 and B*37:04:02 alleles.

Primer mix 21: Specific PCR fragment of 100 bp in the C*07:54 allele. Specific PCR fragment of 225 bp in the C*07:06, 07:18 and 07:19 alleles.

Primer mix 22: Specific PCR fragment of 205 bp in the C*07:84 allele. Specific PCR fragment of 250 bp in the C*07:20 and 07:96 and the C*03:15, 03:32, 03:45, 03:60, 04:03, 04:06, 04:16 and 15:25 alleles.

Primer mix 23: Specific PCR fragment of 155 bp in the C*07:23 allele. Specific PCR fragment of 185 bp in the C*07:21 allele.

Primer mix 25: Specific PCR fragment of 105 bp in the C*07:46 and B*56:06w and B*56:14 alleles. Specific PCR fragment of 150 bp in the C*07:30 allele.

Primer mix 26: Specific PCR fragment of 100 bp in the C*07:32N allele. Specific PCR fragment of 250 bp in the C*07:55N allele. Specific PCR fragment of 300 bp in the C*07:31 and the C*12:03:10 alleles.

Primer mix 28: Specific PCR fragment of 175 bp in the C*07:44 allele. Specific PCR fragment of 260 bp in the C*07:35 and 07:47 alleles.

Primer mix 32: Specific PCR fragment of 245 bp in the C*07:07, 07:09, 07:49 and 07:76 and the C*02:14, 04:42, 06:02:01:01-06:02:01:02, 06:02:03-06:02:05, 06:04-06:10, 06:12-06:37, 15:23 and 18:01-18:03 and in the B*08:17 and B*08:47 alleles. Specific PCR fragment of 285 bp in the C*07:60 and 07:79 and the C*03:39 and 03:67 and the B*08:39 alleles.

Primer mix 33: Specific PCR fragment of 205 bp in the C*07:66 allele. Specific PCR fragment of 230 bp in the C*07:82 allele.

Primer mix 34: Specific PCR fragment of 105 bp in the C*07:67 allele. Specific PCR fragment of 270 bp in the C*07:70 allele.

Primer mix 35: Specific PCR fragment of 320 bp in the C*07:72 allele. Specific PCR fragment of 360 bp in the C*07:71 and 07:80 alleles.

Primer mix 38: Specific PCR fragment of 135 bp in the C*07:58 allele. Specific PCR fragment of 195 bp in the C*07:65 allele. Specific PCR fragment of 405 bp in the C*07:61N allele.

Primer mix 39: Specific PCR fragment of 100 bp in the C*07:74 allele. Specific PCR fragment of 255 bp in the C*07:59 allele.

Primer mix 40: Specific PCR fragment of 205 bp in the C*07:62 allele. Specific PCR fragment of 440 bp in the C*07:07, 07:09, 07:76 and 07:78 and the C*06:17 and 18:01-18:03 alleles.

Primer mix 41: Specific PCR fragment of 140 bp in the C*07:91 allele. Specific PCR fragment of 255 bp in the C*07:88 allele.

Primer mix 43: Specific PCR fragment of 105 bp in the C*07:94 allele. Specific PCR fragment of 140 bp in the C*07:99 allele.

Primer mix 44: Specific PCR fragment of 165 bp in the C*07:98N allele. Specific PCR fragment of 195 bp in the C*07:65 and 07:77 alleles.

Primer mix 46: Specific PCR fragment of 200 bp in the C*07:83 allele. Specific PCR fragment of 350 bp in the C*07:90 allele.

Primer mix 47: Specific PCR fragment of 200 bp in the C*07:86 allele. Specific PCR fragment of 455 bp in the C*07:97 allele.

Primer mix 48: Specific PCR fragment of 220 bp in the C*07:87 allele. Specific PCR fragment of 400 bp in the C*07:85 allele.

3The HLA-C*07 subtyping kit cannot distinguish the C*07:02:01:01-07:02:14 and C*07:48 alleles.

‘w’, may be weakly amplified.

Change in revision R01 compared to R00:

- Primer mix 20 does not amplify the C*07:48 allele. This has been corrected in the specificity and amplification tables. Thus, this lot of the HLA-C*07 subtyping kit cannot distinguish the C*07:02:01:01-07:02:14 and C*07:48 alleles.