Gel Documentation Form and Worksheet

HLA-C*07 (101.615-24/04, -24u/04u) 	Lot No: 03V	Expiry Date: 2016-May-01

Name:__________________

Sample ID:______________

DNA Conc.(ng/ul):_________

Test  Date:_______________

Tested By:_______________

Review Date:_______________

Reviewed By:_______________

Gel Picture

PHOTO DOCUMENT

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

Interpretation:_______________

Failed lanes:    _________

Comments:    ___________________

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

Primer mixes 9, 19, 35, 36, 38, 43, 44 and 59 have a tendency to giving rise to primer oligomer formation.

Primer mixes 13, 17, 20, 23, 24, 30, 37, 40, 44, 45, 48 and 64 may have tendencies of unspecific amplifications, most pronounced in primer mix 24.

Primer mixes 1 and 16 may give rise to a lower yield of HLA-specific PCR product than the other C*07 primer mixes.

In primer mix 32 the positive control band may be weaker than for other HLA-C*07 primer mixes.

Primer mix 77 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by control primer pairs. PCR product sizes range from 75 to 200 base pairs. The PCR product generated by the control primer pair is 430 base pairs.

1HLA-C*07 alleles in bold lettering are listed as confirmed alleles on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.13.1, July 2013.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3Primer mix 1: Specific PCR fragment of 245 bp in the C*07:02:41, 07:07, 07:09, 07:49, 07:76, 07:115, 07:210, 07:223, 07:238, 07:247 and 07:315 alleles. Specific PCR fragment of 425 bp in the C*07:02:21, 07:41 and 07:166 alleles. Specific PCR fragment of 245 and 425 bp in the C*07:01:01:01-07:02:20, 07:02:22-07:02:40, 07:06, 07:08, 07:10-07:33N, *07:35-07:40, 07:42-07:48, 07:50-07:75, 07:77-07:114, 07:116-07:165, 07:167-07:209, 07:211-07:222,  07:224-07:237, 07:239-07:246, 07:248-07:294 and 07:296-07:314 alleles.

Primer mix 5: Specific PCR fragment of 70 bp in the C*07:53 and 07:216 alleles. Specific PCR fragment of 145 bp in the C*07:57 allele. Specific PCR fragment of 245 bp in the C*07:03 allele.

Primer mix 8: Specific PCR fragment of 285 bp in the C*07:45 and the C*02:57 and 04:40 and in the B*38:23 and B*39:73 alleles. Specific PCR fragment of 505 bp in the C*07:06 allele.

Primer mix 9: Specific PCR fragment of 215 bp in the C*07:33N allele. Specific PCR fragment of 425 bp in the C*07:07, 07:09, 07:49, 07:76, 07:210, 07:238, 07:247 and 07:315 and the C*06:17 and 18:01-18:07N alleles.

Primer mix 11: Specific PCR fragment of 275 bp in the C*07:24 and 07:218 alleles. Specific PCR fragment of 300 bp in the C*07:10, 07:43, 07:112, 07:184 and 07:196 and the C*03:04:19 and 03:113 and in the B*35:01:38, B*35:08:04 and B*37:04:02 alleles.

Primer mix 15: Specific PCR fragment of 80 bp in the C* 07:50 allele. Specific PCR fragment of 290 bp in the C*07:14 allele.  Specific PCR fragment of 460 bp in the C*07:89 allele.

Primer mix 18: Specific PCR fragment of 120 bp in the C*07:52 and 07:243 and the C*05:03 alleles. Specific PCR fragment of 230 bp in the C*07:15 and in the B*07:77 and B*07:193 and the B*40:60 and B*40:245 alleles.

Primer mix 19: Specific PCR fragment of 220 bp in the C*07:26, 07:92, 07:96:01-07:96:02 and 07:314 alleles. Specific PCR fragment of 465 bp in the C*07:07, 07:16, 07:51, 07:177 and 07:181 alleles.

Primer mix 20: Specific PCR fragment of 145 bp in the C*07:03, 07:17:01-07:17:02, 07:140-07:141:02, 07:304 and 07:307 and the C*03:04:19, 03:113, 04:01:43, 05:01:17, 05:09:02, 05:17, 06:02:23, 08:15:02, 08:72:02,  and 12:03:10 and in the A*30:01:03 and A*33:03:09 alleles. Specific PCR fragment of 170 bp in the 07:44 and 07:48 alleles. Specific PCR fragment of 195 bp in the C*07:81 and 07:168  alleles. Specific PCR fragment of 145 bp and 195 bp in the B*35:01:38, 35:08:04 and B*37:04:02 alleles.

Primer mix 21: Specific PCR fragment of 100 bp in the C*07:54 allele. Specific PCR fragment of 225 bp in the C*07:06, 07:18 and 07:19 alleles.

Primer mix 22: Specific PCR fragment of 205 bp in the C*07:84 allele. Specific PCR fragment of 250 bp in the C*07:20, 07:96:01-07:96:02, 07:127 and 07:263 and the C*02:49, 03:15, 03:32, 03:45, 03:60, 03:136, 04:03:01-04:03:02, 04:06, 04:16, 04:80, 04:147, 06:70, 15:25, 15:62 and in the A*01:110 and A*11:135 alleles.

Primer mix 23: Specific PCR fragment of 155 bp in the C*07:23 allele. Specific PCR fragment of 185 bp in the C*07:21 allele.

Primer mix 25: Specific PCR fragment of 105 bp in the C*07:46 and in the B*56:08w and B*56:14 alleles. Specific PCR fragment of 150 bp in the C*07:30 allele.

Primer mix 26: Specific PCR fragment of 100 bp in the C*07:32N allele. Specific PCR fragment of 250 bp in the C*07:55N allele. Specific PCR fragment of 300 bp in the C*07:31:01-07:31:02 and 07:177 and the C*06:02:23 and 12:03:10 alleles.

Primer mix 28: Specific PCR fragment of 175 bp in the C*07:44 allele. Specific PCR fragment of 260 bp in the C*07:35 and 07:47 alleles.

Primer mix 30: Specific PCR fragment of 60 bp in the C*07:75 and 07:254 alleles. Specific PCR fragment of 260 bp in the C*07:227N and the C*04:66 alleles.

Primer mix 32: Specific PCR fragment of 245 bp in the C*07:07, 07:09, 07:49, 07:76, 07:210, 07:238, 07:247 and 07:315 and the C*02:14, 04:42, 05:43, 06:02:01:01-06:02:01:02, 06:02:03-06:02:05, 06:02:07-06:02:11, 06:02:13-06:02:27, 06:04-06:10, 06:12-06:51, 06:53:01-06:75, 06:78-06:102, 15:23, 15:63 and 18:01-18:07N and in the B*08:17 and B*08:47alleles. Specific PCR fragment of 285 bp in the C*07:60 and 07:79 and the C*03:39 and 03:67 and the B*08:39 alleles.

Primer mix 33: Specific PCR fragment of 205 bp in the C*07:66 allele. Specific PCR fragment of 230 bp in the C*07:82 allele.

Primer mix 34: Specific PCR fragment of 105 bp in the C*07:67 allele. Specific PCR fragment of 270 bp in the C*07:70 allele.

Primer mix 35: Specific PCR fragment of 320 bp in the C*07:72 allele. Specific PCR fragment of 360 bp in the C*07:71 allele.

Primer mix 36: Specific PCR fragment of 185 bp in the C*07:107 and 07:224 and in the C*06:22 and 16:31 and the B*15:193 alleles. Specific PCR fragment of 265 bp in the C*07:20, 07:64, 07:73, 07:92, 07:96:01-07:96:02 and 07:172 and the C*18:04 alleles.

Primer mix 37: Specific PCR fragment of 250 bp in the C*07:63 and 07:121Q alleles. Specific PCR fragment of 275 bp in the C*07:38:01-07:38:02, 07:68, 07:260 and 07:302 and the C*08:38 and the B*07:77, B*07:193, B*40:60 and B*40:245 alleles.

Primer mix 38: Specific PCR fragment of 135 bp in the C*07:58 allele. Specific PCR fragment of 195 bp in the C*07:65 allele.

Primer mix 39: Specific PCR fragment of 100 bp in the C*07:74 allele. Specific PCR fragment of 195 bp in the C*07:308 allele.  Specific PCR fragment of 255 bp in the C*07:59 allele.

Primer mix 40: Specific PCR fragment of 120 bp in the C*07:235Q allele.  Specific PCR fragment of 205 bp in the C*07:62 allele. Specific PCR fragment of 440 bp in the C*07:07, 07:09, 07:76, 07:78 and 07:315 and in the C*06:17 and 18:01-18:07N alleles.

Primer mix 41: Specific PCR fragment of 140 bp in the C*07:91 allele. Specific PCR fragment of 255 bp in the C*07:88 allele.

Primer mix 42: Specific PCR fragment of 100 bp in the C*07:15 and 07:207 alleles. Specific PCR fragment of 205 bp in the C*07:77 and 07:93 alleles. Specific PCR fragment of 250 bp in the C*07:123, 07:173 and 07:294 and the B*35:01:38 and B*35:08:04 alleles.

Primer mix 43: Specific PCR fragment of 105 bp in the C*07:94 and 07:183 alleles. Specific PCR fragment of 140 bp in the C*07:99 allele. Specific PCR fragment of 175 bp in the C*07:80 allele.

Primer mix 44: Specific PCR fragment of 165 bp in the C*07:98N allele. Specific PCR fragment of 195 bp in the C*07:65 and 07:77 alleles.

Primer mix 45: Specific PCR fragment of 240 bp in the C*07:245 allele. Specific PCR fragment of 475 bp in the C*07:95 and the C*04:34 and 04:122 and in the B*35:178 and B*73:01-73:02 alleles.

Primer mix 46: Specific PCR fragment of 200 bp in the C*07:83 allele. Specific PCR fragment of 315 bp in the C*07:202 allele. Specific PCR fragment of 350 bp in the C*07:90 allele. Specific PCR fragment of 420 bp in the C*07:152N allele.

Primer mix 47: Specific PCR fragment of 200 bp in the C*07:86 allele. Specific PCR fragment of 420 bp in the C*07:152N allele.Specific PCR fragment of 455 bp in the C*07:97 allele.

Primer mix 48: Specific PCR fragment of 155 bp in the C*07:122 allele. Specific PCR fragment of 220 bp in the C*07:87 allele. Specific PCR fragment of 400 bp in the C*07:85 allele.

Primer mix 49: Specific PCR fragment of 155 bp in the C*07:101, 07:148 and 07:161 and the C*04:129, 05:42, 05:46, 06:67, 08:05, 08:21 and 17:05 alleles. Specific PCR fragment of 220 bp in the C*07:100 allele.

Primer mix 50: Specific PCR fragment of 105 bp in the C*07:104N and in the B*15:272N alleles. Specific PCR fragment of 170 bp in the C*07:119 and in the B*07:87 and B*27:55 alleles. Specific PCR fragment of 425 bp in the C*07:124 and 07:155 alleles.

Primer mix 51: Specific PCR fragment of 65 bp in the C*07:126 allele. Specific PCR fragment of 235 bp in the C*07:121Q allele. Specific PCR fragment of 285 bp in the C*07:117 allele.

Primer mix 52: Specific PCR fragment of 65 bp in the C*07:203 allele. Specific PCR fragment of 205 bp in the C*07:118 allele. Specific PCR fragment of 280 bp in the C*07:106 allele.

Primer mix 53: Specific PCR fragment of 110 bp in the C*07:140, 07:151 and 07:154 and in the B*37:04:02 alleles. Specific PCR fragment of 155 bp in the C*07:122 allele.

Primer mix 54: Specific PCR fragment of 120 bp in the C*07:153 and in the B*51:142 alleles. Specific PCR fragment of 235 bp in the C*07:64 allele. Specific PCR fragment of 340 bp in the C*07:125 allele. Specific PCR fragment of 370 bp in the C*07:129 allele.

Primer mix 55: Specific PCR fragment of 150 bp in the C*07:179 and the C*06:07 alleles. Specific PCR fragment of 210 bp in the C*07:252 allele. Specific PCR fragment of 290 bp in the C*07:132 allele.

Primer mix 56: Specific PCR fragment of 210 bp in the C*07:150Q allele. Specific PCR fragment of 255 bp in the C*07:123, 07:143, 07:173 and 07:294 and the B*35:01:38 and B*35:08:04 alleles.

Primer mix 57: Specific PCR fragment of 70 bp in the C*07:01:25 allele. Specific PCR fragment of 175 bp in the C*07:130 and the C*05:01:17, 05:09:02, 05:17 and 08:15:02 alleles.

Primer mix 58: Specific PCR fragment of 80 bp in the C*07:137:01-07:137:02, 07:186, 07:212 and 07:265 and the C*03:162 and in the B*38:43 alleles. Specific PCR fragment of 450 bp in the C*07:164N allele.

Primer mix 60: Specific PCR fragment of 130 bp in the C*07:165, 07:241, 07:265 07:304 and 07:307 and the C*03:04:19, 03:113, 04:01:43, 05:01:17, 05:09:02, 06:02:23, 08:15:02 and 12:03:10 alleles. Specific PCR fragment of 305 bp in the C*07:147 allele. Specific PCR fragment of 370 bp in the C*07:180 allele.

Primer mix 62: Specific PCR fragment of 200 bp in the C*07:84 and 07:205 alleles. Specific PCR fragment of 280 bp in the C *07:106 allele.

Primer mix 63: Specific PCR fragment of 275 bp in the C *07:24 and 07:218 alleles. Specific PCR fragment of 320 bp in the C*07:198N and 07:202 alleles.

Primer mix 64: Specific PCR fragment of 135 bp in the C*07:69 allele. Specific PCR fragment of 225 bp in the C *07:192 allele.

Primer mix 74: Specific PCR fragment of 215 bp in the C*07:33N allele. Specific PCR fragment of 355 bp in the C*07:80 allele.

‘w’, may be weakly amplified.

4The following HLA-C*07 alleles can be distinguished by the different sizes of the specific PCR product:

| Alleles | Primer mix | Alleles | Primer mix |
| --- | --- | --- | --- |
| C*07:51, 07:314 | 19 | C*07:118, 07:203 | 52 |
| C*07:74, 07:308 | 39 | C*07:129, 07:153 | 54 |
| C*07:78, 02:235Q | 40 | C*07:132, 07:179 | 55 |
| C*07:93, 07:207 | 42 | C*07:147, 07:241 | 60 |
| C*07:100, 07:161 | 49 | C*07:164N, 07:212 | 58 |
| C*07:104N, 07:119, 07:124 | 50 | C*07:165, 07:180 | 60 |
| C*07:117, 07:126 | 51 | C*07:227N, 07:254 | 30 |

The HLA-C*07 subtyping kit cannot distinguish the following silent mutations: the C*07:01:01:01-07:01:04, C*07:01:06-07:01:24 and 07:01:26-07:01:34 alleles, the C*07:02:01:01-07:02:43 alleles, the C*07:04:01-07:04:08, the C*07:17:01-07:17:02, the C*07:27:01-07:27:02, the C*07:31:01-07:31:02, the C*07:38:01-07:38:02, the C*07:56:01-07:56:02, the C*07:96:01-07:96:02, the C*07:137:01-07:137:02 or the C*07:141:01-07:141:02 alleles.

5This lot of the HLA-C*07 subtyping kit cannot distinguish the C*07:198N and the C*07:218 alleles.

Change in revision R01 compared to R00:

- The HLA-C*07:24 and 07:218 and the C*03:125, 03:176, 05:81, 06:87, 12:45 and 16:13 alleles are not amplified by primer mix 61. This has been corrected in the Specificity and Interpretation Tables. Thus, this lot of the HLA-C*07 subtyping kit cannot distinguish the C*07:198N and the C*07:218 alleles.

Change in revision R02 compared to R01:

- Primer mix 16 does not amplify the C*07:25 allele. This has been corrected in the Specificity and Interpretation Tables.