Gel Documentation Form and Worksheet

HLA-C*07 (101.615-24/04, -24u/04u) 	Lot No: 04X	Expiry Date: 2017-March-01

Name:__________________

Sample ID:______________

DNA Conc.(ng/ul):_________

Test  Date:_______________

Tested By:_______________

Review Date:_______________

Reviewed By:_______________	_

Interpretation:_______________	                    Failed lanes: ____________                           Comments:________________

Gel Picture

PHOTO DOCUMENT

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 38, 43, 46 and 47 have a tendency to giving rise to primer oligomer formation.

Primer mixes 13, 17, 20, 23, 24, 30, 34, 37, 40, 45, 65, 66, 70, 74, 77 and 83 may have tendencies of unspecific amplifications, most pronounced in primer mix 77.

Primer mix 16 may give rise to a lower yield of HLA-specific PCR product than the other C*07 primer mixes.

In primer mix 3 the positive control band may be weaker than for other HLA-C*07 primer mixes.

Primer mix 87 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by control primer pairs. PCR product sizes range from 75 to 200 base pairs. The PCR product generated by the control primer pair is 430 base pairs.

1HLA-C*07 alleles in bold lettering are listed as confirmed alleles on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.16.0, April 2014.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3Primer mix 1: Specific PCR fragment of 245 bp in the C*07:01:41, 07:02:41, 07:07, 07:09, 07:49, 07:76:01-07:76:02, 07:115, 07:210, 07:223, 07:238, 07:247, 07:315, 07:328, 07:357 and 07:361 alleles. Specific PCR fragment of 425 bp in the C*07:02:21, 07:41, 07:166 and 07:348 alleles. Specific PCR fragment of 245 and 425 bp in the C*07:01:01:01-07:02:20, 07:02:22-07:02:40, 07:02:42-07:06, 07:08, 07:10-07:33N, 07:35-07:40, 07:42-07:48, 07:50-07:75, 07:77-07:114, 07:116-07:165, 07:167-07:209, 07:211-07:222,  07:224-07:237, 07:239-07:246, 07:248-07:294, 07:296-07:314, 07:316-07:327, 07:329N-07:347N, 07:349-07:356, 07:358-07:360 and 07:362-07:374 alleles.

Primer mix 5: Specific PCR fragment of 70 bp in the C*07:53 and 07:216 alleles. Specific PCR fragment of 145 bp in the C*07:57 allele. Specific PCR fragment of 245 bp in the C*07:03 allele.

Primer mix 8: Specific PCR fragment of 285 bp in the C*07:45 and the C*02:57, 04:40 and 06:123 and in the B*38:23 and B*39:73 alleles. Specific PCR fragment of 505 bp in the C*07:06 and 07:343 alleles.

Primer mix 9: Specific PCR fragment of 215 bp in the C*07:33N allele. Specific PCR fragment of 425 bp in the C*07:07, 07:09, 07:49, 07:76:01-07:76:02, 07:210, 07:238, 07:247, 07:315 and 07:328 and in the C*06:17 and 18:01-18:07Nalleles.

Primer mix 11: Specific PCR fragment of 280 bp in the C*07:24, 07:112 and 07:218 alleles. Specific PCR fragment of 305 bp in the C*07:10, 07:43, 07:184, 07:196 and 07:367 and the C*03:04:19 and 03:113 and in the B*35:01:38, 35:08:04, B*37:04:02 and B*57:01:18alleles.

Primer mix 15: Specific PCR fragment of 80 bp in the C*07:50 allele. Specific PCR fragment of 290 bp in the C*07:14 allele.  Specific PCR fragment of 460 bp in the C*07:89 allele.

Primer mix 18: Specific PCR fragment of 120 bp in the C*07:52 and 07:243 and the C*05:03 alleles. Specific PCR fragment of 230 bp in the C*07:15 and 07:337 and the C*02:77 and in the B*07:77 and B*07:193 and the B*40:60 and B*40:245 alleles.

Primer mix 19: Specific PCR fragment of 220 bp in the C*07:26, 07:92, 07:96:01-07:96:02, 07:314, 07:317 and 07:351 alleles. Specific PCR fragment of 465 bp in the C*07:07, 07:16, 07:51, 07:177, 07:181, 07:328 and 07:367 alleles.

Primer mix 20: Specific PCR fragment of 145 bp in the C*07:03, 07:17:01-07:17:02, 07:140-07:141:02, 07:304 and 07:307 and the C*03:04:19, 03:113, 04:01:43, 05:01:17, 05:09:02, 05:17, 06:02:23, 08:15:02, 08:72:02, 12:03:10 and 16:02:11 and in the A*30:01:03, A*33:03:09, B*07:33:03 and B*08:01:25 alleles. Specific PCR fragment of 170 bp in the 07:44 and 07:48 alleles. Specific PCR fragment of 195 bp in the C*07:81 and 07:168 alleles. Specific PCR fragment of 145 bp and 195 bp in the B*35:01:38, 35:08:04, B*37:04:02 and B*57:01:18 alleles.

Primer mix 21: Specific PCR fragment of 100 bp in the C*07:54 allele. Specific PCR fragment of 225 bp in the C*07:06 and 07:18-07:19 alleles.

Primer mix 22: Specific PCR fragment of 205 bp in the C*07:84 allele. Specific PCR fragment of 250 bp in the C*07:20, 07:96:01-07:96:02, 07:127 and 07:263 and the C*02:49, 02:75, 03:15, 03:32, 03:45, 03:60, 03:136, 04:03:01-04:03:02, 04:06, 04:16, 04:80, 04:147, 04:160, 06:70, 15:25 and 15:62 and in the A*01:110, A*11:135 alleles.

Primer mix 23: Specific PCR fragment of 155 bp in the C*07:23 allele. Specific PCR fragment of 185 bp in the C*07:21 allele.

Primer mix 25: Specific PCR fragment of 105 bp in the C*07:46 and 07:356 and in the B*56:08w and B*56:14 alleles. Specific PCR fragment of 150 bp in the C*07:30 allele.

Primer mix 26: Specific PCR fragment of 100 bp in the C*07:32N allele. Specific PCR fragment of 250 bp in the C*07:55N allele. Specific PCR fragment of 300 bp in the C*07:31:01-07:31:02 and 07:177 and the C*06:02:23, 12:03:10 and 16:02:11 alleles.

Primer mix 28: Specific PCR fragment of 175 bp in the C*07:44 allele. Specific PCR fragment of 260 bp in the C*07:35, 07:47 and 07:352 alleles.

Primer mix 32: Specific PCR fragment of 115 bp in the C*07:134 allele. Specific PCR fragment of 155 bp in the C*07:189 allele. Specific PCR fragment of 245 bp in the C*07:07, 07:09, 07:49, 07:76:01-07:76:02, 07:210, 07:238, 07:247, 07:315 and 07:328 and the C*02:14, 04:42, 05:43, 06:02:01:01-06:02:01:03, 06:02:03-06:02:05, 06:02:07-06:02:11, 06:02:13-06:02:31, 06:02:33-06:02:38, 06:04-06:10, 06:12-06:28 06:30-06:51, 06:53:01-06:75, 06:78-06:121, 06:123, 06:125-06:131, 15:23, 15:63, 17:20, 18:01-18:07N, B*08:17, 08:47 alleles. Specific PCR fragment of 285 bp in the C*07:60 and 07:79 and the C*03:39 and 03:67 and in the B*08:39 alleles. Specific PCR fragment of 115 bp and 245 bp in the C*06:29 allele.

Primer mix 33: Specific PCR fragment of 205 bp in the C*07:66 allele. Specific PCR fragment of 230 bp in the C*07:82 allele.

Primer mix 34: Specific PCR fragment of 105 bp in the C*07:67 allele. Specific PCR fragment of 270 bp in the C*07:70 allele.

Primer mix 35: Specific PCR fragment of 320 bp in the C*07:72 allele. Specific PCR fragment of 360 bp in the C*07:71 allele.

Primer mix 37: Specific PCR fragment of 225 bp in the C*07:121Q alleles. Specific PCR fragment of 265 bp in the C*07:38:01-07:38:02, 07:68, 07:260 and 07:302 and the C*08:38 and the B*07:77, B*07:193, B*40:60 and B*40:245 alleles.

Primer mix 39: Specific PCR fragment of 100 bp in the C*07:74 allele. Specific PCR fragment of 195 bp in the C*07:308 allele.  Specific PCR fragment of 255 bp in the C*07:59 allele.

Primer mix 40: Specific PCR fragment of 120 bp in the C*07:235Q allele.  Specific PCR fragment of 205 bp in the C*07:62 allele. Specific PCR fragment of 440 bp in the C*07:07, 07:09, 07:76:01-07:76:02, 07:78, 07:315 and 07:328 and in the C*06:17 and 18:01-18:07N alleles.

Primer mix 41: Specific PCR fragment of 140 bp in the C*07:91 allele. Specific PCR fragment of 170 bp in the C*07:234 allele. Specific PCR fragment of 255 bp in the C*07:88 allele.

Primer mix 42: Specific PCR fragment of 100 bp in the C*07:15 and 07:207 alleles. Specific PCR fragment of 205 bp in the C*07:77 and 07:93 alleles. Specific PCR fragment of 250 bp in the C*07:123, 07:173 and 07:294 and in the B*08:01:25, B*35:01:38, B*35:08:04 and B*57:01:18 alleles. Specific PCR fragment of 100 bp and 250 in the B*07:33:03 allele.

Primer mix 43: Specific PCR fragment of 105 bp in the C*07:94 and 07:183 alleles. Specific PCR fragment of 140 bp in the C*07:99 allele. Specific PCR fragment of 165 bp in the C*07:80 and 07:189 alleles.

Primer mix 44: Specific PCR fragment of 165 bp in the C*07:98N allele. Specific PCR fragment of 195 bp in the C*07:65 and 07:77 alleles.

Primer mix 45: Specific PCR fragment of 240 bp in the C*07:245 allele. Specific PCR fragment of 475 bp in the C*07:95 and 07:342 and the C*04:34 and 04:122 and in the B*35:178 and B*73:01-73:02 alleles.

Primer mix 46: Specific PCR fragment of 315 bp in the C*07:202 allele. Specific PCR fragment of 350 bp in the C*07:90 allele. Specific PCR fragment of 420 bp in the C*07:152N allele.

Primer mix 47: Specific PCR fragment of 200 bp in the C*07:86 allele. Specific PCR fragment of 255 bp in the C*07:222 allele. Specific PCR fragment of 420 bp in the C*07:152N allele.Specific PCR fragment of 455 bp in the C*07:97 allele.

Primer mix 48: Specific PCR fragment of 155 bp in the C*07:122 and 07:373 alleles. Specific PCR fragment of 220 bp in the C*07:87 allele. Specific PCR fragment of 255 bp in the C*07:222. Specific PCR fragment of 400 bp in the C*07:85 allele.

Primer mix 49: Specific PCR fragment of 155 bp in the C*07:101, 07:148 and 07:161 and the C*04:129, 05:42, 05:46, 06:67, 08:05, 08:21 and 17:05 alleles. Specific PCR fragment of 230 bp in the C*07:100, 07:259, 07:316 and 07:338 and the C*01:17, 01:21 and 01:23 alleles.

Primer mix 50: Specific PCR fragment of 105 bp in the C*07:104N and in the B*15:272N alleles. Specific PCR fragment of 170 bp in the C*07:119 and in the B*07:87 and B*27:55 alleles. Specific PCR fragment of 425 bp in the C*07:124 and 07:155 alleles.

Primer mix 51: Specific PCR fragment of 65 bp in the C*07:126 allele. Specific PCR fragment of 235 bp in the C*07:121Q allele. Specific PCR fragment of 285 bp in the C*07:117 allele.

Primer mix 52: Specific PCR fragment of 65 bp in the C*07:203 allele. Specific PCR fragment of 205 bp in the C*07:118 allele. Specific PCR fragment of 280 bp in the C*07:106 allele.

Primer mix 53: Specific PCR fragment of 110 bp in the C*07:140, 07:151 and 07:154 and in the B*08:01:25 and B*37:04:02 alleles. Specific PCR fragment of 155 bp in the C*07:122 and 07:373 alleles.

Primer mix 54: Specific PCR fragment of 120 bp in the C*07:153 and 07:324 and in the B*51:142 alleles. Specific PCR fragment of 235 bp in the C*07:64 and 07:374 alleles. Specific PCR fragment of 340 bp in the C*07:125 and 07:356 alleles. Specific PCR fragment of 370 bp in the C*07:129 allele.

Primer mix 55: Specific PCR fragment of 160 bp in the C*07:179 and 07:234 and the C*06:07 alleles. Specific PCR fragment of 210 bp in the C*07:252 and in the A*11:179 alleles. Specific PCR fragment of 290 bp in the C*07:132 allele. Specific PCR fragment of 530 bp in the C*07:343 alleles.

Primer mix 56: Specific PCR fragment of 210 bp in the C*07:150Q allele. Specific PCR fragment of 255 bp in the C*07:123, 07:143, 07:173 and 07:294 and in the B*07:33:03, B*08:01:25, B*35:01:38, 35:08:04 and B*57:01:18 alleles.

Primer mix 57: Specific PCR fragment of 70 bp in the C*07:01:25 allele. Specific PCR fragment of 175 bp in the C*07:130 and the C*05:01:17, 05:09:02, 05:17 and 08:15:02 alleles. Specific PCR fragment of 230 bp in the C*07:259 allele.

Primer mix 58: Specific PCR fragment of 80 bp in the C*07:137:01-07:137:02, 07:186, 07:212 and 07:265 and the C*03:162, 03:220 and in the B*38:43 and B*49:29 alleles. Specific PCR fragment of 450 bp in the C*07:164N allele.

Primer mix 60: Specific PCR fragment of 130 bp in the C*07:165, 07:241, 07:265, 07:304 and 07:307 and the C*03:04:19, 03:113, 04:01:43, 05:01:17, 05:09:02, 06:02:23, 08:15:02 and 12:03:10 and in the B*07:33:03 alleles. Specific PCR fragment of 305 bp in the C*07:147 allele. Specific PCR fragment of 370 bp in the C*07:180 allele.

Primer mix 61: Specific PCR fragment of 245 bp in the C*07:120 alleles. Specific PCR fragment of 405 bp in the C*14:65  alleles.

Primer mix 62: Specific PCR fragment of 200 bp in the C*07:84 and 07:205 and the A*26:101 alleles. Specific PCR fragment of 280 bp in the C *07:106 allele.

Primer mix 63: Specific PCR fragment of 270 bp in the C *07:24, 07:218 and 07:301 alleles. Specific PCR fragment of 320 bp in the C*07:198N and 07:202 alleles.

Primer mix 64: Specific PCR fragment of 135 bp in the C*07:69 allele. Specific PCR fragment of 230 bp in the C *07:192, 07:316 and 07:338 and the C*01:17, 01:21 and 01:23 alleles.

Primer mix 66: Specific PCR fragment of 150 bp in the C*07:136 allele. Specific PCR fragment of 180 bp in the C*07:128 and 07:232 alleles. Specific PCR fragment of 225 bp in the C*07:279 alleles.

Primer mix 68: Specific PCR fragment of 230 bp in the C*07:149 allele. Specific PCR fragment of 260 bp in the C*07:133 and 07:242 and the C*03:178 alleles.

Primer mix 70: Specific PCR fragment of 220 bp in the C*07:131 and 07:194 alleles. Specific PCR fragment of 405 bp in the C*07:61N allele.

Primer mix 71: Specific PCR fragment of 120 bp in the C*07:347N allele. Specific PCR fragment of 240 bp in the C*07:29, 07:177, 07:184 and 07:248 and the C*03:04:19 and 03:113 and in the B*07:33:03 alleles.

Primer mix 72: Specific PCR fragment of 120 bp in the C*07:22, 07:137:01-07:137:02 and 07:138 and the C*06:02:23 and 12:03:10 and in the A*33:03:09 and A*68:01:12 alleles. Specific PCR fragment of 290 bp in the C*07:250 alleles.

Primer mix 74: Specific PCR fragment of 210 bp in the C*07:33N and 07:293 alleles. Specific PCR fragment of 355 bp in the C*07:80 allele.

Primer mix 80: Specific PCR fragment of 155 bp in the C*07:350N allele. Specific PCR fragment of 315 bp in the C*07:316 and 07:359 and the C*06:111 alleles.

Primer mix 83: Specific PCR fragment of 200 bp in the C*07:229 allele. Specific PCR fragment of 250 bp in the C*07:361 allele.

4The following HLA-C*07 alleles can be distinguished by the different sizes of the HLA-specific PCR product:

| Alleles | Primer mix | Alleles | Primer mix |
| --- | --- | --- | --- |
| C*07:29, 07:347N | 71 | C*07:118, 07:203 | 52 |
| C*07:51, 07:314, 07:351 | 19 | C*07:125, 07:374 | 54 |
| C*07:52, 07:337 | 18 | C*07:129, 07:153 | 54 |
| C*07:61N, 07:194 | 70 | C*07:132, 07:179 | 55 |
| C*07:74, 07:308 | 39 | C*07:147, 07:241 | 60 |
| C*07:78, 02:235Q | 40 | C*07:164N, 07:212 | 58 |
| C*07:93, 07:207 | 42 | C*07:165, 07:180 | 60 |
| C*07:100, 07:161 | 49 | C*07:245, 07:342 | 45 |
| C*07:104N, 07:119, 07:124 | 50 | C*07:350N, 07:359 | 80 |
| C*07:117, 07:126 | 51 |  |  |

The HLA-C*07 subtyping kit cannot distinguish the following silent mutations: the C*07:01:01:01-07:01:04, 07:01:06-07:01:12, 07:01:14-07:01:24, 07:01:26-07:01:35 and 07:01:37-07:01:43 alleles, the C*07:01:05 and 07:01:36 alleles, the C*07:02:01:01-07:02:59 alleles, the C*07:04:01-07:04:09, the C*07:17:01-07:17:02, the C*07:27:01-07:27:02, the C*07:31:01-07:31:02, the C*07:38:01-07:38:02, the C*07:56:01-07:56:02, the C*07:96:01-07:96:02, the C*07:108:01-07:108:02 alleles, the C*07:137:01-07:137:02, the C*07:141:01-07:141:02 alleles, the C*07:172:01-07:172:02 alleles or the C*07:199:01-07:199:02 alleles.

5This lot of the C*07 kit cannot distinguish the C*07:75 and the C*07:02:01:01-07:02:59, 07:105, 07:114, 07:135, 07:144, 07:146, 07:157-07:160, 07:167, 07:169, 07:174-07:175, 07:178, 07:185, 07:187, 07:193, 07:195, 07:208-07:209, 07:211, 07:213, 07:217, 07:221, 07:226, 07:233, 07:239-07:240, 07:244, 07:251, 07:261-07:262, 07:270-07:271, 07:273-07:275, 07:283-07:284, 07:286-07:287, 07:290-07:291, 07:305-07:306, 07:309, 07:312, 07:318-07:322, 07:325-07:327, 07:333-07:334, 07:339-07:341, 07:344-07:346, 07:360, 07:363, 07:368 and 07:372 alleles, nor the C*07:254 and the C*07:01:01:01-07:01:04, 07:01:06-07:01:12, 07:01:14-07:01:24, 07:01:26-07:01:35, 07:01:37-07:01:43, 07:103, 07:111, 07:113, 07:115, 07:156, 07:162-07:163, 07:170, 07:182, 07:188, 07:190, 07:197, 07:200-07:201, 07:204, 07:206, 07:214-07:215, 07:219, 07:223, 07:225, 07:228, 07:236-07:237, 07:246, 07:249, 07:253, 07:255, 07:257, 07:266-07:269, 07:276-07:278, 07:280-07:282, 07:296-07:300, 07:303, 07:310-07:311, 07:313, 07:330-07:332, 07:362, 07:366 and 07:369-07:371 alleles.

‘w’, may be weakly amplified.

Change in revision R01 compared to R00:

- Primer mix 16 does not amplify the C*07:25 allele. This has been corrected in the Specificity and Interpretation Tables.

Changes in revision R02 compared to R01:

  - Primer mix 30 does not amplify the C*07:75 and 07:254 alleles. Thus, this lot of the C*07 kit cannot distinguish the C*07:75 and the C*07:02:01:01-07:02:59, 07:105, 07:114, 07:135, 07:144, 07:146, 07:157-07:160, 07:167, 07:169, 07:174-07:175, 07:178, 07:185, 07:187, 07:193, 07:195, 07:208-07:209, 07:211, 07:213, 07:217, 07:221, 07:226, 07:233, 07:239-07:240, 07:244, 07:251, 07:261-07:262, 07:270-07:271, 07:273-07:275, 07:283-07:284, 07:286-07:287, 07:290-07:291, 07:305-07:306, 07:309, 07:312, 07:318-07:322, 07:325-07:327, 07:333-07:334, 07:339-07:341, 07:344-07:346, 07:360, 07:363, 07:368 and 07:372 alleles, nor the C*07:254 and the C*07:01:01:01-07:01:04, 07:01:06-07:01:12, 07:01:14-07:01:24, 07:01:26-07:01:35, 07:01:37-07:01:43, 07:103, 07:111, 07:113, 07:115, 07:156, 07:162-07:163, 07:170, 07:182, 07:188, 07:190, 07:197, 07:200-07:201, 07:204, 07:206, 07:214-07:215, 07:219, 07:223, 07:225, 07:228, 07:236-07:237, 07:246, 07:249, 07:253, 07:255, 07:257, 07:266-07:269, 07:276-07:278, 07:280-07:282, 07:296-07:300, 07:303, 07:310-07:311, 07:313, 07:330-07:332, 07:362, 07:366 and 07:369-07:371 alleles. This has been corrected in the Specificity and Interpretation Tables.