PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® HLA-C*07 Add-on

Prod. No:	101.863-12/12u

Lot No:    	44S

Expiry Date:    	2015-December-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 5 primer mixes used for HLA-C*07 Add-on SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified HLA-C*07 alleles3,4 |
| --- | --- | --- | --- |
| 16 | 245 bp, 425 bp | 800 bp | *07:01:01:01-07:33N, 07:35-07:294, 07:296-07:307 |
| 2 | 505 bp | 1070 bp | *07:06 |
| 3 | 225 bp | 1070 bp | *07:06, 07:18-07:19 |
| 45 | 80 bp | 1070 bp | *07:11 |
| 55 | 80 bp | 1070 bp | *07:50 |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-C*07 Add-on SSP typings.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most wells, or a band of 800 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-C*07 Add-on subtyping.

3For several HLA Class I alleles 1st and/or 4th exon(s) and beyond, as well as intron nucleotide sequences, are not available. In these instances it is not known whether some of the primers of the SSP sets are completely matched with the target sequences or not. We assume that unknown sequences in these regions are conserved within allelic groups.

4Due to the sharing of sequence motifs between HLA-C alleles all C*07 alleles will be amplified by primer mix 1.

5HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

6The primer pairs in wells 1 will give rise to two HLA-specific PCR fragments for many C*07 alleles.

1The HLA-Cw*0734 allele has been renamed to HLA-C*07:27:02.

The C*07:295 allele has never been assigned.