Gel Documentation Form and Worksheet

HLA-C*12 (101.624-12/12u) 	Lot No: 3D5	Expiry Date: 2018-June-01

Name:__________________

Sample ID:______________

DNA Conc.(ng/ul):_________

Test  Date:_______________

Tested By:_______________

Review Date:_______________

Reviewed By:_______________

Interpretation:___________                           Failed lanes: ____________                          Comments:________________

Gel Picture

PHOTO DOCUMENT

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 1, 5 and 12 may give rise to a lower yield of HLA-specific PCR product than the other C*12 primer mixes.

Primer mixes 8 and 28 have a tendency to giving rise to primer oligomer formation.

Primer mixes 30 and 40 may have tendencies of unspecific amplifications.

Primer mix 45 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 430 base pairs.

1HLA-C*12 alleles in bold lettering are listed as confirmed alleles on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.22.0, October 2015.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3Primer mix 6: Specific PCR fragment of 75 bp in the C*12:48 and 12:102 alleles. Specific PCR fragment of 150 bp in the C*12:06 allele. Specific PCR fragment of 415 bp in the C*12:08 and 12:81 and in the C*03:08, 03:29, 03:31, 03:246, 04:112, 04:169, 05:36, 06:44, 15:88-15:95N, 15:97-15:103, 15:106, 16:64, 16:70 and 16:87 alleles.

Primer mix 8: Specific PCR fragment of 95 bp in the C*12:15 and the C*15:02:14 alleles. Specific PCR fragment of 155 bp in the C*12:40 and in the C*16:14 alleles. Specific PCR fragment of 195 bp in the C*12:80N allele. Specific PCR fragment of 245 bp in the C*12:07 allele.

Primer mix 13: Specific PCR fragment of 105 bp in the C*12:31 and the C*08:01:19 alleles. Specific PCR fragment of 150 bp in the C*12:10:01-12:10:02, 12:155Q and 12:156 and the C*14:02:08 alleles. Specific PCR fragment of 105 and 150 bp in the C*04:01:05 allele.

Primer mix 14: Specific PCR fragment of 100 bp in the *12:16 and 12:147 and the C*01:02:34, 01:21, 02:42, 02:107, 04:140, 04:166, 05:98, 06:05w, 07:02:09, 08:14, 08:80, 08:103, 15:63, 15:113, 16:80 and in the B*67:02 alleles. Specific PCR fragment of 150 bp in the C*12:11, 12:60 and 12:118 allele.

Primer mix 16: Specific PCR fragment of 185 bp in the C*12:13 allele. Specific PCR fragment of 225 bp in the C*12:14:01-12:14:02 and the C*01:60, 04:58, 04:160, 05:23, 05:62, 06:118, 08:07, 08:47, 08:104, 14:17, 15:65 and 17:01:01:01-17:30 alleles.

Primer mix 17: Specific PCR fragment of 130 bp in the C*12:159 allele. Specific PCR fragment of 565 bp in the C*12:03:04, 12:03:09 and the C*01:02:18, 06:02:38, 07:447 and 14:02:08 and in the B*27:05:27 alleles.

Primer mix 18: Specific PCR fragment of 145 bp in the C*04:12 allele. Specific PCR fragment of 245 bp in C*03:53 allele. Specific PCR fragments of 145 bp and 245 bp in the C*12:17 and 12:27 alleles. Specific PCR fragment of 165 and 270 bp in C*12:35 allele. All specific bands may not always be visible.

Primer mix 20: Specific PCR fragment of 105 bp in the C*12:46N allele. Specific PCR fragment of 180 bp in the C*12:22, 12:58 and 12:94 and the C*04:52, 04:55, 05:55, 14:10, 14:48 and 15:12 alleles. Specific PCR fragment of 235 bp in the C*12:19 and 12:139 alleles.

Primer mix 22:  Specific PCR fragment of 100 bp in the C*12:15, 12:23 and 12:99 and the C*15:02:14 alleles. Specific PCR fragments of 590 bp in C*12:21 and the C*08:01:19 alleles.

Primer mix 24: Specific PCR fragment of 135 bp in the C*12:99 allele. Specific PCR fragment of 185 bp in the C*12:43 allele. Specific PCR fragment of 425 bp in the C*12:28 and 12:135 and the C*04:01:05, 06:02:38, 06:76:02 and 07:447 alleles.

Primer mix 25: Specific PCR fragment of 80 bp in the C*12:39N allele. Specific PCR fragment of 155 bp in the C*12:02:06, 12:02:08, 12:02:12, 12:21 and 12:118 and the C*04:01:05, 08:01:19 and 08:02:02 and in the B*27:05:27 and B*40:02:21 alleles.

Primer mix 27: Specific PCR fragment of 100 bp in the C*12:30 and the C*07:214 and 07:429 alleles. Specific PCR fragment of 155 bp in the C*12:03:19, 12:03:22 and 12:155Q and the C*01:04 and 16:02:13 alleles. Specific fragment of 185 bp in the C*12:94 and the C*14:48 alleles. Specific PCR fragment of 205 bp in the C*12:36 allele. Specific PCR fragment of 295 bp in the C*12:101 and the C*16:81 alleles. Specific PCR fragment of 570 bp in the A*02:466, A*02:504 and A*02:596 alleles.

Primer mix 28: Specific PCR fragment of 275 bp in the C*12:50 and the C*01:32, 02:56, 03:102, 03:263, 04:180:01, 06:20, 07:81, 07:168, 07:450 and 08:123 alleles. Specific PCR fragment of 350 bp in C*12:45 and 12:166 and the C*05:81, 06:87, 07:24, 07:218, 14:65, 16:13 and 16:61 alleles.

Primer mix 29: Specific PCR fragment of 125 bp in the C*12:38 and 12:104N and in the B*15:181N and B*57:50 alleles. Specific PCR fragment of 185 bp in the C*12:42Q and in the B*46:51Q allele. Specific PCR fragments of 210 bp in C*12:29 and 12:86 alleles.

Primer mix 30: Specific PCR fragment of 85 bp in the C*12:32 and 12:102 and the C*06:41 alleles. Specific PCR fragment of 190 bp in the C*12:143 allele. Specific PCR fragment of 230 bp in C*12:34 allele.

Primer mix 31: Specific PCR fragment of 120 bp in the C*12:47 and in the A*11:197, A*26:67 and A*68:95 alleles. Specific PCR fragment of 145 bp in the C*12:84N and 12:123 alleles. Specific PCR fragment of 190 bp in the C*12:42Q and 12:80N and in the B*46:51Q alleles. Specific PCR fragment of 180 bp in the C*12:62 and in the A*02:335 allelles.

Primer mix 32: Specific PCR fragment of 50 bp in the C*12:54 and the C*01:59, 02:65, 03:130, 03:140, 03:243, 04:114, 05:20, 06:82, 07:49, 07:210, 07:238, 07:247, 07:403,  14:04, 14:64, 15:85 and 16:57 and in the A*68:46 and in the B*07:253 alleles. Specific PCR fragment of 115 bp in C*12:37 and the C*07:204 and in the A*02:211:01, A*02:594, A*24:261 and A*68:76:01-68:76:02 alleles. Specific PCR fragment of 180 bp in the C*12:62 and in the A*02:335 alleles.

Primer mix 34: Specific PCR fragment of 235 bp in the C*12:109 and the C*03:171, 03:211:01, 04:144, 05:93, 06:73, 08:20 and 08:40 alleles. Specific PCR fragment of 260 bp in the C*12:125 allele.

Primer mix 35: Specific PCR fragment of 195 bp in the C*12:110 and 12:143 alleles. Specific PCR fragment of 260 bp in the C*12:125 allele.

Primer mix 37: Specific PCR fragment of 225 bp in the C*12:108 allele. Specific PCR fragment of 285 bp in the C*12:73 and in the B*53:34 and B*58:21 alleles.

‘w’, might be weakly amplified.

4The HLA-C*12 primer set cannot separate the C*12:09, C*05:16, C*05:85 and C*05:107 alleles, the C*12:16 and C*01:21 alleles, the C*12:18:02 and C*08:21 alleles, the C*12:153 and C*16:15:02 alleles, or the C*12:33 and the C*02:05:01-02:05:03 and 02:17 alleles. These alleles can be distinguished by the HLA-C low resolution kit and the HLA-C*01, HLA-C*02, HLA-C*05, HLA-C*08 or HLA-C*16 kit, respectively.

5The following HLA-C*12 alleles can be distinguished by the different sizes of the HLA-specific PCR product:

| Alleles | Primer mix |
| --- | --- |
| C*12:03:09, 12:159 | 17 |
| C*12:06, 12:48, 12:81 | 6 |
| C*12:29, 12:38 | 29 |
| C*12:30, 12:36 | 27 |
| C*12:32, 12:34 | 30 |
| C*12:45, 12:50 | 28 |
| C*12:46N, 12:139 | 20 |
| C*12:86, 12:104N | 29 |