PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® HLA-C*14

Prod. No:	101.625-06/06u

Lot No:    	20L

Expiry Date:    	2013-September-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 23 primer mixes used for HLA-C*14 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified HLA-C*14 alleles3 | Other amplified HLA Class I alleles4 |
| --- | --- | --- | --- | --- |
| 1 | 145 bp | 800 bp | *14:02:01-14:02:02, 14:02:04-14:02:07, 14:04-14:07N, 14:09, 14:11-14:29 | *07:64 |
| 2 | 145 bp | 1070 bp | *14:02:03, 14:03, 14:08, 14:10 | *04:01:01:01-04:01:22, 04:01:24-04:01:26, 04:04:01-04:05, 04:07-04:15:03, 04:17-04:20, 04:23-04:25, 04:26w, 04:27-04:41, 04:43-04:55, 04:57-04:70, 04:72-04:79, 04:81-04:84, 15:36 |
| 3 | 245 bp | 800 bp | *14:04, 14:12 | *04:01:01:01-04:01:26, 04:04:01-04:05, 04:07-04:10, 04:12-04:15:03, 04:17-04:20, 04:23-04:25, 04:26w, 04:27-04:28, 04:30-04:35, 04:37-04:41, 04:43-04:54, 04:56-04:70, 04:72-04:79, 04:81-04:84, 15:36 |
| 4 | 140 bp | 1070 bp | *14:02:03, 14:03, 14:08, 14:10, 14:22 | *01:21, 02:12w, 02:27:01-02:27:02, 04:11, 04:29, 04:36, 04:55, 07:02:09, 08:01:01-08:09, 08:11-08:46, 12:02:01-12:03:03, 12:03:05-12:03:08, 12:03:10-12:03:15, 12:06-12:08, 12:10:01-12:20, 12:22-12:26, 12:28-12:32, 12:34-12:40, 12:42Q-12:51, 15:07, 15:21w, 15:25, 16:01:01, 16:01:03-16:01:06, 16:04:01, 16:06-16:08, 16:10-16:11, 16:13-16:18, 16:20-16:24, 16:26-16:32, B*35:08:02, B*67:02 |
| 5 | 210 bp | 1070 bp | *14:02:01-14:04, 14:07N, 14:10-14:14, 14:17-14:27, 14:29 |  |
| 65 | 130 bp | 1070 bp | *14:05, 14:09 | *03:02:01-03:04:20, 03:06-03:13, 03:15-03:24, 03:26, 03:28-03:43:01, 03:44-03:57, 03:59-03:85, 03:88, 03:91-03:93, 03:95-03:98, 03:100-03:101, 03:104-03:108, 04:15:01-04:15:03, 04:17, 04:80, 06:02:08, 16:01:06 |
| 75 | 130 bp | 1070 bp | *14:04 | *01:17, 01:23, 03:15, 03:27, 03:38:01-03:38:02, 03:69, 04:01:23, 06:02:01:01-06:02:01:02, 06:02:03-06:04, 06:06-06:43, 06:45-06:58, 07:01:01-07:02:07, 07:02:10-07:20, 07:22-07:33N, 07:35-07:38, 07:41-07:68, 07:70-07:71, 07:73-07:151, 07:153-07:155, 07:157-07:171, 12:03:04, 12:03:09, 12:04:01, 17:01:01:01-17:01:03, 17:01:05-17:08, 18:01-18:04, B*07:13, B*07:15 |
| 8 | 210 bp | 1070 bp | *14:03, 14:10, 14:22 | *01:48, 02:12w, 02:27:01-02:27:02, 03:02:01-03:04:16, 03:04:18-03:06, 03:08-03:09, 03:10w, 03:12-03:14, 03:16-03:17, 03:19-03:28, 03:29w, 03:30-03:36, 03:38:01-03:38:02, 03:40-03:44, 03:46-03:49, 03:51-03:63, 03:65-03:98, 03:100-03:106, 03:108, 07:20, 07:96, 07:127, 15:07, 15:21w, 15:25, 15:43, B*55:30 |
| 9 | 140 bp | 1070 bp | *14:02:01, 14:02:03-14:08, 14:10-14:16, 14:18-14:27, 14:29 | *03:17, 03:71, 04:37, 06:02:08, 16:01:06 |
| 105,6 | 100 bp, 210 bp | 1070 bp | *14:06, 14:15, 14:21N |  |
| 11 | 205 bp | 1070 bp | *14:07N |  |
| 127 | 225 bp, 320 bp | 1070 bp | *14:08, 14:23 |  |
| 135,8 | 125 bp, 285 bp | 1070 bp | *14:10, 14:14 | *04:52, 04:55 |
| 149 | 200 bp, 290 bp | 1070 bp | *14:11, 14:15 |  |
| 15 | 150 bp | 800 bp | *14:12 | *01:14, 02:02:13, 03:07, 03:10, 03:15, 03:29, 03:45, 04:01:23, 05:01:12, 06:02:01:01-06:02:01:02, 06:02:03-06:04, 06:06-06:10, 06:12-06:58, 07:07, 07:09, 07:76, 12:04:01, 17:01:01:01-17:01:03, 17:01:05-17:08, 18:01-18:04 |
| 16 | 410 bp | 1070 bp | *14:13 | *03:15, 03:32, 03:45, 03:60, 04:01:01:01-04:01:04, 04:01:06, 04:01:08, 04:01:10-04:01:26, 04:03-04:04:01, 04:05-04:10, 04:12-04:20, 04:23-04:32, 04:34-04:54, 04:56-04:84, 06:02:08, 18:01-18:04 |
| 1710 | 150 bp, 210 bp | 1070 bp | *14:16, 14:22, 14:27 | *02:17 |
| 185,11 | 90 bp, 185 bp | 1070 bp | *14:17, 14:26 |  |
| 195,12 | 85 bp, 185 bp | 1070 bp | *14:18, 14:29 |  |
| 205,13 | 105 bp, 140 bp | 1070 bp | *14:19, 14:28 | *03:88 |
| 21 | 230 bp | 1070 bp | *14:20 |  |
| 225 | 95 bp | 1070 bp | *14:24 |  |
| 23 | 230 bp | 1070 bp | *14:25 | *02:17 |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-C*14 resolution SSP typings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective length of the specific PCR product(s) of the alleles amplified by these primer mixes is given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most wells, or a band of 800 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-C*14 SSP subtyping.

In addition, wells number 3 and 15 contain the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3The C*14:18 and 14:29 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 19.

4Due to the sharing of sequence motifs between HLA-C alleles non-HLA-C*14 alleles will be amplified by primer mixes 1 to 4, 6 to 9, 13, 15 to 17, 20 and 23. In addition, primer mix 4 will amplify the B*35:0802 and B*67:02 alleles, primer mix 7 will amplify the B*07:13 and B*07:15 alleles and primer mix 8 will amplify the B*55:30 allele.

5Specific PCR fragments shorter than 130 base pairs have a lower intensity and are less sharp than longer PCR bands.

6Primer mix 10: specific PCR product of 100 bp in the C*14:21N allele. Specific PCR product of 210 bp in the C*14:06 and 14:15 alleles.

7Primer mix 12: specific PCR product of 225 bp in the C*14:08 allele. Specific PCR product of 320 bp in the C*14:23 allele.

8Primer mix 13: specific PCR product of 125 bp in the C*14:14 allele. Specific PCR product of 285 bp in the C*14:10 and the C*04:52 and 04:55 alleles.

9Primer mix 14: specific PCR product of 200 bp in the C*14:15 allele. Specific PCR product of 290 bp in the C*14:11 allele.

10Primer mix 17: specific PCR product of 150 bp in the C*14:22 and 14:27 alleles. Specific PCR product of 210 bp in the C*14:16 and the C*02:17 alleles.

11Primer mix 18: specific PCR product of 90 bp in the C*14:26 allele. Specific PCR product of 185 bp in the C*14:17 allele.

12Primer mix 19: specific PCR product of 85 bp in the C*14:18 allele. Specific PCR product of 185 bp in the C*14:29 allele.

13Primer mix 20: specific PCR product of 105 bp in the C*14:19 and the C*03:88 alleles. Specific PCR product of 140 bp in the C*14:28 allele.

‘w’, may be weakly amplified.

1The HLA-Cw*1401 nucleotide sequence has been shown to be identical to C*14:02.

2The C*14:18 and 14:29 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 19.

3Primer mix 10: specific PCR product of 100 bp in the C*14:21N allele. Specific PCR product of 210 bp in the C*14:06 and 14:15 alleles.

Primer mix 12: specific PCR product of 225 bp in the C*14:08 allele. Specific PCR product of 320 bp in the C*14:23 allele.

Primer mix 13: specific PCR product of 125 bp in the C*14:14 allele. Specific PCR product of 285 bp in the C*14:10 and the C*04:52 and 04:55 alleles.

Primer mix 14: specific PCR product of 200 bp in the C*14:15 allele. Specific PCR product of 290 bp in the C*14:11 allele.

Primer mix 17: specific PCR product of 150 bp in the C*14:22 and 14:27 alleles. Specific PCR product of 210 bp in the C*14:16 and the C*02:17 alleles.

Primer mix 18: specific PCR product of 90 bp in the C*14:26 allele. Specific PCR product of 185 bp in the C*14:17 allele.

Primer mix 19: specific PCR product of 85 bp in the C*14:18 allele. Specific PCR product of 185 bp in the C*14:29 allele.

Primer mix 20: specific PCR product of 105 bp in the C*14:19 and the C*03:88 alleles. Specific PCR product of 140 bp in the C*14:28 allele.

‘w’, may be weakly amplified.