PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® HLA-C*14

Prod. No:	101.625-06/06u

Lot No:    	12S

Expiry Date:    	2015-October-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 24 primer mixes used for HLA-C*14 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified HLA-C*14 alleles3 | Other amplified HLA Class I alleles4 |
| --- | --- | --- | --- | --- |
| 1 | 145 bp | 800 bp | *14:02:01-14:02:02, 14:02:04-14:02:11, 14:04-14:07N, 14:09, 14:11-14:33, 14:36-14:37, 14:39-14:40, 14:42-14:49 | *07:64 |
| 26 | 145 bp | 1070 bp | *14:02:03, 14:03, 14:08, 14:10, 14:35N, 14:38, 14:41 | *04:01:01:01-04:01:22, 04:01:24-04:01:48, 04:04:01-04:05, 04:07-04:15:03, 04:17-04:20, 04:23-04:25, 04:26w, 04:27-04:41, 04:43-04:55, 04:57-04:70, 04:72-04:79, 04:81-04:89, 04:90w, 04:91-04:102, 04:104-04:106, 04:108-04:144, 05:78, 12:55, 15:36 |
| 3 | 245 bp | 800 bp | *14:04, 14:12, 14:49 | *04:01:01:01-04:01:28, 04:01:30-04:01:48, 04:04:01-04:05, 04:07-04:10, 04:12-04:15:03, 04:17-04:20, 04:23-04:25, 04:26w, 04:27-04:28, 04:30-04:35, 04:37-04:41, 04:43-04:54, 04:56-04:70, 04:72-04:79, 04:81-04:89, 04:90w, 04:91-04:102, 04:104-04:106, 04:108-04:144, 05:78, 15:36 |
| 4 | 140 bp | 1070 bp | *14:02:03, 14:03, 14:08, 14:10, 14:22, 14:35N, 14:38, 14:41 | *01:21, 02:12w, 02:27:01-02:27:02, 03:04:25, 04:11, 04:29, 04:36, 04:55, 07:02:09, 08:01:01-08:09, 08:11-08:63, 08:65-08:74, 12:02:01-12:03:03, 12:03:05-12:03:08, 12:03:10-12:03:23, 12:03:24w, 12:03:25, 12:06-12:08, 12:10:01-12:20, 12:22-12:26, 12:28-12:32, 12:34-12:40, 12:42Q-12:53, 12:55-12:59, 12:61-12:71, 12:72w, 12:73-12:94, 15:07, 15:21w, 15:25, 16:01:01, 16:01:03-16:01:13, 16:04:01, 16:06-16:08, 16:10-16:11, 16:13-16:18, 16:20-16:24, 16:26-16:36, 16:37w, 16:38-16:45, 16:49-16:55, B*35:08:02, B*35:08:05, B*67:02 |
| 5 | 210 bp | 1070 bp | *14:02:01-14:04, 14:07N, 14:10-14:14, 14:17-14:27, 14:29-14:34, 14:36-14:44, 14:46-14:49 |  |
| 6 | 130 bp | 1070 bp | *14:05, 14:09 | *03:02:01-03:02:07, 03:02:09-03:04:26, 03:04:28-03:04:29, 03:04:31-03:04:35, 03:06-03:13, 03:15-03:24, 03:26, 03:28-03:40:01, 03:40:03-03:43:01, 03:44-03:57, 03:59-03:85, 03:88, 03:91-03:93, 03:95-03:98, 03:100-03:101, 03:104-03:118, 03:120-03:121N, 03:123-03:127, 03:129-03:134, 03:136-03:142, 03:144-03:152, 03:154-03:166, 03:168-03:177, 03:179-03:185, 04:15:01-04:15:03, 04:17, 04:80, 04:100, 06:02:08, 08:01:07, 08:02:07, 08:33:02, 12:03:20, 16:01:06, B*15:78:03 |
| 7 | 130 bp | 1070 bp | *14:04, 14:49 | *01:17, 01:23, 03:15, 03:27, 03:38:01-03:38:02, 03:69, 03:130, 03:136, 03:163, 04:01:23, 06:02:01:01-06:02:01:02, 06:02:03-06:04, 06:06-06:43, 06:45-06:60, 06:62-06:76:01, 06:77-06:96, 07:01:01:01-07:02:07, 07:02:10-07:02:28, 07:02:30-07:20, 07:22-07:33N, 07:35-07:38:02, 07:41-07:68, 07:70-07:71, 07:73-07:151, 07:153-07:155, 07:157-07:176, 07:178-07:209, 07:211-07:222, 07:223w, 07:224-07:237, 07:239-07:294, 07:296-07:299, 12:03:04, 12:03:09, 12:04:01, 17:01:01:01-17:01:03, 17:01:05-17:19, 18:01-18:06, B*07:13, B*07:15, B*07:160, B*42:18 |
| 8 | 210 bp | 1070 bp | *14:03, 14:10, 14:22, 14:35N, 14:38, 14:41 | *01:48, 02:12w, 02:27:01-02:27:02, 03:02:01-03:03:14, 03:03:15w, 03:03:16-03:04:16, 03:04:18-03:06, 03:08-03:09, 03:10w, 03:12-03:14, 03:16-03:17, 03:19-03:28, 03:29w, 03:30-03:36, 03:38:01-03:38:02, 03:40:01-03:44, 03:46-03:49, 03:51-03:63, 03:65-03:98, 03:100-03:106, 03:108-03:114, 03:116:01-03:129, 03:131-03:133, 03:135-03:139, 03:141-03:162, 03:164-03:185, 07:20, 07:96:01-07:96:02, 07:127, 07:263, 15:07, 15:21w, 15:25, 15:43, 16:34, B*55:30 |
| 9 | 140 bp | 1070 bp | *14:02:01, 14:02:03-14:08, 14:10-14:16, 14:18-14:27, 14:29-14:49 | *03:17, 03:71, 04:37, 06:02:08, 12:03:20, 16:01:06 |
| 105,7 | 100 bp, 145, 210 bp | 1070 bp | *14:06, 14:15, 14:21N, 14:47N |  |
| 11 | 205 bp | 1070 bp | *14:07N | B*15:258N |
| 128 | 225 bp, 320 bp | 1070 bp | *14:08, 14:23 |  |
| 135,9 | 125 bp, 285 bp | 1070 bp | *14:10, 14:14 | *04:52, 04:55 |
| 1410 | 200 bp, 290 bp | 1070 bp | *14:11, 14:15 |  |
| 15 | 150 bp | 800 bp | *14:12, 14:49 | *01:14, 02:02:13, 03:07, 03:10, 03:15, 03:29, 03:45, 03:163, 04:01:23, 05:01:12, 05:29:02, 06:02:01:01-06:02:01:02, 06:02:03-06:04, 06:06-06:10, 06:12-06:76:01, 06:77-06:81, 06:83-06:96, 07:07, 07:09, 07:76, 12:04:01, 15:02:10, 17:01:01:01-17:01:03, 17:01:05-17:19, 18:01-18:06, B*35:222 |
| 16 | 410 bp | 1070 bp | *14:13, 14:49 | *03:15, 03:32, 03:45, 03:60, 03:136, 04:01:01:01-04:01:04, 04:01:06, 04:01:08, 04:01:10-04:01:32, 04:01:34-04:01:48, 04:03-04:04:01, 04:05-04:10, 04:12-04:20, 04:23-04:32, 04:34-04:54, 04:56-04:94:01, 04:95N-04:97, 04:99-04:106, 04:108-04:144, 06:02:08, 07:02:32, 07:242, 18:01-18:06 |
| 1711 | 150 bp, 210 bp | 1070 bp | *14:16, 14:22, 14:27, 14:47N | *02:17 |
| 185,12 | 90 bp, 185 bp | 1070 bp | *14:17, 14:26 |  |
| 195,13 | 85 bp, 185 bp | 1070 bp | *14:18, 14:29 | *04:140 |
| 205,14 | 105 bp, 140 bp | 1070 bp | *14:19, 14:28, 14:35N | *03:88 |
| 215,15 | 125 bp, 230 bp | 1070 bp | *14:20, 14:32 |  |
| 225,16 | 95 bp, 250 bp | 1070 bp | *14:24:01-14:24:02, 14:31 | *03:23 |
| 235,17 | 125 bp, 230 bp | 1070 bp | *14:25, 14:32 | *02:17 |
| 24 | 145 bp | 1070 bp | *14:33-14:34 | *07:267 |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-C*14 resolution SSP typings.

When the primers in a primer mix can give rise to HLA-specific PCR products of more than one length this is indicated if the size difference is more than 20 base pairs. Size differences of 20 base pairs or less are not given. For high resolution SSP kits the respective lengths of the HLA-specific PCR product(s) are given for the alleles amplified by these primer mixes.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most wells, or a band of 800 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-C*14 SSP subtyping.

In addition, wells number 3 and 15 contain the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3For several HLA Class I alleles 1st and/or 4th exon(s) and beyond, as well as intron nucleotide sequences, are not available. In these instances it is not known whether some of the primers of the SSP sets are completely matched with the target sequences or not. We assume that unknown sequences in these regions are conserved within allelic groups.

4Due to the sharing of sequence motifs between HLA-C alleles non-HLA-C*14 alleles will be amplified by primer mixes 1 to 4, 6 to 9, 13, 15 to 17, 19, 20 and 22 to 24. In addition, primer mix 4 will amplify the B*35:08:02, B*35:08:05 and B*67:02 alleles, primer mix 6 will amplify the B*15:78:03 allele, primer mix 7 will amplify the B*07:13, B*07:15, B*07:160 and B*42:18 alleles, primer mix 8 will amplify the B*55:30 allele, primer mix 11 will amplify the B*15:258N allele and primer mix 15 will amplify the B*35:222 allele.

5HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

6Primer mix 2 may have tendency of unspecific amplification.

7Primer mix 10: specific PCR product of 100 bp in the C*14:21N allele. Specific PCR product of 145 bp in the C*14:47N allele. Specific PCR product of 210 bp in the C*14:06 and 14:15 alleles.

8Primer mix 12: specific PCR product of 225 bp in the C*14:08 allele. Specific PCR product of 320 bp in the C*14:23 allele.

9Primer mix 13: specific PCR product of 125 bp in the C*14:14 allele. Specific PCR product of 285 bp in the C*14:10 and the C*04:52 and 04:55 alleles.

10Primer mix 14: specific PCR product of 200 bp in the C*14:15 allele. Specific PCR product of 290 bp in the C*14:11 allele.

11Primer mix 17: specific PCR product of 150 bp in the C*14:22, 14:27 and 14:47N alleles. Specific PCR product of 210 bp in the C*14:16 and the C*02:17 alleles.

12Primer mix 18: specific PCR product of 90 bp in the C*14:26 allele. Specific PCR product of 185 bp in the C*14:17 allele.

13Primer mix 19: specific PCR product of 85 bp in the C*14:18 and the C*04:140 alleles. Specific PCR product of 185 bp in the C*14:29 allele.

14Primer mix 20: specific PCR product of 105 bp in the C*14:19 and the C*03:88 alleles. Specific PCR product of 140 bp in the C*14:28 and 14:35N alleles.

15Primer mix 21: specific PCR product of 125 bp in the C*14:32 allele. Specific PCR product of 230 bp in the C*14:20 allele.

16Primer mix 22: specific PCR product of 95 bp in the C*14:24:01-14:24:02 and the C*03:23 alleles. Specific PCR product of 250 bp in the C*14:31 allele.

17Primer mix 23: specific PCR product of 125 bp in the C*14:32 allele. Specific PCR product of 230 bp in the C*14:25 and the C*02:17 alleles.

‘w’, may be weakly amplified.

1The HLA-Cw*1401 nucleotide sequence has been shown to be identical to C*14:02:01.

2HLA-C*14 alleles in bold lettering are listed as confirmed alleles on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.11.0, January 2013.

3The C*14:18 and 14:29 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 19.

4The C*14:24 and 14:31 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 22.

6Primer mix 10: specific PCR product of 100 bp in the C*14:21N allele. Specific PCR product of 145 bp in the C*14:47N allele. Specific PCR product of 210 bp in the C*14:06 and 14:15 alleles.

Primer mix 12: specific PCR product of 225 bp in the C*14:08 allele. Specific PCR product of 320 bp in the C*14:23 allele.

Primer mix 13: specific PCR product of 125 bp in the C*14:14 allele. Specific PCR product of 285 bp in the C*14:10 and the C*04:52 and 04:55 alleles.

Primer mix 14: specific PCR product of 200 bp in the C*14:15 allele. Specific PCR product of 290 bp in the C*14:11 allele.

Primer mix 17: specific PCR product of 150 bp in the C*14:22, 14:27 and 14:47N alleles. Specific PCR product of 210 bp in the C*14:16 and the C*02:17 alleles.

Primer mix 18: specific PCR product of 90 bp in the C*14:26 allele. Specific PCR product of 185 bp in the C*14:17 allele.

Primer mix 19: specific PCR product of 85 bp in the C*14:18 and the C*04:140 alleles. Specific PCR product of 185 bp in the C*14:29 allele.

Primer mix 20: specific PCR product of 105 bp in the C*14:19 and the C*03:88 alleles. Specific PCR product of 140 bp in the C*14:28  and 14:35N alleles.

Primer mix 21: specific PCR product of 125 bp in the C*14:32 allele. Specific PCR product of 230 bp in the C*14:20 allele.

Primer mix 22: specific PCR product of 95 bp in the C*14:24:01-14:24:02 and the C*03:23 alleles. Specific PCR product of 250 bp in the C*14:31 allele.

Primer mix 23: specific PCR product of 125 bp in the C*14:32 allele. Specific PCR product of 230 bp in the C*14:25 and the C*02:17 alleles.

‘w’, may be weakly amplified.

Change in revision R01 compared to R00:

- The HLA-C*14:41 allele is amplified by primer mix 2. This has been corrected in the specificity and interpretation tables.