PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® HLA-C*15

Prod. No:	101.626-12/12u

Lot No:    	06M

Expiry Date:    	2013-October-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 24 primer mixes used for HLA-C*15 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified HLA-C*15 alleles3 | Other amplified HLA Class I alleles4 |
| --- | --- | --- | --- | --- |
| 1 | 325 bp | 800 bp | *15:02:01-15:03, 15:05:01-15:13, 15:15-15:24, 15:26-15:29, 15:31-15:39, 15:41-15:48 | *02:06, 02:23, 02:36, 03:81, 07:123, 12:15, 16:20 |
| 2 | 175 bp | 1070 bp | *15:02:01-15:09, 15:12-15:13, 15:15, 15:18-15:19, 15:21-15:24, 15:26, 15:28-15:42, 15:44-15:48 | *03:08, 03:29, 03:31, 07:20, 07:96 |
| 35,9 | 70 bp, 195 bp | 1070 bp | *15:03, 15:16, 15:25w, 15:28 | *02:12w, 03:15, 03:27w, 03:38:01w-03:38:02w, 03:69w, 04:03, 04:06, 04:16, 04:80, 07:20w, 07:96w, 07:127w |
| 4 | 315 bp | 1070 bp | *15:04, 15:09, 15:19-15:20, 15:27, 15:30 | *01:04, 01:21, 02:02:01-02:02:03, 02:02:05-02:05, 02:08-02:18, 02:20-02:21, 02:24-02:40, 02:42-02:46, 03:02:01-03:02:07, 03:14-03:16, 03:33, 03:36, 03:40, 03:42-03:43:02, 03:60, 03:71, 03:84, 03:89, 03:95, 03:108, 04:54, 05:04, 06:02:01:01-06:02:01:02, 06:02:03-06:08, 06:10-06:33, 06:35-06:37, 06:39-06:58, 07:01:01-07:03, 07:05-07:10, 07:14-07:27:02, 07:30-07:33N, 07:35-07:40, 07:42-07:44, 07:46-07:62, 07:64-07:100, 07:102-07:138, 07:140-07:141, 07:143-07:171, 08:09, 08:11, 12:02:01-12:14:02, 12:16-12:30, 12:32-12:51, 14:02:01-14:25, 14:27-14:29, 16:01:01-16:02:06, 16:04:01, 16:06-16:09, 16:11-16:32 |
| 55,6 | 100 bp | 1070 bp | *15:05:01-15:05:05, 15:22-15:23, 15:29, 15:36, 15:46 |  |
| 610 | 305 bp, 345 bp, 370 bp | 1070 bp | *15:06:01-15:06:03, 15:22, 15:26, 15:37, 15:40 |  |
| 7 | 140 bp | 800 bp | *15:07, 15:21, 15:25 | *01:21, 02:12, 02:27:01-02:27:02, 04:11, 04:29, 04:36, 04:55, 07:02:09, 08:01:01-08:09, 08:11-08:46, 12:02:01-12:03:03, 12:03:05-12:03:08, 12:03:10-12:03:15, 12:06-12:08, 12:10:01-12:20, 12:22-12:26, 12:28-12:32, 12:34-12:40, 12:42Q-12:51, 14:02:03, 14:03, 14:08, 14:10, 14:22, 16:01:01, 16:01:03-16:01:06, 16:04:01, 16:06-16:08, 16:10-16:11, 16:13-16:18, 16:20-16:24, 16:26-16:32, B*35:08:02, B*67:02 |
| 811 | 160 bp, 185 bp | 1070 bp | *15:08, 15:19 | *02:06, 02:47, 12:15, B*07:78w, B*13:18w, B*13:31w, B*13:41w, B*15:73w, B*54:10w, 54:20w, B*55:09w , B*55:21w, B*55:37 |
| 912 | 135 bp, 305 bp | 1070 bp | *15:11, 15:18, 15:23 | *02:02:01-02:02:03, 02:02:05-02:02:12, 02:02:14-02:20, 02:22-02:25Q, 02:27:01-02:38N, 02:40, 02:42-02:44, 02:46-02:47, 04:03, 04:06, 04:42, 04:80, 05:26, 05:43, 06:05, 07:02:09, 08:37, 12:16, 16:21 |
| 10 | 170 bp | 1070 bp | *15:10:01-15:10:02 | *02:08, 03:18, 03:64, 04:01:01:01-04:01:02, 04:01:04-04:01:26, 04:04:01-04:05, 04:07-04:15:03, 04:17-04:20, 04:23-04:79, 04:81-04:84, 05:01:01:01-05:01:15, 05:03-05:53, 06:28, 07:64, 07:73, 07:92, 08:01:01-08:13, 08:15:01-08:46, 12:02:01-12:15, 12:17-12:51, 14:02:01-14:02:07, 14:04-14:09, 14:11-14:17, 14:19-14:21N, 14:23-14:29, 16:01:01-16:02:06, 16:04:01, 16:06-16:32, 17:01:01:01-17:08, 18:04 |
| 11 | 315 bp | 800 bp | *15:02:01-15:03, 15:07-15:08, 15:10:01-15:13, 15:15-15:18, 15:21, 15:26, 15:28, 15:31-15:35, 15:37-15:39, 15:41-15:45, 15:47-15:48 | *02:06, 02:47, 03:12, 03:19, 03:102, 12:15 |
| 125,13 | 100 bp, 240 bp | 1070 bp | *15:07, 15:21, 15:25, 15:28, 15:43 | *02:12, 02:27:01-02:27:02, 03:02:01-03:04:16, 03:04:18-03:06, 03:08-03:14, 03:16-03:17, 03:19-03:38:02, 03:40-03:44, 03:46-03:66, 03:68-03:98, 03:100-03:108, 07:96 |
| 135,7,14 | 125 bp, 185 bp | 800 bp | *15:12, 15:24 | *04:52, 04:55, 05:47, 14:10 |
| 1415 | 130 bp, 440 bp | 800 bp | *15:11, 15:13, 15:16-15:17, 15:43 | *02:06, 02:47, 03:12, 03:19, 03:102, 12:15, B*46:11, B*46:18, B*56:08, B*56:14 |
| 155,6 | 85 bp | 1070 bp | *15:02:01-15:06:03, 15:08-15:10:02, 15:12-15:13, 15:15, 15:18-15:19, 15:21-15:24, 15:26, 15:28-15:42, 15:44-15:47 | *03:29, 05:36, 06:44, 07:07, 07:09 |
| 165,16 | 90 bp, 165 bp, 345 bp | 1070 bp | *15:15, 15:26-15:27 |  |
| 178,17 | 140 bp, 215 bp, 295 bp | 1070 bp | *15:34, 15:36, 15:39 | *01:30 |
| 185,7,18 | 85 bp, 160 bp | 1070 bp | *15:31, 15:33 | *07:123, B*15:200, B*58:05 |
| 197,19 | 165 bp, 320 bp, 355 bp | 1070 bp | *15:30, 15:42, 15:46 |  |
| 205,20 | 120 bp, 235 bp | 1070 bp | *15:44-15:45 |  |
| 2121 | 170 bp, 445 bp | 1070 bp | *15:35, 15:47 | *02:35, 05:21 |
| 22 | 225 bp | 800 bp | *15:38 |  |
| 236,22 | 175 bp, 545 bp | 1070 bp | *15:29, 15:48 | *08:22 |
| 2423 | 175 bp, 380 bp | 1070 bp | *15:32Q, 15:41 |  |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of C*15 SSP typings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective lengths of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most wells, or a band of 800 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-C*15 SSP subtyping.

In addition, wells number 7, 11, 13, 14 and 22 contain the primer pair giving rise to the shorter, 800 bp, internal positive control band.

In the presence of a specific amplification the intensity of the control band often decreases.

3The C*15:31 and C*15:33 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 18.

The C*15:32Q and C*15:41 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 24.

The C*15:34 and C*15:39 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 17.

The C*15:35 and C*15:47 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 21.

The C*15:44 and C*15:45 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 20.

4Due to the sharing of sequence motifs between HLA Class I alleles non-HLA-C*15 alleles will be amplified by most primer mixes. In addition, primer mix 7 amplifies the B*35:08:02 and B*67:02 alleles, primer mix 8 weakly amplifies the B*07:78, B*13:18, B*13:31, B*13:41, B*15:73, B*54:10, B*54:20, B*55:09, B*55:21 and B*55:37 alleles, primer mix 14 amplifies the B*46:11, B*46:18, B*56:08 and B*56:14 alleles and primer mix 18 amplifies the B*15:200 and B*58:05 alleles.

5Short specific PCR fragments are less intense and not as sharp as longer specific bands.

6Primer mixes 5, 15 and 23 may give a lower yield of HLA-specific PCR products than the other HLA-C*15 primer mixes.

7Primer mixes 13, 18 and 19 may have tendencies of primer oligomer formation.

8Primer mix 17 may have a tendency of giving rise to nonspecific amplifications.

9Primer mix 3: Specific PCR fragment of 70 bp in the C*15:28 allele. Specific PCR fragment of 195 bp in the C*15:03, 15:16 and 15:25w and in the C*02:12w, 03:15, 03:27w, 03:38:01w-03:38:02w, 03:69w, 04:03, 04:06, 04:16, 04:80, 07:20w, 07:96w and 07:127w alleles.

10Primer mix 6: Specific PCR fragment of 305 bp in the C*15:06:01-15:06:03 and C*15:40 alleles. Specific PCR fragment of 345 bp in the C*15:26 allele. Specific PCR fragment of 370 bp in the C*15:22 and 15:37 alleles.

11Primer mix 8: Specific PCR fragment of 160 bp in the C*15:08 and the C*02:06, 02:47, 12:15, B*07:78w, B*13:18w, 13:31w and 13:41w and in the B*15:73w, B*54:10w, B*54:20w, B*55:09w, B*55:21w and B*55:37w  alleles. Specific PCR fragment of 185 bp in the C*15:19 allele.

12Primer mix 9: Specific PCR fragment of 135 bp in the C*15:11 and 15:23 and the C*02:02:01-02:02:03, 02:02:05-02:02:12, 02:02:14-02:20, 02:22-02:25Q, 02:27:01-02:38N, 02:40, 02:42-02:44, 02:46-02:47, 04:03, 04:06, 04:42, 04:80, 05:26, 05:43, 06:05, 07:02:09, 08:37, 12:16 and 16:21 alleles. Specific PCR fragment of 305 bp in the C*15:18 allele.

13Primer mix 12: Specific PCR fragment of 100 bp in the C*15:28 allele. Specific PCR fragment of 240 bp in the C*15:07, 15:21, 15:25 and 15:43 and the C*02:12, 02:27:01-02:27:02, 03:02:01-03:04:16, 03:04:18-03:06, 03:08-03:14, 03:16-03:17, 03:19-03:38:02, 03:40-03:44, 03:46-03:66, 03:68-03:98, 03:100-03:108 and 07:96 alleles.

14Primer mix 13: Specific PCR fragment of 125 bp in the C*15:24 allele. Specific PCR fragment of 185 bp in the C*15:12 and C*04:52, 04:55 and 14:10 alleles.

15Primer mix 14: Specific PCR fragment of 130 bp in the C*15:13 allele. Specific PCR fragment of 440 bp the C*15:11, 15:16-15:17 and 15:43 and the C*02:06, 02:47, 03:12, 03:19, 03:102 and 12:15 and in the B*46:11, B*46:18, B*56:08 and B*56:14 alleles.

16Primer mix 16: Specific PCR fragment of 90 bp in the C*15:27 allele. Specific PCR fragment of 165 bp in the C*15:15 allele. Specific PCR fragment of 345 bp in the C*15:26 allele.

17Primer mix 17: Specific PCR fragment of 140 bp in the C*15:34 allele. Specific PCR fragment of 215 bp in the C *15:36 allele. Specific PCR fragment of 295 bp in the C*15:39 and the C*01:30 alleles.

18Primer mix 18: Specific PCR fragment of 85 bp in the C*15:33 allele. Specific PCR fragment of 160 bp in the C*15:31 and the C*07:123 and in the B*15:200 and B*58:05 alleles.

19Primer mix 19: Specific PCR fragment of 165 bp in the C*15:42 allele. Specific PCR fragment of 320 bp in the C*15:30 allele. Specific PCR fragment of 355 bp in the C*15:46 allele.

20Primer mix 20: Specific PCR fragment of 120 bp in the C*15:44 allele. Specific PCR fragment of 235 bp in the C*15:45 allele.

21Primer mix 21: Specific PCR fragment of 170 bp in the C*15:47 allele. Specific PCR fragment of 445 bp in the C*15:35 and the C*02:35 and 05:21 alleles.

22Primer mix 23: Specific PCR fragment of 175 bp in the C*15:48 allele. Specific PCR fragment of 545 bp in the C*15:29 and the C*08:22 alleles.

23Primer mix 24: Specific PCR fragment of 175 bp in the C*15:32Q allele. Specific PCR fragment of 380 bp in the C*15:41 allele.

‘w’, might be weakly amplified.

1The HLA-C*15:01 nucleotide sequence has been shown to be identical to C*15:02.

The HLA-C*15:14 nucleotide sequence has been renamed C*15:10:02.

2The C*15:31 and C*15:33 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 18.

3The C*15:32Q and C*15:41 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 24.

4The C*15:34 and C*15:39 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 17.

5The C*15:35 and C*15:47 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 21.

6The C*15:44 and C*15:45 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 20.

7Primer mix 3: Specific PCR fragment of 70 bp in the C*15:28 allele. Specific PCR fragment of 195 bp in the C*15:03, 15:16 and 15:25w and in the C*02:12w, 03:15, 03:27w, 03:38:01w-03:38:02w, 03:69w, 04:03, 04:06, 04:16, 04:80, 07:20w, 07:96w and 07:127w alleles.

Primer mix 6: Specific PCR fragment of 305 bp in the C*15:06:01-15:06:03 and C*15:40 alleles. Specific PCR fragment of 345 bp in the C*15:26 allele. Specific PCR fragment of 370 bp in the C*15:22 and 15:37 alleles.

Primer mix 8: Specific PCR fragment of 160 bp in the C*15:08 and the C*02:06, 02:47, 12:15, B*07:78w, B*13:18w, 13:31w and 13:41w and in the B*15:73w, B*54:10w, B*54:20w, B*55:09w, B*55:21w and B*55:37w  alleles. Specific PCR fragment of 185 bp in the C*15:19 allele.

Primer mix 9: Specific PCR fragment of 135 bp in the C*15:11 and 15:23 and the C*02:02:01-02:02:03, 02:02:05-02:02:12, 02:02:14-02:20, 02:22-02:25Q, 02:27:01-02:38N, 02:40, 02:42-02:44, 02:46-02:47, 04:03, 04:06, 04:42, 04:80, 05:26, 05:43, 06:05, 07:02:09, 08:37, 12:16 and 16:21 alleles. Specific PCR fragment of 305 bp in the C*15:18 allele.

Primer mix 12: Specific PCR fragment of 100 bp in the C*15:28 allele. Specific PCR fragment of 240 bp in the C*15:07, 15:21, 15:25 and 15:43 and the C*02:12, 02:27:01-02:27:02, 03:02:01-03:04:16, 03:04:18-03:06, 03:08-03:14, 03:16-03:17, 03:19-03:38:02, 03:40-03:44, 03:46-03:66, 03:68-03:98, 03:100-03:108 and 07:96 alleles.

Primer mix 13: Specific PCR fragment of 125 bp in the C*15:24 allele. Specific PCR fragment of 185 bp in the C*15:12 and C*04:52, 04:55 and 14:10 alleles.

Primer mix 14: Specific PCR fragment of 130 bp in the C*15:13 allele. Specific PCR fragment of 440 bp the C*15:11, 15:16-15:17 and 15:43 and the C*02:06, 02:47, 03:12, 03:19, 03:102 and 12:15 and in the B*46:11, B*46:18, B*56:08 and B*56:14 alleles.

Primer mix 16: Specific PCR fragment of 90 bp in the C*15:27 allele. Specific PCR fragment of 165 bp in the C*15:15 allele. Specific PCR fragment of 345 bp in the C*15:26 allele.

Primer mix 17: Specific PCR fragment of 140 bp in the C*15:34 allele. Specific PCR fragment of 215 bp in the C *15:36 allele. Specific PCR fragment of 295 bp in the C*15:39 and the C*01:30 alleles.

Primer mix 18: Specific PCR fragment of 85 bp in the C*15:33 allele. Specific PCR fragment of 160 bp in the C*15:31 and the C*07:123 and in the B*15:200 and B*58:05 alleles.

Primer mix 19: Specific PCR fragment of 165 bp in the C*15:42 allele. Specific PCR fragment of 320 bp in the C*15:30 allele. Specific PCR fragment of 355 bp in the C*15:46 allele.

Primer mix 20: Specific PCR fragment of 120 bp in the C*15:44 allele. Specific PCR fragment of 235 bp in the C*15:45 allele.

Primer mix 21: Specific PCR fragment of 170 bp in the C*15:47 allele. Specific PCR fragment of 445 bp in the C*15:35 and the C*02:35 and 05:21 alleles.

Primer mix 23: Specific PCR fragment of 175 bp in the C*15:48 allele. Specific PCR fragment of 545 bp in the C*15:29 and the C*08:22 alleles.

Primer mix 24: Specific PCR fragment of 175 bp in the C*15:32Q allele. Specific PCR fragment of 380 bp in the C*15:41 allele.

‘w’, might be weakly amplified.