PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® HLA-Cw*01

Prod. No:	101.621-12/12u

Lot No:    	03H

Expiry Date:    	2012-March-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 24 primer mixes used for HLA-Cw*01 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified HLA-Cw*01 alleles3 | Other amplified HLA Class I alleles4 |
| --- | --- | --- | --- | --- |
| 15 | 90 bp | 800 bp | *010201-010209, 0106-0111, 0113-0120, 0122, 0123, 0125-0133 |  |
| 25,7 | 90 bp, 270 bp | 1070 bp | *0103, 0115, 0124 | *0358, 0437 |
| 38 | 150 bp, 265 bp | 800 bp | *0104, 0114 |  |
| 46,9 | 210 bp, 240 bp | 1070 bp | *0105, 0116 | *0737, B*0777 |
| 510 | 150 bp, 200 bp | 800 bp | *0106, 0120 |  |
| 6 | 195 bp | 1070 bp | *0107 |  |
| 76,11 | 150 bp, 195 bp | 1070 bp | *0108, 0120 |  |
| 8 | 210 bp | 1070 bp | *0104, 0109, 0122 | *0623 |
| 9 | 230 bp | 800 bp | *0110 |  |
| 1012 | 210 bp, 290 bp | 1070 bp | *0111, 0122 |  |
| 11 | 345 bp | 800 bp | *0112, 0134 |  |
| 12 | 155 bp | 1070 bp | *0113 | *0509, 0517, 0815, B*1533 |
| 13 | 155 bp | 800 bp | *010201-0120, 0122-0134 | B*5418 |
| 146 | 240 bp | 1070 bp | *0117, 0121, 0123 |  |
| 15 | 230 bp | 800 bp | *0118 |  |
| 1613 | 255 bp, 295 bp | 1070 bp | *0119, 0123 | A*0124 |
| 175 | 75 bp | 800 bp | *0124, 0125 |  |
| 186,14 | 195 bp, 260 bp | 800 bp | *0126, 0134 |  |
| 195 | 100 bp | 1070 bp | *0127 |  |
| 205 | 110 bp | 800 bp | *0128 | *0359 |
| 215,15 | 125 bp, 245 bp | 800 bp | *0129, 0133 |  |
| 226 | 255 bp | 1070 bp | *0130 |  |
| 23 | 235 bp | 1070 bp | *0131 |  |
| 24 | 250 bp | 1070 bp | *0132 |  |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-Cw*01 SSP typings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective lengths of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may  be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most wells, or a band of 800 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-Cw*01 SSP subtyping.

In addition, wells number 3, 5, 9, 11, 13, 15, 17, 18, 20 and 21 contain the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3The Cw*0129 and Cw*0133 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 21.

4Due to the sharing of sequence motifs between HLA Class I alleles some non-HLA-Cw*01 alleles will be amplified by primer mixes 2, 4, 8, 12 and 20. In addition, the B*0777 allele will be amplified by primer mix 4, the B*1533 allele will be amplified by primer mix 12, the B*5418 allele will be amplified by primer mix 13 and the A*0124 allele will be amplified by primer mix 16.

5Short specific PCR fragments are less intense and not as sharp as longer specific bands.

6Primer mixes 4, 7, 14, 18 and 22 may have a tendency of giving rise to nonspecific amplifications.

7Primer mix 2: Specific PCR product of 90 bp in the Cw*0103 and Cw*0124 and in the Cw*0358 and Cw*0437 alleles. Specific PCR product of 270 bp in the Cw*0115 allele.

8Primer mix 3: Specific PCR product of 150 bp in the Cw*0104 allele. Specific PCR product of 265 bp in the Cw*0114 allele.

9Primer mix 4: Specific PCR product of 210 bp in the Cw*0105 and Cw*0737 and the B*0777 alleles. Specific PCR product of 240 bp in the Cw*0116 allele.

10Primer mix 5: Specific PCR product of 150 bp in the Cw*0120 allele. Specific PCR product of 200 bp in the Cw*0106 allele.

11Primer mix 7: Specific PCR product of 150 bp in the Cw*0120 allele. Specific PCR product of 195 bp in the Cw*0108 allele.

12Primer mix 10: Specific PCR fragment of 210 bp in the Cw*0122 allele. Specific PCR fragments of 290 bp in the Cw*0111 allele.

13Primer mix 16: Specific PCR product of 255 bp in the Cw*0119 allele. Specific PCR product of 295 bp in the Cw*0123 and the A*0124 alleles.

14Primer mix18: Specific PCR product of 195 bp in the Cw*0126 allele. Specific PCR product of 260 bp in the Cw*0134 allele.

15Primer mix 21: Specific PCR fragment of 125 bp in the Cw*0133 allele. Specific PCR fragments of 245 bp in the Cw*0129 allele.

1The HLA-Cw*0101 nucleotide sequence has been shown to be identical to Cw*0102.

2The Cw*0129 and Cw*0133 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 21.

3Primer mix 2: Specific PCR product of 90 bp in the Cw*0103 and Cw*0124 and in the Cw*0358 and Cw*0437 alleles. Specific PCR product of 270 bp in the Cw*0115 allele.

Primer mix 3: Specific PCR product of 150 bp in the Cw*0104 allele. Specific PCR product of 265 bp in the Cw*0114 allele.

Primer mix 4: Specific PCR product of 210 bp in the Cw*0105 and Cw*0737 and the B*0777 alleles. Specific PCR product of 240 bp in the Cw*0116 allele.

Primer mix 5: Specific PCR product of 150 bp in the Cw*0120 allele. Specific PCR product of 200 bp in the Cw*0106 allele.

Primer mix 7: Specific PCR product of 150 bp in the Cw*0120 allele. Specific PCR product of 195 bp in the Cw*0108 allele.

Primer mix 10: Specific PCR fragment of 210 bp in the Cw*0122 allele. Specific PCR fragments of 290 bp in the Cw*0111 allele.

Primer mix 16: Specific PCR product of 255 bp in the Cw*0119 allele. Specific PCR product of 295 bp in the Cw*0123 and the A*0124 alleles.

Primer mix18: Specific PCR product of 195 bp in the Cw*0126 allele. Specific PCR product of 260 bp in the Cw*0134 allele.

Primer mix 21: Specific PCR fragment of 125 bp in the Cw*0133 allele. Specific PCR fragments of 245 bp in the Cw*0129 allele.