PHOTO DOCUMENT

Olerup SSP AB

Olerup SSPTM HLA-Cw*03

Prod. No:	101.611-12 – including Taq pol.

101.611-12u– without Taq pol.

Lot No:    	Y28

Expiry Date:    	2009-September-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ug/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 31 primer mixes used for HLA-Cw*03 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified HLA-Cw*03 alleles3 | Other amplified HLA Class I alleles4 |
| --- | --- | --- | --- | --- |
| 15 | 280 bp | 800 bp | 030201-030203, 030401-0310, 0314-0317, 0319, 0323-0329, 0332-033802, 0340-0342, 0344-0346 | 020201-020203, 020205-0203, 0204weakly, 0205-0213, 0214weakly, 0215-0220, 0403, 0406, 0416, 0603, 150201-1511, 1513, 1515-1521 |
| 2 | 210 bp | 1070 bp | 030201-030203, 0314, 0315, 0333, 0336, 0340, 0342, 0343 |  |
| 3 | 280 bp | 800 bp | 030301-030305, 031101-0313, 0320N-0322Q, 0330, 0331, 0343 | 1512 |
| 4 | 275 bp | 1070 bp | 030301-030407, 0306-0312, 0314, 0318-0324, 0326, 0328-0332, 0334, 0337-0340, 0342-0346 |  |
| 5 | 275 bp | 1070 bp | 0305, 0313, 0317, 0325, 0327, 0335 | B*1596 |
| 66,9 | 90, 215 bp | 800 bp | 0306, 0321 | B*1596 |
| 76,10 | 110, 185 bp | 1070 bp | 030301-030305, 0308, 031101-0313, 0318, 0320N-0322Q, 0329-0331, 0339, 0343 | 070101-070109, 0706, 0707, 0709, 0716, 0718-0720, 0722, 0724, 0726, 0728, 0730, 0735, 0736, 0740, 0744, 0752, 0753 |
| 8 | 250 bp | 1070 bp | 0309, 0322Q, 0342 |  |
| 9 | 240 bp | 1070 bp | 0307, 0315, 0345 | 020201-020203, 020205-0203, 0204weakly, 0205-0211, 0213, 0214weakly, 0215-0220, 0403, 0406, 0416, 0603, 150201-1506, 1508-1513, 1515-1520 |
| 107 | 250 bp | 800 bp | 0307, 0310, 0315, 0329, 0345 | 020201weakly, 020202-020203, 020205-0203, 0204weakly, 0205-0213, 0214weakly, 0215-0220, 0403, 0406, 0416, 0603, 150201-1506, 1508-1513, 1515-1521 |
| 116,7 | 415 bp | 1070 bp | 0308, 0329, 0331 | 1208 |
| 12 | 260 bp | 1070 bp | 031101, 0318, 0339 | 1512 |
| 136,11 | 80, 140 bp | 1070 bp | 030201-030203, 0315-0317, 0333, 0336, 0340, 0342, 0343 | 140201, 140203-1409 |
| 146 | 210 bp | 1070 bp | 0312, 0319 |  |
| 156 | 135 bp | 1070 bp | 030201-0317, 0319-033802, 0340-0346 | 0603weakly |
| 16 | 150 bp | 1070 bp | 0307, 0310, 0315, 0329, 0345 | 0114, 06020101-0604, 0606-0610, 0612-0616N, 0707, 0709, 120401,1701-1704, 1801-1803 |
| 176,7 | 90 bp | 1070 bp | 0320N |  |
| 18 | 225 bp | 1070 bp | 0323 |  |
| 196,12 | 105, 260 bp | 800 bp | 0324, 0333 |  |
| 20 | 165 bp | 1070 bp | 0325 |  |
| 21 | 210 bp | 1070 bp | 0326 |  |
| 225 | 195 bp | 800 bp | 0315, 0327, 033801-033802 | 0212, 0403, 0406, 0416, 0720, 1503, 1516 |
| 236 | 90 bp | 1070 bp | 0328 |  |
| 24 | 200 bp | 1070 bp | 0330 |  |
| 2513 | 295, 345 bp | 1070 bp | 0332, 033801-033802, 0345 | 0614, 0710, 0743, 1503, 1516 |
| 266,14 | 70, 105, 265 bp | 1070 bp | 0313, 0334-0336 | B*1596 |
| 276 | 80 bp | 1070 bp | 031101-031102, 0337 | 0603 |
| 288 | 170 bp | 1070 bp | 0316, 0341 | 04010101-040104, 0403-0420, 0423-0431, 140201, 140203-1409, 1801-1803 |
| 29 | 285 bp | 1070 bp | 0339 |  |
| 30 | 165 bp | 1070 bp | 0346 |  |
| 31 | 210 bp | 800 bp | 0344 |  |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-Cw*03 SSP subtypings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 bp or more. Size differences shorter than 20 bp are not given. For high resolution SSP kits the length of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control band may sometimes be observed. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most tubes, or a band of 800 base pairs, for some tubes.

Tube number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-Cw*03 subtyping.

In addition, tubes number 3, 6, 11, 19, 22 and 31 contain the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

4Due to the sharing of sequence motifs between HLA-Cw alleles the primer pairs in tubes 1, 3, 7, 9 to 13, 15, 16 and 22 will amplify some non-HLA-Cw*03 alleles. In addition, the B*1596 allele will be amplified by primer mixes 5, 6 and 26

5The primer pairs in tubes 1 and 22 may give rise to less specific PCR product than the other HLA-Cw*03 primer pairs.

6Short specific PCR fragments are less intense and not as sharp as longer specific bands.

7Primer mixes 10, 11 and 17 may give rise to primer oligomer formation, most pronounced for primer mix 11.

8Primer mix 28 may give rise to nonspecific amplifications.

9Primer mix 6: Specific PCR fragment of 90 bp in the Cw*0321 and B*1596 alleles. Specific PCR fragment of 215 bp in the Cw*0306 alleles.

10Primer mix 7: Specific PCR fragment of 110 bp in the Cw*0308, 0329 and 0331 and the Cw*070101-070109, 0706, 0707, 0709, 0716, 0718-0720, 0722, 0724, 0726, 0728, 0730, 0735, 0736, 0740, 0744, 0752 and 0753 alleles. Specific PCR fragment of 185 bp in the Cw*030301-030305, 031101-0313, 0318, 0320N to 0322Q, 0329 to 0331, 0339 and 0343 alleles.

11Primer mix 13: Specific PCR fragment of 80 bp in the Cw*030201-030203, 0315, 0316, 0333, 0336, 0340, 0342 and 0343 and the Cw*140201, 140203 to 1409 alleles. Specific PCR fragment of 140 bp in the Cw*0317 allele.

12Primer mix 19: Specific PCR fragment of 105 bp in the Cw*0324 allele. Specific PCR fragment of 260 bp in the Cw*0333 allele.

13Primer mix 25: Specific PCR fragment of 295 bp in the Cw*0332 and 0345 and the Cw*0614, 0710 and 0743 alleles. Specific PCR fragment of 345 bp in the Cw*033801-033802 and the Cw*1503 and 1516 alleles.

14Primer mix 26: Specific PCR fragment of 70 bp in the Cw*0334 allele. Specific PCR fragment of 105 bp in the Cw*0336 allele. Specific PCR fragment of 265 bp in the Cw*0313 and 0335 and the B*1596 alleles.

1The nucleotide sequence of the HLA-Cw*0301 allele has been shown to be identical to Cw*0304.

2Primer mix 6: Specific PCR fragment of 90 bp in the Cw*0321 and B*1596 alleles. Specific PCR fragment of 215 bp in the Cw*0306 alleles.

Primer mix 7: Specific PCR fragment of 110 bp in the Cw*0308, 0329 and 0331 and the Cw*070101-070109, 0706, 0707, 0709, 0716, 0718-0720, 0722, 0724, 0726, 0728, 0730, 0735, 0736, 0740, 0744, 0752 and 0753 alleles. Specific PCR fragment of 185 bp in the Cw*030301-030305, 031101-0313, 0318, 0320N to 0322Q, 0329 to 0331, 0339 and 0343 alleles.

Primer mix 13: Specific PCR fragment of 80 bp in the Cw*030201-030203, 0315, 0316, 0333, 0336, 0340, 0342 and 0343 and the Cw*140201, 140203 to 1409 alleles. Specific PCR fragment of 140 bp in the Cw*0317 allele.

Primer mix 19: Specific PCR fragment of 105 bp in the Cw*0324 allele. Specific PCR fragment of 260 bp in the Cw*0333 allele.

Primer mix 25: Specific PCR fragment of 295 bp in the Cw*0332 and 0345 and the Cw*0614, 0710 and 0743 alleles. Specific PCR fragment of 345 bp in the Cw*033801-033802 and the Cw*1503 and 1516 alleles.

Primer mix 26: Specific PCR fragment of 70 bp in the Cw*0334 allele. Specific PCR fragment of 105 bp in the Cw*0336 allele. Specific PCR fragment of 265 bp in the Cw*0313 and 0335 and the B*1596 alleles.

‘w’, might be weakly amplified.