PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® HLA-Cw*04

Prod. No:	101.612-12/12u

Lot No:    	72F

Expiry Date:    	2011-April-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 24 primer mixes used for HLA-Cw*04 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified HLA-Cw*04 alleles | Other amplified HLA Class I alleles3 |
| --- | --- | --- | --- | --- |
| 1 | 250 bp | 800 bp | 04010101-040106, 0403, 0405, 0407-0412, 041501-0418, 0419weakly, 0420, 0423-0433, 0435-0437 | 010201-0121, 020201-020203, 020205-020206, 0204-0215, 0217, 0219-0223, 050101-050104, 0503-0510, 0512-0516, 0518-0523, 06020101-06020102, 060203, 0603, 0605-0619, 0802, 0805, 0807, 0812, 0815, 0817-0819, 120201-1213, 1215-1217, 1221, 140201-1405, 1407N-1411, 1508, 160401, 1801-1803 |
| 2 | 220 bp | 1070 bp | 04010101-040106, 040401-0405, 0407-0409N, 0412weakly, 0413-041502, 0417-0420, 0423-0435, 0437 | 0764 |
| 34 | 145 bp | 1070 bp | 04010101-040106, 040401-0405, 0408-041502, 0417-0420, 0423-0426, 0428-0437 |  |
| 4 | 210 bp | 1070 bp | 0403, 0406 | 0212 |
| 55 | 250 bp | 1070 bp | 040401-040402, 0406, 0413, 0434 | 0122, 0203, 0216, 0218, 0511, 0517, 0604, 080101-080102, 0803, 0804, 0806, 0808-0811, 0813, 0814, 0816, 1214, 1218, 1220, 1406, 150201-1507, 1509-1513, 1515-1521, 1701-1705, B*5802 |
| 64 | 95 bp | 1070 bp | 0405 |  |
| 74 | 145 bp | 1070 bp | 0407, 0427, 0432 | 0334, 0764, 1803 |
| 85 | 270 bp | 1070 bp | 0408, 0434 | 1515, B*5802 |
| 94,5,6 | 110, 180, 220 bp | 1070 bp | 0409N, 0430, 0435 |  |
| 107 | 195, 220 bp | 1070 bp | 0410, 0411, 0431, 0436 |  |
| 11 | 180 bp | 1070 bp | 04010101-040106, 0403-0410, 0412-0420, 0423-0426, 0428-0432, 0434-0437 | 0605 |
| 124,8 | 125, 180 bp | 1070 bp | 0411, 0429, 0433, 0436 |  |
| 13 | 225 bp | 1070 bp | 0412 |  |
| 149 | 155, 190 bp | 1070 bp | 0413, 0416, 0418 | 0708 |
| 156,10 | 170, 190 bp | 1070 bp | 0414, 0428 |  |
| 164 | 130 bp | 1070 bp | 041501, 0437 | 0305, 0325, 0327, 1409 |
| 174,6 | 90 bp | 800 bp | 0417 |  |
| 18 | 225 bp | 1070 bp | 0419 |  |
| 194,11 | 120, 155 bp | 1070 bp | 0420, 0435, 0437 |  |
| 20 | 540 bp | 1070 bp | 041502, 0417 |  |
| 214 | 85 bp | 1070 bp | 0423 |  |
| 224 | 130 bp | 1070 bp | 0424 |  |
| 234 | 85 bp | 1070 bp | 0425 |  |
| 24 | 180 bp | 1070 bp | 0426, 0430 |  |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-Cw*04 SSP typings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective lengths of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may  be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most wells, or a band of 800 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-Cw*04 subtyping.

In addition, well number 17 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3Due to the sharing of sequence motifs between HLA-Cw alleles some non-HLA-Cw*04 alleles will be amplified by primer mixes 1, 2, 4, 5, 7, 8, 11, 14 and 16. In addition, HLA-B*5802 will be amplified by primer mixes 5 and 8.

4Short specific PCR fragments are less intense and not as sharp as longer specific bands.

5The primers in primer mix 5, 8 and 9 may have a tendency of giving rise to nonspecific amplifications. The specific PCR product in primer mix 9 may also be less intense than the other Cw*04 specific PCR products.

6Primer mix 9: Specific PCR fragment of 110 bp in the Cw*0435 allele. Specific PCR fragment of 180 bp in the Cw*0430 allele. Specific PCR fragment of 220 bp in the Cw*0409N allele.

7Primer mix 10: Specific PCR fragmentof 195 bp in the Cw*0431 allele. Specific PCR fragment of 220 bp in the Cw*0410, Cw*0411 and Cw*0436 alleles.

8Primer mix 12: Specific PCR fragment of 125 bp in the Cw*0429 and Cw*0436 alleles. Specific PCR fragment of 180 bp in the Cw*0433 allele. Specific PCR fragment of 125 and 180 bp in the Cw*0411 allele.

9Primer mix 14: Specific PCR fragment of 155 bp in the Cw*0416 allele. Specific PCR fragment of 190 bp in the Cw*0413, Cw*0418 and Cw*0708 alleles.

10Primer mix 15: Specific PCR fragment of 170 bp in the Cw*0414 allele. Specific PCR fragment of 190 bp in the Cw*0428 allele.

11Primer mix 19: Specific PCR fragment of 120 bp in the Cw*0435 and Cw*0437 alleles. Specific PCR fragment of 155 bp in the Cw*0420 allele.

1The nucleotide sequence of the Cw*0402 allele has been shown to be identical to Cw*04010101.

The nucleotide sequence of the Cw*0421 allele has been renamed to Cw*041502.

The nucleotide sequence of the Cw*0422 allele has been shown to be identical to Cw*0421.

2Primer mix 9: Specific PCR fragment of 110 bp in the Cw*0435 allele. Specific PCR fragment of 180 bp in the Cw*0430 allele. Specific PCR fragment of 220 bp in the Cw*0409N allele.

Primer mix 10: Specific PCR fragmentof 195 bp in the Cw*0431 allele. Specific PCR fragment of 220 bp in the Cw*0410, Cw*0411 and Cw*0436 alleles.

Primer mix 12: Specific PCR fragment of 125 bp in the Cw*0429 and Cw*0436 alleles. Specific PCR fragment of 180 bp in the Cw*0433 allele. Specific PCR fragment of 125 and 180 bp in the Cw*0411 allele.

Primer mix 14: Specific PCR fragment of 155 bp in the Cw*0416 allele. Specific PCR fragment of 190 bp in the Cw*0413, Cw*0418 and Cw*0708 alleles.

Primer mix 15: Specific PCR fragment of 170 bp in the Cw*0414 allele. Specific PCR fragment of 190 bp in the Cw*0428 allele.

Primer mix 19: Specific PCR fragment of 120 bp in the Cw*0435 and Cw*0437 alleles. Specific PCR fragment of 155 bp in the Cw*0420 allele.

‘w’, may be weakly amplified.