PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® HLA-Cw*06

Prod. No:	101.614-12/12u

Lot No:    	95F

Expiry Date:    	2011-June-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 16 primer mixes used for HLA-Cw*06 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified HLA-Cw*06 alleles | Other amplified HLA Class I alleles3 |
| --- | --- | --- | --- | --- |
| 1 | 240 bp | 800 bp | 06020101-06020102, 060203, 0604-0616N, 0618-0620, 0622, 0623 | 0339, 1216 |
| 2 | 220 bp | 800 bp | 06020101-06020102, 060203, 0603, 0607- 0613, 0615-0620, 0622, 0623 | 0104, 0109, 0205, 0217, 12030101-1207, 1211-1213, 1215, 160401 |
| 34 | 135 bp | 1070 bp | 0603, 0618 | 030201-0317, 0319-033802, 0340-0350, 0357, 0358 |
| 4 | 250 bp | 1070 bp | 0604 | 0122, 0511, 0517, 080101-080102, 0803, 0804, 0806, 0808-0811, 0813, 0814, 0816, 0820-0822, 121401-121402, 1218, 1220, 1406, 150201-1507, 1509-1513, 1515-1521, 1701-1705, B*5802 |
| 56 | 165, 220 bp | 800 bp | 0605, 0618, 0623 | 0104, 0109, 0221, 05010101-050104, 0503-0525, 0810, 1221, 1705 |
| 65 | 250 bp | 1070 bp | 0605, 0606 | 010201-0103, 0105-0108, 0110-0120, 0123-0127, 04010101-040106, 0403, 0405, 0407-0412, 041501-0420, 0423-0433, 0435-0438, 0440, 050101-050104, 0503-0507N, 0509, 0510, 0512-0516, 0518-0525, 0802, 0805, 0807, 0812, 0815, 0817-0819, 0823, 1209, 140201-1405, 1407N-1411, 1508, 1801, 1802 |
| 74,7 | 110, 185 bp | 1070 bp | 0607, 0619 |  |
| 8 | 240 bp | 1070 bp | 0608 | 0110, 0205, 0217 |
| 98 | 165, 210, 435 bp | 1070 bp | 0609, 0617, 0623 | 0104, 0109, 0222, 0508, 0707, 0709, 1801-1803, B*9537 |
| 105 | 190 bp | 800 bp | 0610, 0622 |  |
| 114,9 | 130, 185 bp | 1070 bp | 0611, 0622 | 070401-070403, 0711, 0712, 0745, 0763, 0768 |
| 1210 | 200, 275 bp | 1070 bp | 0612, 0619, 0620 |  |
| 13 | 155 bp | 1070 bp | 0613 |  |
| 14 | 295 bp | 1070 bp | 0614 | 0332, 0345, 0710, 0743 |
| 15 | 360 bp | 800 bp | 0615 |  |
| 1611 | 235, 275 bp | 1070 bp | 0616N, 0620 |  |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-Cw*06 SSP typings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective lengths of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most wells, or a band of 800 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-Cw*06 subtyping.

In addition, wells number 2, 5, 10 and 15 contain the primer pair giving rise to the shorter, 800 bp, internal positive control in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3Due to the sharing of sequence motifs between HLA-Cw alleles non-HLA-Cw*06 alleles will be amplified by primer mixes 1 to 6, 8, 9, 11 and 14. In addition, the B*5802 allele will be amplified by primer mix 4 and the B*9537 allele will be amplified by primer mix 9.

4Short specific PCR fragments are less intense and not as sharp as longer specific bands.

5Primer mixes 6 and 10 have a tendency of giving rise to non-specific amplifications.

6Primer mix 5: Specific PCR fragment of 165 bp in the Cw*0605 and Cw*050101 to 050104, 0503 to 0525, 0810, 1221 and 1705 alleles. Specific PCR fragment of 220 bp in the Cw*0618 and 0623 and Cw*0104, 0109 and 0221 alleles.

7Primer mix 7: Specific PCR fragment of 110 bp in the Cw*0607 allele. Specific PCR fragment of 185 bp in the Cw*0619 allele.

8Primer mix 9: Specific PCR fragment of 165 bp in the Cw*0609 and Cw*0222, 0508 and B*9537 alleles. Specific PCR fragment of 210 bp in the Cw*0623 and Cw*0104 and 0109 alleles.

Specific PCR fragment of 435 bp in the Cw*0617 and Cw*0707, 0709, 1801 and 1802 alleles. Specific PCR fragments of 165 and 435 bp in the Cw*1803 allele.

9Primer mix 11: Specific PCR fragment of 130 bp in the Cw*0611 and Cw*070401 to 070403, 0711, 0712, 0745, 0763 and 0768 alleles. Specific PCR fragment of 185 bp in the Cw*0622 allele.

10Primer mix 12: Specific PCR fragment of 200 bp in the Cw*0612 and 0619 alleles. Specific PCR fragment of 275 bp in the Cw*0620 allele.

11Primer mix 16: Specific PCR fragment of 235 bp in the Cw*0616N allele. Specific PCR fragment of 275 bp in the Cw*0620 allele.

1The sequence of the Cw*0601 allele has been shown to be identical to Cw*0602.

The sequence of the Cw*060202 allele has been renamed to Cw*0617.

2Primer mix 5: Specific PCR fragment of 165 bp in the Cw*0605 and Cw*050101 to 050104, 0503 to 0525, 0810, 1221 and 1705 alleles. Specific PCR fragment of 220 bp in the Cw*0618 and 0623 and Cw*0104, 0109 and 0221 alleles.

Primer mix 7: Specific PCR fragment of 110 bp in the Cw*0607 allele. Specific PCR fragment of 185 bp in the Cw*0619 allele.

Primer mix 9: Specific PCR fragment of 165 bp in the Cw*0609 and Cw*0222, 0508 and B*9537 alleles. Specific PCR fragment of 210 bp in the Cw*0623 and Cw*0104 and 0109 alleles.

Specific PCR fragment of 435 bp in the Cw*0617 and Cw*0707, 0709, 1801 and 1802 alleles. Specific PCR fragments of 165 and 435 bp in the Cw*1803 allele.

Primer mix 11: Specific PCR fragment of 130 bp in the Cw*0611 and Cw*070401 to 070403, 0711, 0712, 0745, 0763 and 0768 alleles. Specific PCR fragment of 185 bp in the Cw*0622 allele.

Primer mix 12: Specific PCR fragment of 200 bp in the Cw*0612 and 0619 alleles. Specific PCR fragment of 275 bp in the Cw*0620 allele.

Primer mix 16: Specific PCR fragment of 235 bp in the Cw*0616N allele. Specific PCR fragment of 275 bp in the Cw*0620 allele.