PHOTO DOCUMENT

Olerup SSP AB

Olerup SSPTM HLA-Cw*07

Prod. No:	101.615-24/04 – including Taq pol.

101.615-24u/04u– without Taq pol.

Lot No:    	X93

Expiry Date:    	2009-June-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ug/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 32 primer mixes used for HLA-Cw*07 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified HLA-Cw*07 alleles3 | Other amplified HLA-Class I alleles4 |
| --- | --- | --- | --- | --- |
| 15 | 80 bp | 800 bp | 070101-0705, 0706weakly, 0707-0710?, 0712-0717?, 0718weakly, 0719weakly, 0720-0725?, 0726, 0727-0729?, 0730, 0731?, 0732N, 0733N-0735?, 0736, 0737-0741?, 0742, 0743-0748?, 0749 |  |
| 2 | 290 bp | 800 bp | 070101-070104, 070106-070107, 0705-0707, 0709, 0714, 0716, 0718, 0720-0722, 0724, 0726-0728, 0730, 0731, 0734-0736, 0740, 0743, 0744 | 0517 |
| 35 | 110 bp | 800 bp | 070101-070107, 0706, 0709, 0718-0722, 0724, 0726, 0728, 0730, 0735, 0736, 0740, 0744 |  |
| 4 | 185 bp | 1070 bp | 07020101-07020103, 0710, 0713, 0715, 0719, 0723, 0725, 0729, 0732N, 0733N, 0737-0739, 0742, 0746-0749 | B*4060 |
| 56 | 245 bp | 800 bp | 0703 |  |
| 65 | 100 bp | 1070 bp | 070401-070402, 0711, 0712 |  |
| 7 | 190 bp | 1070 bp | 0705 | B*2736 |
| 86,7,8 | 285, 645 bp | 1070 bp | 0706, 0745 | A*0284 |
| 96,9 | 240, 425 bp | 800 bp | 0707, 0709, 0729, 0749 | 1801-1803 |
| 10 | 185 bp | 1070 bp | 0708 |  |
| 1110 | 275, 305 bp | 800 bp | 0710, 0724, 0743 |  |
| 125 | 130 bp | 1070 bp | 0712, 0741 | 050101-050103, 0503-0508, 0510-0516, 080101-0806, 0808-0814, A*2910 |
| 135,7 | 80 bp | 1070 bp | 0711 |  |
| 14 | 240 bp | 800 bp | 0713, 0728 | 0517 |
| 15 | 290 bp | 800 bp | 0714 |  |
| 16 | 395 bp | 800 bp | 07020101-0705, 0708, 0710-0715, 0717, 0723, 0725, 0727, 0729, 0731-0734, 0737-0739, 0741-0743, 0745-0749 | 1801-1803 |
| 175 | 130 bp | 1070 bp | 070401-070402, 0711 | 0807 |
| 18 | 230 bp | 800 bp | 0715 | B*4060 |
| 1911 | 220, 465 bp | 800 bp | 0707, 0716, 0726, 0733N |  |
| 205,12 | 145, 170 bp | 1070 bp | 0703, 0717, 0744, 0748 | 0517 |
| 217 | 225 bp | 800 bp | 0706, 0718, 0719 |  |
| 226 | 250 bp | 800 bp | 0720 | 0315, 0332, 0403, 0406, 0416 |
| 2313 | 155, 185 bp | 1070 bp | 0721, 0723 |  |
| 245 | 130 bp | 1070 bp | 0722, 0725 |  |
| 255,14 | 105, 150 bp | 1070 bp | 0730, 0746 | B*5608weakly, B*5614weakly |
| 265,15 | 100, 300 bp | 1070 bp | 0731, 0732N |  |
| 275,16 | 120, 290 bp | 1070 bp | 070107, 0714, 0734, 0737 |  |
| 28 | 260 bp | 800 bp | 0735, 0747 |  |
| 29 | 255 bp | 1070 bp | 0736, 0742 |  |
| 30 | 275 bp | 1070 bp | 0738 | B*4060 |
| 31 | 415 bp | 1070 bp | 0739, 0740 |  |
| 32 | 245 bp | 1070 bp | 0707, 0709, 0749 | 0214, 06020101-060202, 0604-0616N, 1801-1803, B*0806, B*0817 |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-Cw*07 SSP subtypings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 bp or more. Size differences shorter than 20 bp are not given. For high resolution SSP kits the length of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers, e.g. the 3’-primer in primer mixes 13 and 21.

PCR fragments longer than the control band may sometimes be observed. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most tubes, or a band of 800 base pairs, for some tubes.

Tube number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-Cw*07 subtyping.

In addition, tubes number 2, 3, 5, 9, 11, 14, 15, 16, 18, 19, 21, 22 and 28 contain the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3For several HLA-Cw*07 alleles 1st, 5th, 6th and 7th exon nucleotide sequences are not available. In these instances it is not known whether some of the primers of the SSP sets are completely matched with the target sequences or not. We assume that unknown sequences in the 1st, 5th, 6th and 7th exons are conserved within allelic groups.

## 4Due to the sharing of sequence motifs between HLA-Cw alleles non-HLA-Cw*07 alleles will be amplified by primer mixes 2, 9, 12, 14, 16, 17, 20, 22 and 32. In addition, the A*0284 allele will be amplified by primer mix 8, the A*2910 allele will be amplified by primer mix 12, the B*0806 and B*0817 alleles will be amplified by primer mix 32, the B*2736 allele will be amplified by primer mix 7, the B*4060 allele by primer mixes 4, 18 and 30 and the B*5608 and B*5614 alleles will be weakly amplified by primer mix 25.

5Short specific PCR fragments are less intense and not as sharp as longer specific bands.

6Primer mixes 5, 8, 9 and 22 have a tendency of primer oligomer formation.

7Primer mixes 8, 13 and 21 have a tendency of giving rise to nonspecific amplifications.

8Primer mix 8: Specific PCR fragment of 285 bp in the Cw*0745 and A*0284 alleles. Specific PCR fragmentof 645 bp in the Cw*0706 allele.

9Primer mix 9: Specific PCR fragment of 240 bp in the Cw*0729 allele. Specific PCR fragment of 425 bp in the Cw*0707, 0709 and 0749 and the Cw*1801-1802 alleles.

10Primer mix 11: Specific PCR fragment of 275 bp in the Cw*0724 allele. Specific PCR fragment of 305 bp in the Cw*0710 and 0743 alleles.

11Primer mix 19: Specific PCR fragment of 220 bp in the Cw*0726 and 0733N alleles. Specific PCR fragment of 465 bp in the Cw*0707 and 0716 alleles.

12Primer mix 20: Specific PCR fragment of 145 bp in the Cw*0703 and 0717 and Cw*0517 alleles. Specific PCR fragment of 170 bp in the Cw*0744 and 0748 alleles.

13Primer mix 23:	 Specific PCR fragment of 155 bp in the Cw*0723 allele. Specific PCR fragment of 185 bp in the Cw*0721 allele.

14Primer mix 25: Specific PCR fragment of 105 bp in the Cw*0746 and B*5606 and 5614 alleles. Specific PCR fragment of 150 bp in the Cw*0730 allele.

15Primer mix 26: Specific PCR fragment of 100 bp in the Cw*0732N allele. Specific PCR fragment of 300 bp in the Cw*0731 allele.

16Primer mix 27: Specific PCR fragment of 120 bp in the Cw*0737 allele. Specific PCR fragment of 290 bp in the Cw*070107, 0714 and 0734 alleles.

‘?’, nucleotide sequence information not available for the primer matching sequence.

1Primer mix 8: Specific PCR fragment of 285 bp in the Cw*0745 and A*0284 alleles. Specific PCR fragmentof 645 bp in the Cw*0706 allele.

Primer mix 9: Specific PCR fragment of 240 bp in the Cw*0729 allele. Specific PCR fragment of 425 bp in the Cw*0707, 0709 and 0749 and the Cw*1801-1802 alleles.

Primer mix 11: Specific PCR fragment of 275 bp in the Cw*0724 allele. Specific PCR fragment of 305 bp in the Cw*0710 and 0743 alleles.

Primer mix 19: Specific PCR fragment of 220 bp in the Cw*0726 and 0733N alleles. Specific PCR fragment of 465 bp in the Cw*0707 and 0716 alleles.

Primer mix 20: Specific PCR fragment of 145 bp in the Cw*0703 and 0717 and Cw*0517 alleles. Specific PCR fragment of 170 bp in the Cw*0744 and 0748 alleles.

Primer mix 23:	 Specific PCR fragment of 155 bp in the Cw*0723 allele. Specific PCR fragment of 185 bp in the Cw*0721 allele.

Primer mix 25: Specific PCR fragment of 105 bp in the Cw*0746 and B*5606 and 5614 alleles. Specific PCR fragment of 150 bp in the Cw*0730 allele.

Primer mix 26: Specific PCR fragment of 100 bp in the Cw*0732N allele. Specific PCR fragment of 300 bp in the Cw*0731 allele.

Primer mix 27: Specific PCR fragment of 120 bp in the Cw*0737 allele. Specific PCR fragment of 290 bp in the Cw*070107, 0714 and 0734 alleles.

‘w’, might be weakly amplified.

‘?’, nucleotide sequence information not available for the primer matching sequence.