PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® HLA-Cw*12

Prod. No:	101.624-12/12u

Lot No:    	75G

Expiry Date:    	2012-January-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 24 primer mixes used for HLA-Cw*12 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified HLA-Cw*12 alleles | Other amplified HLA Class I alleles3 |
| --- | --- | --- | --- | --- |
| 15 | 235 bp | 800 bp | *120201-120402, 1206-1208, 1210-1215, 1217-1220, 1222-1228 | *0212, 0315, 0327, 033801, 033802, 0353, 0369, 0403, 0406, 0416, 0603, 0726, 0805, 0821, 0825, 1503, 1516, 1615, 17010101- 1705, B*0713, B*0715, B*6702 |
| 24,6 | 100 bp | 1070 bp | *120201-120203, 1208, 1210, 121402, 1216-1218, 1222, 1227 | *0104, 0121 |
| 3 | 220 bp | 800 bp | *12030101-1207, 1211-1213, 1215, 1223, 1225, 1226, 1228 | *0104, 0109, 0205, 0217, 06020101, 06020102, 060203-0603, 0607-0613, 0615-0628, 160401 |
| 4 | 340 bp | 1070 bp | *120401-1205, 1209, 1221 | *0114, 020201-020203, 020205-0211, 0213-0226, 0228-0231, 0307, 0315, 0345, 04010101-040112, 0403-0410, 0412-0420, 0423-0428, 0430-0435, 0437-0448, 05010101-050110, 0503-0530, 06020101, 06020102, 060203-0610, 0612-0628, 0707, 0709, 0749, 0776, 0810, 1404, 1412, 150201-1506, 1508-1513, 1515-1520, 1522, 1523, 1602, 1609, 1612, 17010101-1705, 1801-1803 |
| 54 | 130 bp | 1070 bp | *1205, 1209, 1221 | *020201-020203, 020205-020207, 0203-0211, 0213-0231, 0410, 0411, 0436, 05010101-050110, 0503-0530, 0605, 080101-0804, 0806-0820, 0822-0824, 0826N-0829, 140203, 1403, 1408, 1410, 150201-150204, 1504-1513, 1515, 1517-1523, 160101, 160103, 1602, 160401, 1606-1614, 1616, 1617 |
| 68 | 150 bp, 415 bp | 1070 bp | *1206, 1208 | *0308, 0329, 0331 |
| 74 | 140 bp | 800 bp | *120402, 1205, 1209, 1221 | *020201-020203, 020205-020207, 0203-0211, 0213-0226, 0228-0231, 04010101-040109, 040111, 040112, 0403-0410, 0412-0420, 0423-0428, 0430-0435, 0437-0448, 05010101-050110, 0503-0530, 0605, 0810, 150201-1506, 1508-1513, 1515-1520, 1522, 1523, 1602, 1609, 1612 |
| 84,9 | 95 bp, 245 bp | 1070 bp | *1207, 1215 |  |
| 94 | 95 bp | 1070 bp | *120201-120402, 1206-1208, 1210-1220, 1222-1227 | *0117, 0121, 0212, 0327, 033801, 033802, 0433, 0707, 0716, 0751, 0805, 0821, 0825, 1404, 1503, 1516, 1615, 17010101-1705, B*0713, B*6702 |
| 10 | 155 bp | 1070 bp | *12030101, 12030102, 120303-120307, 120402-1207, 1211-1213, 1215, 1220, 1223-1225, 1228 |  |
| 11 | 220 bp | 1070 bp | *1209, 1224 | *010201-0103, 0106-0108, 0110-0120, 0123-0133, 0358, 0437, 0516, 0605, 0606, 0812, 140201-1405, 1407N, 1410-1414, B*1403 |
| 124 | 140 bp | 1070 bp | *120201-120303, 120305-120308, 1206-1208, 1210-1220, 1222-1226, 1228 | *0121, 0212w, 0227, 0411, 0429, 0436, 070209, 080101-0809, 0811-0829, 140203, 1403, 1408, 1410, 1507, 1521w, 160101, 160103, 160401, 1606-1608, 1610, 1611, 1613-1617, B*350802, B*6702 |
| 13 | 150 bp | 1070 bp | *1210 |  |
| 144,10 | 100 bp, 145 bp | 1070 bp | *1211, 1216 | *0121, 0605w, 070209, 0814, B*6702 |
| 154,7, 11 | 95 bp, 140 bp | 800 bp | *1212, 1215 | B*350802 |
| 1612 | 185 bp, 225 bp | 1070 bp | *1213-121402 | *0523, 0807, 17010101-1705 |
| 17 | 555 bp | 1070 bp | *120304 |  |
| 184,13 | 140 bp, 245 bp | 1070 bp | *1217, 1227 | *0353, 0412 |
| 194 | 100 bp | 1070 bp | *121402, 1218, 1225 | B*4029, B*4432 |
| 2014 | 170 bp, 220 bp | 1070 bp | *1219, 1222 |  |
| 21 | 250 bp | 1070 bp | *121401, 121402, 1218, 1220 | *0122, 0511, 0517, 0527, 0604, 080101, 080102, 080301-0804, 0806, 0808-0811, 0813, 0814, 0816, 0820-0822, 0824, 0826N, 1406, 150201-1507, 1509-1513, 1515-1523, 17010101-1705, B*5802 |
| 224,15 | 110 bp, 580 bp | 800 bp | *1221, 1223 |  |
| 23 | 150 bp | 1070 bp | *1226 |  |
| 24 | 425 bp | 1070 bp | *1228 | *040105 |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-Cw*12 high resolution SSP typings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective length of the specific PCR product(s) of the alleles amplified by these primer mixes is given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most wells, or a band of 800 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-Cw*12 SSP subtyping.

In addition, wells number 3, 7, 15 and 22 contain the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3The HLA-Cw*12 primer set cannot separate the Cw*1209 and Cw*0516 alleles or the Cw*1216 and Cw*0121 alleles. These alleles can be distinguished by the HLA-Cw low resolution kit and the HLA-Cw*05 or HLA-Cw*01 kit, respectively.

Due to the sharing of sequence motifs between HLA Class I alleles some non-HLA-Cw*12 alleles will be amplified by primer mixes 1 to 7, 9, 11, 12, 14, 16, 18, 21 and 24. In addition, primer mix 1 amplifies the B*0713, B*0715 and B*6702 alleles, primer mix 9 amplifies the B*0713 and B*6702 alleles, primer mix 11 amplifies the B*1403 allele, primer mix 12 amplifies the B*350802 and B*6702 alleles, primer mix 14 amplifies the B*6702, primer mix 15 amplifies the B*350802, primer mix 19 amplifies the B*4029 and B*4432 alleles and primer mix 21 amplifies the B*5802 allele.

4Specific PCR fragments shorter than 150 base pairs have a lower intensity and are less sharp than longer PCR bands.

5Primer mix 1 may yield less specific PCR products than the other Cw*12 primer mixes.

6Primer mixes 2 has a tendency of giving rise to nonspecific amplifications.

7Primer mix 15 may give rise to a long unspecific amplification product. This band should be disregarded when interpreting the typing results.

8Primer mix 6: Specific PCR fragment of 150 bp in the Cw*1206 allele. Specific PCR fragment of 415 bp in the Cw*1208 and Cw*0308, Cw*0329 and Cw*0331 alleles.

9Primer mix 8: Specific PCR fragment of 95 bp in the Cw*1215 allele. Specific PCR fragment of 245 bp in the Cw*1207 allele.

10Primer mix 14: Specific PCR fragment of 100 bp in the Cw*1216 and Cw*0121, Cw*0605w, Cw*070209, Cw*0814 and B*6702 alleles. Specific PCR fragment of 145 bp in the Cw*1211 allele.

11Primer mix 15: Specific PCR fragment of 95 bp in the Cw*1215 allele. Specific PCR fragment of 140 bp in the Cw*1212 and B*350802 alleles.

12Primer mix 16: Specific PCR fragment of 185 bp in the Cw*1213 allele. Specific PCR fragment of 225 bp in Cw*121401 and Cw*121402 and the Cw*0523, Cw*0807 and Cw*17010101-1705alleles.

13Primer mix 18: Specific PCR fragment of 140 bp in the Cw*0412 allele. Specific PCR fragment of 245 bp in Cw*0353 allele. Specific PCR fragments of 140 bp and 245 bp in the Cw*1217 and  Cw*1227 alleles

14Primer mix 20: Specific PCR fragment of 170 bp in the Cw*1222 allele. Specific PCR fragment of 220 bp in Cw*1219  allele.

15Primer mix 22: Specific PCR fragment of 110 bp in the 1223 allele. Specific PCR fragments of 110 bp and 580 bp in Cw* 1221 allele.

‘w’, might be weakly amplified.

1The HLA-Cw*1201 nucleotide sequence has been shown to be identical to Cw*120202.

The HLA-Cw*12 primer set cannot separate the Cw*1209 and Cw*0516 alleles or the Cw*1216 and Cw*0121 alleles. These alleles can be distinguished by the HLA-Cw low resolution kit and the HLA-Cw*05 or HLA-Cw*01 kit, respectively.2Primer mix 6: Specific PCR fragment of 150 bp in the Cw*1206 allele. Specific PCR fragment of 410 bp in the Cw*1208 and Cw*0308, Cw*0329 and Cw*0331 alleles.

Primer mix 6: Specific PCR fragment of 150 bp in the Cw*1206 allele. Specific PCR fragment of 415 bp in the Cw*1208 and Cw*0308, Cw*0329 and Cw*0331 alleles.

Primer mix 8: Specific PCR fragment of 95 bp in the Cw*1215 allele. Specific PCR fragment of 245 bp in the Cw*1207 allele.

Primer mix 14: Specific PCR fragment of 100 bp in the Cw*1216 and Cw*0121, Cw*0605w, Cw*070209, Cw*0814 and B*6702 alleles. Specific PCR fragment of 145 bp in the Cw*1211 allele.

Primer mix 15: Specific PCR fragment of 95 bp in the Cw*1215 allele. Specific PCR fragment of 140 bp in the Cw*1212 and B*350802 alleles.

Primer mix 16: Specific PCR fragment of 185 bp in the Cw*1213 allele. Specific PCR fragment of 225 bp in Cw*121401 and Cw*121402 and the Cw*0523, Cw*0807 and Cw*17010101-1705alleles.

Primer mix 18: Specific PCR fragment of 140 bp in the Cw*0412 allele. Specific PCR fragment of 245 bp in Cw*0353 allele. Specific PCR fragments of 140 bp and 245 bp in the Cw*1217 and  Cw*1227 alleles

Primer mix 20: Specific PCR fragment of 170 bp in the Cw*1222 allele. Specific PCR fragment of 220 bp in Cw*1219  allele.

Primer mix 22: Specific PCR fragment of 110 bp in the 1223 allele. Specific PCR fragments of 110 bp and 580 bp in Cw* 1221 allele.

‘w’, may be weakly amplified.