PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® HLA-Cw*15

Prod. No:	101.626-12/12u

Lot No:    	75F

Expiry Date:    	2011-March-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 16 primer mixes used for HLA-Cw*15 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified HLA-Cw*15 alleles | Other amplified HLA Class I alleles3 |
| --- | --- | --- | --- | --- |
| 1 | 330 bp | 800 bp | 150201-1503, 150501-1513, 1515-1521 | 0206, 1215 |
| 2 | 175 bp | 1070 bp | 150201-1509, 1512, 1513, 1515, 1518, 1519, 1521 | 0308, 0329, 0331, 0720 |
| 3 | 195 bp | 1070 bp | 1503, 1516 | 0212weakly, 0315weakly, 0327weakly, 033801-033802weakly, 0403, 0406, 0416weakly, 0720weakly |
| 4 | 320 bp | 1070 bp | 1504, 1509, 1519, 1520 | 0104, 0121, 020201-020203, 020205-0205, 0208-0218, 0220, 0221, 030201-030203, 0314-0316, 0333, 0336, 0340, 0342, 034301-034302, 0504, 06020101-06020102, 060203-0608, 0610-0619, 070101-0703, 0705-0710, 0714-072702, 0730-0733N, 0735-0740, 0742-0744, 0746-0762, 0764, 0809, 0811, 120201-1214, 1216-1221, 140201-1411, 160101-1602, 160401, 1606-1609, 1611, 1612 |
| 54,5 | 100 bp | 1070 bp | 150501-150504 |  |
| 6 | 320 bp | 1070 bp | 1506 | 010201-010206, 0105-0120, 0122, 0219, 030301-0305, 0307-031102, 0313, 0317, 0318, 0320N-0325, 0327-0332, 0334, 0335, 0337-0339, 0341, 0344-0350, 0729 |
| 74 | 140 bp | 800 bp | 1507, 1521 | 0121, 0212 weakly, 0411, 0429, 0436, 080101-0809, 0811-0819, 120201-120303, 120305-120306, 1206-1208, 1210-1220, 140203, 1403, 1408, 1410, 160101, 160103, 160401, 1606-1608, 1610, 1611 B*350802, B*6702 |
| 86 | 160, 185 bp | 1070 bp | 1508, 1519 | 0206, 1215, B*1318weakly, B*1573weakly, B*5410weakly, B*5509weakly, B*5521weakly |
| 94,7 | 125, 310 bp | 1070 bp | 1511, 1518 | 020201-020203, 020205-0220, 0222, 0223, 0403, 0406 |
| 10 | 165 bp | 1070 bp | 151001-151002 | 0208, 0318, 04010101-040102, 040104-040106, 040401-0405, 0407-041502, 0417-0420, 0423-0437, 05010101-050104, 0503-0523, 080101-0813, 0815-0819, 120201-1215, 1217-1221, 140201-140204, 1404-1409, 1411, 160101-1602, 160401, 1606-1612, 1701-1705 |
| 11 | 320 bp | 800 bp | 150201-1503, 1507, 1508, 151001-1513, 1515-1518, 1521 | 0206, 0312, 0319, 1215 |
| 12 | 240 bp | 1070 bp | 1507, 1521 | 0212, 030201-0306, 0308-0314, 0316, 0317, 0319-033802, 0340-0344, 0346-0350 |
| 13 | 185 bp | 1070 bp | 1512 | 1410 |
| 144,5,8 | 145, 435 bp | 800 bp | 1513, 1516, 1517 | 0319, B*4611, B*4618, B*5614 |
| 154 | 85 bp | 1070 bp | 150201-1506, 1508-151002, 1512, 1513, 1515, 1518, 1519, 1521 | 0329, 0707, 0709 |
| 16 | 165 bp | 1070 bp | 1515 |  |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of Cw*15 high resolution SSP typings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective length of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may  be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most wells, or a band of 800 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-Cw*15 SSP subtyping.

In addition, wells number 7, 11 and 14 contain the primer pair giving rise to the shorter, 800 bp, internal positive control band.

In the presence of a specific amplification the intensity of the control band often decreases.

3Due to the sharing of sequence motifs between HLA Class I alleles some non-HLA-Cw*15 alleles will be amplified by all primer mixes. In addition, the B*350802 and the B*6702 alleles will be amplified by primer mix 7, the B*1318, B*1573, B*5410, B*5509 and B*5521 alleles will be weakly amplified by primer mix 8, the B*4611, B*4618 and B*5614 alleles will be amplified by primer mix 14.

4Short specific PCR fragments are less intense and not as sharp as longer specific bands.

5Primer mixes 5 and 14 may yield somewhat less specific PCR products than the other HLA-Cw*15 primer mixes.

6Primer mix 8: Specific PCR fragment of the 160 bp in the Cw*1508, and the Cw*0206, Cw*1215, B*1318, B*1573, B*5410, B*5509 and B*5521 alleles. Specific PCR fragment of 185 bp in the Cw*1519 allele.

7Primer mix 9: Specific PCR fragment of 125 bp in the Cw*1511 and the Cw*020201-020203, Cw*020205-020206, Cw*0203-0220, Cw*0222, Cw*0223, Cw*0403 and Cw*0406 alleles. Specific PCR fragment of 310 bp in the Cw*1518 allele.

8Primer mix 14: Specific PCR fragment of 145 bp in the Cw*1513 allele. Specific PCR fragment of 435 bp the Cw*1516 and Cw*1517 and in the Cw*0319, B*4611, B*4618 and B*5614 alleles.

1The HLA-Cw*1501 nucleotide sequence has been shown to be identical to Cw*1502.

The HLA-Cw*1514 nucleotide sequence has been renamed to B*151002.

2Primer mix 8: Specific PCR fragment of the 160 bp in the Cw*1508, Cw*0206 and Cw*1215 and the B*1318, B*1573, B*5410, B*5509 and B*5521 alleles. Specifci PCR fragment of 185 bp in the Cw*1519 allele.

Primer mix 9: Specific PCR fragment of 125 bp in the Cw*1511 and the Cw*020201-020203, Cw*020205-020206, Cw*0203-0220, Cw*0222, Cw*0223, Cw*0403 and Cw*0406 alleles. Specific PCR fragment of 310 bp in the Cw*1518 allele.

Primer mix 14: Specific PCR fragment of 145 bp in the Cw*1513 allele. Specific PCR fragment of 435 bp the Cw*1516 and Cw*1517 and in the Cw*0319, B*4611, B*4618 and B*5614 alleles.

‘w’, might be weakly amplified.